## README
Research Data — Cannabinoid HPLC-MS/MS method development and validation (hemaPEN cDBS)
Protti M, Noreen S, Mandrioli R, Di Lecce R, Iattoni CP, Castrovilli MC, Mercolini L
Journal of Cannabis Research
DATASET CONTENTS
Sheet | Description
MRM_Parameters | MRM transitions, collision/cone voltages and retention times for all analytes and IS.
Solvent_Screening | EY%, ME%, IS-ME% for 7 solvent conditions (screening phase, 200 µL per disc). 6 donors × LQC+HQC × 9 analytes.
EY_ME_Optimised | EY%, ME%, IS-ME% for the final validated extraction protocol (100 µL FA/MeOH). 6 donors × LQC+HQC × 9 analytes.
Calibration_Curves | Analyte/IS peak-area ratios vs nominal concentration (7 levels × 5 days) with back-calculated concentrations.
Precision_Accuracy | Measured concentrations at 4 QC levels × 6 replicates × 5 days.
HCT_Independence | IS-normalised EY% at low/normal/high hematocrit (0.26/0.41/0.55), LQC+HQC, 6 donors.
Stability | Measured concentrations and % deviation from freshly prepared reference under all stability conditions (8 conditions × LQC+HQC × 6 replicates × 9 analytes).
Dilution_Integrity | Samples at 3×ULOQ diluted 1:5 with blank extract, 6 replicates, with bias.
Selectivity_Carryover | Interference check on 6 blank donors and carryover after the highest calibrator (and at 5000 ng/mL for 7-COOH-CBD).
Application_hemaPEN_vs_SPE | Authentic whole blood specimens (n=12): hemaPEN cDBS vs SPE plasma-based reference (whole-blood-equivalent), detection frequency, concentration ranges and bias.
NOTE | All peak areas are in instrument counts (Waters Quattro Micro, MassLynx 4.1). Concentrations in ng/mL.
ANALYTES | Protocol / LLOQ (ng/mL)
CBD | 1-disc — LLOQ 2.0 ng/mL
7-OH-CBD | 1-disc — LLOQ 2.0 ng/mL
7-COOH-CBD | 1-disc — LLOQ 5.0 ng/mL
6α-OH-CBD | 4-disc — LLOQ 0.5 ng/mL
CBN | 4-disc — LLOQ 0.5 ng/mL
CBG | 4-disc — LLOQ 0.5 ng/mL
Δ9-THC | 4-disc — LLOQ 0.5 ng/mL
11-OH-THC | 4-disc — LLOQ 0.5 ng/mL
THC-COOH | 1-disc — LLOQ 2.0 ng/mL
INTERNAL STANDARDS
IS4: CBD-d3 | IS for CBD; ESI+; m/z 318.2→196.1/262.2
IS5: 7-OH-CBD-d3 | IS for 7-OH-CBD and 6α-OH-CBD (surrogate); ESI−; m/z 332.2→264.2/271.1
IS6: 7-COOH-CBD-d3 | IS for 7-COOH-CBD; ESI−; m/z 346.1→182.2/234.2
IS7: CBN-d3 | IS for CBN; ESI+; m/z 314.3→223.3/241.2
IS8: CBG-d3 | IS for CBG; ESI+; m/z 320.3→196.1/194.1
IS1: Δ9-THC-d3 | IS for Δ9-THC; ESI+; m/z 318.2→196.2/262.2
IS2: 11-OH-THC-d3 | IS for 11-OH-THC; ESI+; m/z 334.2→316.3/196.1
IS3: THC-COOH-d3 | IS for THC-COOH; ESI−; m/z 346.4→300.2/247.9

## MRM_Parameters
MRM transitions & MS parameters (Table 1). CE = quantifier collision energy; qualifier in same function.
Compound | Polarity | Precursor (m/z) | Quantifier (m/z) | Qualifier (m/z) | CE (eV) | CV (V) | RT (min)
Δ9-THC | ESI+ | 315.2 | 193.2 | 259.2 | 30 | 30 | 11.3
11-OH-THC | ESI+ | 331.2 | 313.3 | 193.1 | 20 | 20 | 8.1
THC-COOH | ESI− | 343.4 | 297.2 | 244.9 | 30 | 30 | 7.7
CBD | ESI+ | 315.2 | 193 | 259.2 | 30 | 50 | 9.8
7-OH-CBD | ESI− | 329.2 | 261.2 | 268.1 | 20 | 20 | 6.5
7-COOH-CBD | ESI− | 343.1 | 179.2 | 231.2 | 20 | 20 | 5.6
6α-OH-CBD | ESI− | 329.2 | 158.2 | 173.1 | 30 | 30 | 6.2
CBN | ESI+ | 311.3 | 223.3 | 241.2 | 22 | 42 | 12.1
CBG | ESI+ | 317.3 | 193.1 | 191.1 | 30 | 40 | 10.2

## Solvent_Screening
Solvent Screening — Extraction Yield and Matrix Effect (6 Donors, LQC and HQC, Sets A/B/C, 200 µL extraction volume)
Solvent | Analyte | QC Level | Donor | Set A Area (cts) | Set B Area (cts) | Set C Area (cts) | EY% | ME% | IS Area (cts) | IS-ME%
FA/MeOH 0.1% | CBD | LQC | 1 | 1604 | 1406 | 1574 | 89.3 | 101.9 | 48715 | 92.6
FA/MeOH 0.1% | CBD | LQC | 2 | 1366 | 1415 | 1565 | 90.4 | 87.3 | 49105 | 76.5
FA/MeOH 0.1% | CBD | LQC | 3 | 1398 | 1437 | 1548 | 92.8 | 90.3 | 45516 | 99.4
FA/MeOH 0.1% | CBD | LQC | 4 | 1337 | 1409 | 1588 | 88.7 | 84.2 | 47674 | 108.7
FA/MeOH 0.1% | CBD | LQC | 5 | 1627 | 1426 | 1583 | 90.1 | 102.8 | 47135 | 102.8
FA/MeOH 0.1% | CBD | LQC | 6 | 1293 | 1508 | 1591 | 94.8 | 81.3 | 46241 | 110.1
FA/MeOH 0.1% | CBD | HQC | 1 | 36179 | 37411 | 41753 | 89.6 | 86.7 | 47833 | 100
FA/MeOH 0.1% | CBD | HQC | 2 | 37288 | 39368 | 42885 | 91.8 | 86.9 | 45461 | 112.3
FA/MeOH 0.1% | CBD | HQC | 3 | 40808 | 40105 | 41733 | 96.1 | 97.8 | 46791 | 97.4
FA/MeOH 0.1% | CBD | HQC | 4 | 39222 | 40268 | 43299 | 93 | 90.6 | 46277 | 110.8
FA/MeOH 0.1% | CBD | HQC | 5 | 41231 | 34973 | 39473 | 88.6 | 104.5 | 48520 | 106.8
FA/MeOH 0.1% | CBD | HQC | 6 | 38952 | 37034 | 42617 | 86.9 | 91.4 | 47569 | 104.8
FA/MeOH 0.1% | 7-OH-CBD | LQC | 1 | 1409 | 1280 | 1437 | 89.1 | 98.1 | 42932 | 103.6
FA/MeOH 0.1% | 7-OH-CBD | LQC | 2 | 1308 | 1343 | 1403 | 95.7 | 93.2 | 45278 | 102.2
FA/MeOH 0.1% | 7-OH-CBD | LQC | 3 | 1313 | 1345 | 1443 | 93.2 | 91 | 44006 | 100.4
FA/MeOH 0.1% | 7-OH-CBD | LQC | 4 | 1350 | 1418 | 1453 | 97.6 | 92.9 | 44533 | 99.4
FA/MeOH 0.1% | 7-OH-CBD | LQC | 5 | 1377 | 1350 | 1455 | 92.8 | 94.6 | 44079 | 102.3
FA/MeOH 0.1% | 7-OH-CBD | LQC | 6 | 1473 | 1356 | 1405 | 96.5 | 104.8 | 44992 | 108.8
FA/MeOH 0.1% | 7-OH-CBD | HQC | 1 | 34405 | 33746 | 37330 | 90.4 | 92.2 | 42692 | 117.4
FA/MeOH 0.1% | 7-OH-CBD | HQC | 2 | 35367 | 36712 | 38523 | 95.3 | 91.8 | 44625 | 104.7
FA/MeOH 0.1% | 7-OH-CBD | HQC | 3 | 38199 | 36391 | 39088 | 93.1 | 97.7 | 44300 | 88.5
FA/MeOH 0.1% | 7-OH-CBD | HQC | 4 | 37334 | 34785 | 39172 | 88.8 | 95.3 | 42256 | 96
FA/MeOH 0.1% | 7-OH-CBD | HQC | 5 | 37403 | 34627 | 37679 | 91.9 | 99.3 | 44386 | 90.3
FA/MeOH 0.1% | 7-OH-CBD | HQC | 6 | 40262 | 36411 | 39236 | 92.8 | 102.6 | 45038 | 102.7
FA/MeOH 0.1% | 7-COOH-CBD | LQC | 1 | 4009 | 3993 | 4307 | 92.7 | 93.1 | 51617 | 100.6
FA/MeOH 0.1% | 7-COOH-CBD | LQC | 2 | 4454 | 4254 | 4609 | 92.3 | 96.6 | 53290 | 91.4
FA/MeOH 0.1% | 7-COOH-CBD | LQC | 3 | 4393 | 4004 | 4366 | 91.7 | 100.6 | 50329 | 99.7
FA/MeOH 0.1% | 7-COOH-CBD | LQC | 4 | 4114 | 4018 | 4181 | 96.1 | 98.4 | 49636 | 101
FA/MeOH 0.1% | 7-COOH-CBD | LQC | 5 | 3992 | 4010 | 4354 | 92.1 | 91.7 | 50621 | 112.7
FA/MeOH 0.1% | 7-COOH-CBD | LQC | 6 | 4402 | 4166 | 4371 | 95.3 | 100.7 | 53643 | 109.2
FA/MeOH 0.1% | 7-COOH-CBD | HQC | 1 | 116426 | 105730 | 114179 | 92.6 | 102 | 52890 | 102
FA/MeOH 0.1% | 7-COOH-CBD | HQC | 2 | 121775 | 110262 | 109713 | 100.5 | 111 | 51630 | 105.1
FA/MeOH 0.1% | 7-COOH-CBD | HQC | 3 | 104295 | 111220 | 113028 | 98.4 | 92.3 | 51263 | 90.4
FA/MeOH 0.1% | 7-COOH-CBD | HQC | 4 | 121780 | 108960 | 116910 | 93.2 | 104.2 | 52337 | 97.5
FA/MeOH 0.1% | 7-COOH-CBD | HQC | 5 | 114141 | 108964 | 116539 | 93.5 | 97.9 | 50967 | 92.2
FA/MeOH 0.1% | 7-COOH-CBD | HQC | 6 | 120689 | 107067 | 116251 | 92.1 | 103.8 | 51478 | 101.2
FA/MeOH 0.1% | 6α-OH-CBD | LQC | 1 | 286 | 296 | 326 | 90.8 | 87.7 | 40419 | 95.6
FA/MeOH 0.1% | 6α-OH-CBD | LQC | 2 | 212 | 281 | 308 | 91.2 | 68.8 | 38883 | 107.5
FA/MeOH 0.1% | 6α-OH-CBD | LQC | 3 | 297 | 298 | 327 | 91 | 90.8 | 39860 | 118.7
FA/MeOH 0.1% | 6α-OH-CBD | LQC | 4 | 293 | 271 | 311 | 87.2 | 94.2 | 38011 | 109.1
FA/MeOH 0.1% | 6α-OH-CBD | LQC | 5 | 268 | 288 | 333 | 86.5 | 80.5 | 38129 | 102.1
FA/MeOH 0.1% | 6α-OH-CBD | LQC | 6 | 281 | 264 | 312 | 84.6 | 90.1 | 38670 | 96.8
FA/MeOH 0.1% | 6α-OH-CBD | HQC | 1 | 16564 | 15079 | 17213 | 87.6 | 96.2 | 38406 | 97
FA/MeOH 0.1% | 6α-OH-CBD | HQC | 2 | 16075 | 14466 | 16919 | 85.5 | 95 | 37361 | 84.6
FA/MeOH 0.1% | 6α-OH-CBD | HQC | 3 | 17637 | 15066 | 17218 | 87.5 | 102.4 | 38623 | 106.4
FA/MeOH 0.1% | 6α-OH-CBD | HQC | 4 | 16032 | 15756 | 16815 | 93.7 | 95.3 | 38023 | 98.7
FA/MeOH 0.1% | 6α-OH-CBD | HQC | 5 | 14804 | 15258 | 17358 | 87.9 | 85.3 | 37067 | 89.5
FA/MeOH 0.1% | 6α-OH-CBD | HQC | 6 | 15783 | 14677 | 16584 | 88.5 | 95.2 | 38716 | 103.3
FA/MeOH 0.1% | CBN | LQC | 1 | 289 | 282 | 327 | 86.2 | 88.4 | 39009 | 101.9
FA/MeOH 0.1% | CBN | LQC | 2 | 259 | 288 | 322 | 89.4 | 80.4 | 39391 | 109.4
FA/MeOH 0.1% | CBN | LQC | 3 | 296 | 307 | 333 | 92.3 | 88.9 | 40373 | 118.4
FA/MeOH 0.1% | CBN | LQC | 4 | 330 | 301 | 346 | 87.1 | 95.4 | 39653 | 98.7
FA/MeOH 0.1% | CBN | LQC | 5 | 302 | 304 | 334 | 90.9 | 90.4 | 40828 | 101.6
FA/MeOH 0.1% | CBN | LQC | 6 | 355 | 313 | 327 | 95.6 | 108.6 | 40597 | 91.3
FA/MeOH 0.1% | CBN | HQC | 1 | 14656 | 14909 | 16865 | 88.4 | 86.9 | 42627 | 97.9
FA/MeOH 0.1% | CBN | HQC | 2 | 16999 | 15114 | 17554 | 86.1 | 96.8 | 40180 | 108.4
FA/MeOH 0.1% | CBN | HQC | 3 | 15537 | 15070 | 17048 | 88.4 | 91.1 | 38415 | 92.4
FA/MeOH 0.1% | CBN | HQC | 4 | 16264 | 15853 | 17832 | 88.9 | 91.2 | 41485 | 95.6
FA/MeOH 0.1% | CBN | HQC | 5 | 16937 | 16152 | 18128 | 89.1 | 93.4 | 38718 | 89.9
FA/MeOH 0.1% | CBN | HQC | 6 | 18794 | 16147 | 17862 | 90.4 | 105.2 | 41533 | 92
FA/MeOH 0.1% | CBG | LQC | 1 | 300 | 256 | 297 | 86.3 | 101 | 37163 | 107
FA/MeOH 0.1% | CBG | LQC | 2 | 218 | 242 | 278 | 87.1 | 78.4 | 35720 | 101
FA/MeOH 0.1% | CBG | LQC | 3 | 215 | 228 | 280 | 81.3 | 76.8 | 36577 | 110.1
FA/MeOH 0.1% | CBG | LQC | 4 | 232 | 257 | 297 | 86.5 | 78.1 | 34965 | 96
FA/MeOH 0.1% | CBG | LQC | 5 | 282 | 243 | 285 | 85.3 | 98.9 | 34460 | 104.8
FA/MeOH 0.1% | CBG | LQC | 6 | 249 | 251 | 284 | 88.3 | 87.7 | 34167 | 101.7
FA/MeOH 0.1% | CBG | HQC | 1 | 14114 | 13157 | 15871 | 82.9 | 88.9 | 34265 | 82.3
FA/MeOH 0.1% | CBG | HQC | 2 | 12673 | 13934 | 16165 | 86.2 | 78.4 | 34395 | 95.5
FA/MeOH 0.1% | CBG | HQC | 3 | 15240 | 13259 | 15581 | 85.1 | 97.8 | 35605 | 95.8
FA/MeOH 0.1% | CBG | HQC | 4 | 13167 | 12988 | 15208 | 85.4 | 86.6 | 35213 | 83.7
FA/MeOH 0.1% | CBG | HQC | 5 | 13418 | 14685 | 16226 | 90.5 | 82.7 | 34991 | 112.8
FA/MeOH 0.1% | CBG | HQC | 6 | 15596 | 14690 | 16377 | 89.7 | 95.2 | 34800 | 121.7
FA/MeOH 0.1% | Δ9-THC | LQC | 1 | 327 | 350 | 378 | 92.6 | 86.5 | 44121 | 81.6
FA/MeOH 0.1% | Δ9-THC | LQC | 2 | 350 | 336 | 385 | 87.3 | 90.9 | 44932 | 102.1
FA/MeOH 0.1% | Δ9-THC | LQC | 3 | 366 | 337 | 371 | 90.9 | 98.7 | 45134 | 96.1
FA/MeOH 0.1% | Δ9-THC | LQC | 4 | 357 | 335 | 369 | 90.8 | 96.7 | 42563 | 115.9
FA/MeOH 0.1% | Δ9-THC | LQC | 5 | 344 | 330 | 369 | 89.4 | 93.2 | 44910 | 108.1
FA/MeOH 0.1% | Δ9-THC | LQC | 6 | 362 | 335 | 391 | 85.6 | 92.6 | 44221 | 102.1
FA/MeOH 0.1% | Δ9-THC | HQC | 1 | 18091 | 17210 | 20200 | 85.2 | 89.6 | 45779 | 102.2
FA/MeOH 0.1% | Δ9-THC | HQC | 2 | 18211 | 18283 | 19938 | 91.7 | 91.3 | 45343 | 94.8
FA/MeOH 0.1% | Δ9-THC | HQC | 3 | 18884 | 19059 | 20761 | 91.8 | 91 | 45391 | 96.6
FA/MeOH 0.1% | Δ9-THC | HQC | 4 | 19875 | 19531 | 21001 | 93 | 94.6 | 43639 | 94.3
FA/MeOH 0.1% | Δ9-THC | HQC | 5 | 19672 | 17975 | 19411 | 92.6 | 101.3 | 46119 | 103.8
FA/MeOH 0.1% | Δ9-THC | HQC | 6 | 19114 | 18334 | 20281 | 90.4 | 94.2 | 44439 | 99.5
FA/MeOH 0.1% | 11-OH-THC | LQC | 1 | 334 | 314 | 358 | 87.7 | 93.3 | 43046 | 98.5
FA/MeOH 0.1% | 11-OH-THC | LQC | 2 | 336 | 344 | 357 | 96.4 | 94.1 | 41357 | 89.4
FA/MeOH 0.1% | 11-OH-THC | LQC | 3 | 358 | 310 | 346 | 89.6 | 103.5 | 43095 | 95.7
FA/MeOH 0.1% | 11-OH-THC | LQC | 4 | 349 | 324 | 343 | 94.5 | 101.7 | 43885 | 100
FA/MeOH 0.1% | 11-OH-THC | LQC | 5 | 338 | 321 | 346 | 92.9 | 97.7 | 42137 | 93.6
FA/MeOH 0.1% | 11-OH-THC | LQC | 6 | 316 | 328 | 352 | 93.1 | 89.8 | 43695 | 113.8
FA/MeOH 0.1% | 11-OH-THC | HQC | 1 | 16543 | 17176 | 18390 | 93.4 | 90 | 39867 | 99.4
FA/MeOH 0.1% | 11-OH-THC | HQC | 2 | 18294 | 17984 | 19655 | 91.5 | 93.1 | 42092 | 104.7
FA/MeOH 0.1% | 11-OH-THC | HQC | 3 | 18104 | 17596 | 18759 | 93.8 | 96.5 | 41690 | 106.2
FA/MeOH 0.1% | 11-OH-THC | HQC | 4 | 20852 | 17989 | 18896 | 95.2 | 110.4 | 43419 | 111.2
FA/MeOH 0.1% | 11-OH-THC | HQC | 5 | 17220 | 17023 | 18852 | 90.3 | 91.3 | 42249 | 109.2
FA/MeOH 0.1% | 11-OH-THC | HQC | 6 | 15937 | 17076 | 19382 | 88.1 | 82.2 | 42134 | 101.5
FA/MeOH 0.1% | THC-COOH | LQC | 1 | 1611 | 1575 | 1642 | 95.9 | 98.1 | 49485 | 85.6
FA/MeOH 0.1% | THC-COOH | LQC | 2 | 1602 | 1582 | 1710 | 92.5 | 93.7 | 51907 | 96.7
FA/MeOH 0.1% | THC-COOH | LQC | 3 | 1725 | 1513 | 1674 | 90.4 | 103 | 49639 | 99.9
FA/MeOH 0.1% | THC-COOH | LQC | 4 | 1669 | 1505 | 1636 | 92 | 102 | 50923 | 97.5
FA/MeOH 0.1% | THC-COOH | LQC | 5 | 1506 | 1594 | 1623 | 98.2 | 92.8 | 49215 | 97.8
FA/MeOH 0.1% | THC-COOH | LQC | 6 | 1660 | 1670 | 1685 | 99.1 | 98.5 | 52632 | 97.7
FA/MeOH 0.1% | THC-COOH | HQC | 1 | 44297 | 42726 | 44646 | 95.7 | 99.2 | 49943 | 100.5
FA/MeOH 0.1% | THC-COOH | HQC | 2 | 44610 | 42023 | 44848 | 93.7 | 99.5 | 47434 | 103.6
FA/MeOH 0.1% | THC-COOH | HQC | 3 | 41250 | 44878 | 45515 | 98.6 | 90.6 | 49190 | 100.5
FA/MeOH 0.1% | THC-COOH | HQC | 4 | 47490 | 43980 | 46198 | 95.2 | 102.8 | 50708 | 110.1
FA/MeOH 0.1% | THC-COOH | HQC | 5 | 42379 | 42550 | 45266 | 94 | 93.6 | 49062 | 97.5
FA/MeOH 0.1% | THC-COOH | HQC | 6 | 42568 | 45456 | 46242 | 98.3 | 92.1 | 47890 | 107.3
MeOH | CBD | LQC | 1 | 1516 | 1248 | 1561 | 82.3 | 97.1 | 48710 | 105.9
MeOH | CBD | LQC | 2 | 1417 | 1146 | 1654 | 80.9 | 85.7 | 47861 | 119.6
MeOH | CBD | LQC | 3 | 1567 | 1190 | 1608 | 75.9 | 97.5 | 49019 | 106
MeOH | CBD | LQC | 4 | 1384 | 1087 | 1531 | 78.5 | 90.4 | 47643 | 112.8
MeOH | CBD | LQC | 5 | 1430 | 1118 | 1581 | 78.2 | 90.4 | 48806 | 87.3
MeOH | CBD | LQC | 6 | 1176 | 925 | 1625 | 78.7 | 72.4 | 48299 | 121.2
MeOH | CBD | HQC | 1 | 33651 | 25762 | 42341 | 76.6 | 79.5 | 48040 | 83.2
MeOH | CBD | HQC | 2 | 43003 | 32918 | 41295 | 76.5 | 104.1 | 47554 | 107.8
MeOH | CBD | HQC | 3 | 28961 | 22461 | 42522 | 77.6 | 68.1 | 50083 | 82.4
MeOH | CBD | HQC | 4 | 49298 | 39543 | 41856 | 80.2 | 117.8 | 49534 | 83.8
MeOH | CBD | HQC | 5 | 33431 | 25226 | 40820 | 75.5 | 81.9 | 50274 | 87.2
MeOH | CBD | HQC | 6 | 45596 | 35102 | 41185 | 77 | 110.7 | 46894 | 102.2
MeOH | 7-OH-CBD | LQC | 1 | 1409 | 1138 | 1462 | 80.8 | 96.4 | 43934 | 105.6
MeOH | 7-OH-CBD | LQC | 2 | 1201 | 991 | 1436 | 82.5 | 83.6 | 47240 | 99.7
MeOH | 7-OH-CBD | LQC | 3 | 1402 | 1130 | 1436 | 80.6 | 97.6 | 45875 | 100.8
MeOH | 7-OH-CBD | LQC | 4 | 1230 | 994 | 1442 | 80.8 | 85.3 | 41999 | 99.3
MeOH | 7-OH-CBD | LQC | 5 | 1148 | 919 | 1406 | 80.1 | 81.7 | 44869 | 106.6
MeOH | 7-OH-CBD | LQC | 6 | 1436 | 1136 | 1499 | 79.1 | 95.8 | 46368 | 97.8
MeOH | 7-OH-CBD | HQC | 1 | 43442 | 35354 | 39403 | 81.4 | 110.3 | 43353 | 101.1
MeOH | 7-OH-CBD | HQC | 2 | 42205 | 34397 | 39971 | 81.5 | 105.6 | 45463 | 100.9
MeOH | 7-OH-CBD | HQC | 3 | 32161 | 25811 | 39194 | 80.3 | 82.1 | 43779 | 96.9
MeOH | 7-OH-CBD | HQC | 4 | 30826 | 26059 | 39560 | 84.5 | 77.9 | 45687 | 100.4
MeOH | 7-OH-CBD | HQC | 5 | 42147 | 34315 | 38743 | 81.4 | 108.8 | 43593 | 99.9
MeOH | 7-OH-CBD | HQC | 6 | 39010 | 30738 | 37839 | 78.8 | 103.1 | 44246 | 98.1
MeOH | 7-COOH-CBD | LQC | 1 | 4334 | 3572 | 4484 | 82.4 | 96.7 | 53864 | 97.1
MeOH | 7-COOH-CBD | LQC | 2 | 4064 | 3457 | 4261 | 85.1 | 95.4 | 54934 | 102.4
MeOH | 7-COOH-CBD | LQC | 3 | 4442 | 3884 | 4385 | 87.4 | 101.3 | 52380 | 102.2
MeOH | 7-COOH-CBD | LQC | 4 | 4007 | 3379 | 4341 | 84.3 | 92.3 | 50330 | 97.4
MeOH | 7-COOH-CBD | LQC | 5 | 4560 | 3866 | 4352 | 84.8 | 104.8 | 50047 | 104.6
MeOH | 7-COOH-CBD | LQC | 6 | 4168 | 3609 | 4392 | 86.6 | 94.9 | 52698 | 98
MeOH | 7-COOH-CBD | HQC | 1 | 90625 | 77354 | 115454 | 85.4 | 78.5 | 50944 | 105.5
MeOH | 7-COOH-CBD | HQC | 2 | 93060 | 79277 | 108112 | 85.2 | 86.1 | 50981 | 104.6
MeOH | 7-COOH-CBD | HQC | 3 | 104255 | 90432 | 115076 | 86.7 | 90.6 | 53225 | 104.3
MeOH | 7-COOH-CBD | HQC | 4 | 103438 | 84999 | 109212 | 82.2 | 94.7 | 53368 | 99.3
MeOH | 7-COOH-CBD | HQC | 5 | 113939 | 98585 | 113443 | 86.5 | 100.4 | 50437 | 94.1
MeOH | 7-COOH-CBD | HQC | 6 | 118117 | 100313 | 116218 | 84.9 | 101.6 | 53442 | 102.2
MeOH | 6α-OH-CBD | LQC | 1 | 298 | 232 | 332 | 77.9 | 89.8 | 39196 | 94.7
MeOH | 6α-OH-CBD | LQC | 2 | 272 | 200 | 319 | 73.5 | 85.3 | 38225 | 99.5
MeOH | 6α-OH-CBD | LQC | 3 | 298 | 226 | 316 | 75.8 | 94.3 | 37972 | 101.2
MeOH | 6α-OH-CBD | LQC | 4 | 297 | 230 | 321 | 77.4 | 92.5 | 36834 | 104.5
MeOH | 6α-OH-CBD | LQC | 5 | 332 | 255 | 329 | 76.8 | 100.9 | 36602 | 95
MeOH | 6α-OH-CBD | LQC | 6 | 302 | 233 | 303 | 77.2 | 99.7 | 37956 | 96.9
MeOH | 6α-OH-CBD | HQC | 1 | 15088 | 11278 | 16638 | 74.7 | 90.7 | 37286 | 99.7
MeOH | 6α-OH-CBD | HQC | 2 | 15777 | 11287 | 15669 | 71.5 | 100.7 | 39336 | 102.8
MeOH | 6α-OH-CBD | HQC | 3 | 17357 | 12967 | 16673 | 74.7 | 104.1 | 37586 | 100.5
MeOH | 6α-OH-CBD | HQC | 4 | 16351 | 12099 | 16662 | 74 | 98.1 | 38543 | 104.2
MeOH | 6α-OH-CBD | HQC | 5 | 14666 | 11074 | 16734 | 75.5 | 87.6 | 39710 | 103.6
MeOH | 6α-OH-CBD | HQC | 6 | 14298 | 10684 | 16787 | 74.7 | 85.2 | 38490 | 102.4
MeOH | CBN | LQC | 1 | 282 | 220 | 334 | 78 | 84.4 | 40873 | 98
MeOH | CBN | LQC | 2 | 388 | 288 | 332 | 74.2 | 116.9 | 40415 | 112.8
MeOH | CBN | LQC | 3 | 374 | 291 | 332 | 77.8 | 112.7 | 40735 | 98.7
MeOH | CBN | LQC | 4 | 395 | 309 | 333 | 78.2 | 118.6 | 40594 | 99.4
MeOH | CBN | LQC | 5 | 254 | 195 | 347 | 76.8 | 73.2 | 41326 | 101.6
MeOH | CBN | LQC | 6 | 323 | 249 | 337 | 77.1 | 95.8 | 39094 | 98.8
MeOH | CBN | HQC | 1 | 13591 | 10292 | 18077 | 75.7 | 75.2 | 39707 | 99.3
MeOH | CBN | HQC | 2 | 15160 | 11150 | 17818 | 73.5 | 85.1 | 41650 | 99.2
MeOH | CBN | HQC | 3 | 10727 | 8520 | 17526 | 79.4 | 61.2 | 41913 | 100.4
MeOH | CBN | HQC | 4 | 16960 | 13681 | 17666 | 80.7 | 96 | 41211 | 91.7
MeOH | CBN | HQC | 5 | 22141 | 17144 | 18427 | 77.4 | 120.2 | 38332 | 92.2
MeOH | CBN | HQC | 6 | 15828 | 11992 | 17783 | 75.8 | 89 | 41024 | 101.4
MeOH | CBG | LQC | 1 | 205 | 151 | 294 | 73.7 | 69.7 | 33222 | 102.8
MeOH | CBG | LQC | 2 | 205 | 144 | 291 | 70.2 | 70.4 | 33350 | 94.4
MeOH | CBG | LQC | 3 | 256 | 184 | 281 | 71.9 | 91.1 | 35110 | 97.8
MeOH | CBG | LQC | 4 | 253 | 184 | 289 | 72.7 | 87.5 | 35580 | 96.7
MeOH | CBG | LQC | 5 | 218 | 152 | 282 | 69.7 | 77.3 | 35051 | 102.3
MeOH | CBG | LQC | 6 | 240 | 174 | 302 | 72.5 | 79.5 | 34272 | 95.3
MeOH | CBG | HQC | 1 | 14456 | 10374 | 15878 | 71.8 | 91 | 34122 | 110.4
MeOH | CBG | HQC | 2 | 16415 | 11429 | 16494 | 69.6 | 99.5 | 34863 | 100.6
MeOH | CBG | HQC | 3 | 17930 | 13111 | 16286 | 73.1 | 110.1 | 34907 | 110.4
MeOH | CBG | HQC | 4 | 11457 | 8007 | 15234 | 69.9 | 75.2 | 36709 | 99.5
MeOH | CBG | HQC | 5 | 9739 | 6671 | 15760 | 68.5 | 61.8 | 35035 | 97
MeOH | CBG | HQC | 6 | 17418 | 12540 | 15794 | 72 | 110.3 | 35392 | 107.6
MeOH | Δ9-THC | LQC | 1 | 279 | 213 | 379 | 76.3 | 73.6 | 45753 | 117.7
MeOH | Δ9-THC | LQC | 2 | 373 | 284 | 383 | 76.1 | 97.4 | 43779 | 91
MeOH | Δ9-THC | LQC | 3 | 344 | 265 | 379 | 77 | 90.8 | 46637 | 88.2
MeOH | Δ9-THC | LQC | 4 | 272 | 211 | 365 | 77.6 | 74.5 | 44794 | 74.3
MeOH | Δ9-THC | LQC | 5 | 445 | 342 | 377 | 76.9 | 118 | 43510 | 101.1
MeOH | Δ9-THC | LQC | 6 | 337 | 265 | 380 | 78.6 | 88.7 | 44866 | 111.3
MeOH | Δ9-THC | HQC | 1 | 15373 | 12126 | 20170 | 78.9 | 76.2 | 44763 | 113.9
MeOH | Δ9-THC | HQC | 2 | 21084 | 15661 | 20437 | 74.3 | 103.2 | 45640 | 91.2
MeOH | Δ9-THC | HQC | 3 | 17256 | 13229 | 20184 | 76.7 | 85.5 | 43562 | 106.7
MeOH | Δ9-THC | HQC | 4 | 14845 | 11571 | 20083 | 77.9 | 73.9 | 45656 | 88.9
MeOH | Δ9-THC | HQC | 5 | 20964 | 15853 | 20231 | 75.6 | 103.6 | 47066 | 84.9
MeOH | Δ9-THC | HQC | 6 | 18033 | 14261 | 20157 | 79.1 | 89.5 | 44591 | 116.9
MeOH | 11-OH-THC | LQC | 1 | 338 | 260 | 341 | 76.9 | 99.1 | 43196 | 103.5
MeOH | 11-OH-THC | LQC | 2 | 275 | 220 | 330 | 80 | 83.3 | 39729 | 99.4
MeOH | 11-OH-THC | LQC | 3 | 333 | 272 | 354 | 81.7 | 94.1 | 40136 | 100.6
MeOH | 11-OH-THC | LQC | 4 | 295 | 244 | 345 | 82.7 | 85.5 | 41800 | 102.8
MeOH | 11-OH-THC | LQC | 5 | 312 | 243 | 347 | 77.9 | 89.9 | 40271 | 91.1
MeOH | 11-OH-THC | LQC | 6 | 395 | 315 | 348 | 79.7 | 113.5 | 44300 | 103.5
MeOH | 11-OH-THC | HQC | 1 | 15570 | 12984 | 19085 | 83.4 | 81.6 | 43640 | 104.6
MeOH | 11-OH-THC | HQC | 2 | 14462 | 11294 | 18409 | 78.1 | 78.6 | 42172 | 98.4
MeOH | 11-OH-THC | HQC | 3 | 18577 | 14225 | 19337 | 76.6 | 96.1 | 40680 | 98.1
MeOH | 11-OH-THC | HQC | 4 | 20016 | 16752 | 19130 | 83.7 | 104.6 | 43891 | 99.8
MeOH | 11-OH-THC | HQC | 5 | 18404 | 14809 | 19589 | 80.5 | 94 | 41189 | 102.6
MeOH | 11-OH-THC | HQC | 6 | 14898 | 11502 | 17940 | 77.2 | 83 | 43351 | 101.5
MeOH | THC-COOH | LQC | 1 | 1723 | 1444 | 1725 | 83.8 | 99.9 | 49668 | 98.3
MeOH | THC-COOH | LQC | 2 | 1713 | 1484 | 1720 | 86.6 | 99.6 | 50423 | 94.7
MeOH | THC-COOH | LQC | 3 | 1493 | 1296 | 1667 | 86.8 | 89.6 | 50271 | 95
MeOH | THC-COOH | LQC | 4 | 1749 | 1407 | 1799 | 80.4 | 97.2 | 50031 | 96
MeOH | THC-COOH | LQC | 5 | 1639 | 1372 | 1683 | 83.7 | 97.4 | 51130 | 97.4
MeOH | THC-COOH | LQC | 6 | 1804 | 1535 | 1706 | 85.1 | 105.7 | 48724 | 96.6
MeOH | THC-COOH | HQC | 1 | 42159 | 35390 | 44171 | 83.9 | 95.4 | 51794 | 98.6
MeOH | THC-COOH | HQC | 2 | 46301 | 38595 | 44636 | 83.4 | 103.7 | 50629 | 98.6
MeOH | THC-COOH | HQC | 3 | 35128 | 29169 | 43337 | 83 | 81.1 | 50337 | 107.4
MeOH | THC-COOH | HQC | 4 | 52847 | 45841 | 47315 | 86.7 | 111.7 | 53837 | 97.8
MeOH | THC-COOH | HQC | 5 | 39449 | 33074 | 44134 | 83.8 | 89.4 | 51536 | 99.5
MeOH | THC-COOH | HQC | 6 | 43304 | 36285 | 46720 | 83.8 | 92.7 | 49320 | 99.2
NH4OH/MeOH 0.1% | CBD | LQC | 1 | 1309 | 971 | 1643 | 74.2 | 79.7 | 46967 | 94.6
NH4OH/MeOH 0.1% | CBD | LQC | 2 | 1186 | 928 | 1546 | 78.2 | 76.7 | 50669 | 101.9
NH4OH/MeOH 0.1% | CBD | LQC | 3 | 996 | 765 | 1589 | 76.8 | 62.7 | 50713 | 106.2
NH4OH/MeOH 0.1% | CBD | LQC | 4 | 1333 | 942 | 1562 | 70.7 | 85.3 | 47787 | 103.6
NH4OH/MeOH 0.1% | CBD | LQC | 5 | 1167 | 867 | 1560 | 74.3 | 74.8 | 47960 | 100.3
NH4OH/MeOH 0.1% | CBD | LQC | 6 | 1222 | 911 | 1575 | 74.5 | 77.6 | 48434 | 98.7
NH4OH/MeOH 0.1% | CBD | HQC | 1 | 30234 | 23026 | 43816 | 76.2 | 69 | 47673 | 98.4
NH4OH/MeOH 0.1% | CBD | HQC | 2 | 38012 | 29338 | 41942 | 77.2 | 90.6 | 46924 | 90.3
NH4OH/MeOH 0.1% | CBD | HQC | 3 | 26157 | 20164 | 42549 | 77.1 | 61.5 | 45749 | 96.1
NH4OH/MeOH 0.1% | CBD | HQC | 4 | 29218 | 22903 | 41384 | 78.4 | 70.6 | 47948 | 91.9
NH4OH/MeOH 0.1% | CBD | HQC | 5 | 37864 | 28255 | 42426 | 74.6 | 89.2 | 46537 | 98.9
NH4OH/MeOH 0.1% | CBD | HQC | 6 | 35551 | 26406 | 43194 | 74.3 | 82.3 | 48124 | 106.1
NH4OH/MeOH 0.1% | 7-OH-CBD | LQC | 1 | 1144 | 898 | 1449 | 78.5 | 79 | 44817 | 106.1
NH4OH/MeOH 0.1% | 7-OH-CBD | LQC | 2 | 1309 | 1037 | 1515 | 79.2 | 86.4 | 42945 | 101.5
NH4OH/MeOH 0.1% | 7-OH-CBD | LQC | 3 | 1291 | 1037 | 1449 | 80.3 | 89.1 | 43614 | 92.9
NH4OH/MeOH 0.1% | 7-OH-CBD | LQC | 4 | 1225 | 961 | 1403 | 78.4 | 87.3 | 43007 | 95.4
NH4OH/MeOH 0.1% | 7-OH-CBD | LQC | 5 | 1456 | 1175 | 1458 | 80.7 | 99.9 | 46472 | 104.6
NH4OH/MeOH 0.1% | 7-OH-CBD | LQC | 6 | 1145 | 923 | 1441 | 80.6 | 79.5 | 42882 | 95.1
NH4OH/MeOH 0.1% | 7-OH-CBD | HQC | 1 | 38962 | 30265 | 38758 | 77.7 | 100.5 | 46722 | 96.7
NH4OH/MeOH 0.1% | 7-OH-CBD | HQC | 2 | 40672 | 32720 | 41350 | 80.4 | 98.4 | 44565 | 96.6
NH4OH/MeOH 0.1% | 7-OH-CBD | HQC | 3 | 32604 | 25363 | 38999 | 77.8 | 83.6 | 44256 | 97.9
NH4OH/MeOH 0.1% | 7-OH-CBD | HQC | 4 | 24336 | 19409 | 38246 | 79.8 | 63.6 | 45504 | 104.3
NH4OH/MeOH 0.1% | 7-OH-CBD | HQC | 5 | 34648 | 27378 | 38604 | 79 | 89.8 | 42680 | 96.1
NH4OH/MeOH 0.1% | 7-OH-CBD | HQC | 6 | 40110 | 31662 | 40313 | 78.9 | 99.5 | 42782 | 105.9
NH4OH/MeOH 0.1% | 7-COOH-CBD | LQC | 1 | 3961 | 3578 | 4188 | 90.3 | 94.6 | 52629 | 102.7
NH4OH/MeOH 0.1% | 7-COOH-CBD | LQC | 2 | 4116 | 3635 | 4264 | 88.3 | 96.5 | 52313 | 101.6
NH4OH/MeOH 0.1% | 7-COOH-CBD | LQC | 3 | 3634 | 3350 | 4371 | 92.2 | 83.1 | 52197 | 109.1
NH4OH/MeOH 0.1% | 7-COOH-CBD | LQC | 4 | 4051 | 3878 | 4289 | 95.7 | 94.5 | 52763 | 98.6
NH4OH/MeOH 0.1% | 7-COOH-CBD | LQC | 5 | 3658 | 3298 | 4081 | 90.2 | 89.6 | 51793 | 94
NH4OH/MeOH 0.1% | 7-COOH-CBD | LQC | 6 | 4141 | 3760 | 4268 | 90.8 | 97 | 51858 | 100.4
NH4OH/MeOH 0.1% | 7-COOH-CBD | HQC | 1 | 114961 | 106039 | 113544 | 92.2 | 101.2 | 53152 | 104.8
NH4OH/MeOH 0.1% | 7-COOH-CBD | HQC | 2 | 103410 | 90697 | 114214 | 87.7 | 90.5 | 54867 | 100.4
NH4OH/MeOH 0.1% | 7-COOH-CBD | HQC | 3 | 120786 | 110591 | 114785 | 91.6 | 105.2 | 50356 | 94.1
NH4OH/MeOH 0.1% | 7-COOH-CBD | HQC | 4 | 115313 | 108033 | 116528 | 93.7 | 99 | 53812 | 96.5
NH4OH/MeOH 0.1% | 7-COOH-CBD | HQC | 5 | 86073 | 78400 | 116557 | 91.1 | 73.8 | 47782 | 104.1
NH4OH/MeOH 0.1% | 7-COOH-CBD | HQC | 6 | 123164 | 108063 | 117221 | 87.7 | 105.1 | 53486 | 100.9
NH4OH/MeOH 0.1% | 6α-OH-CBD | LQC | 1 | 253 | 195 | 317 | 77.1 | 79.8 | 37719 | 91.5
NH4OH/MeOH 0.1% | 6α-OH-CBD | LQC | 2 | 216 | 166 | 305 | 76.9 | 70.8 | 39932 | 97.6
NH4OH/MeOH 0.1% | 6α-OH-CBD | LQC | 3 | 247 | 190 | 304 | 76.9 | 81.2 | 37510 | 102.2
NH4OH/MeOH 0.1% | 6α-OH-CBD | LQC | 4 | 267 | 206 | 326 | 77.2 | 81.9 | 36338 | 101.4
NH4OH/MeOH 0.1% | 6α-OH-CBD | LQC | 5 | 220 | 174 | 313 | 79.1 | 70.3 | 37817 | 101.4
NH4OH/MeOH 0.1% | 6α-OH-CBD | LQC | 6 | 211 | 156 | 304 | 73.9 | 69.4 | 38024 | 105.9
NH4OH/MeOH 0.1% | 6α-OH-CBD | HQC | 1 | 15478 | 12414 | 16666 | 80.2 | 92.9 | 38304 | 106.6
NH4OH/MeOH 0.1% | 6α-OH-CBD | HQC | 2 | 8873 | 6672 | 16567 | 75.2 | 53.6 | 39022 | 101.9
NH4OH/MeOH 0.1% | 6α-OH-CBD | HQC | 3 | 12462 | 9346 | 16941 | 75 | 73.6 | 37850 | 100.9
NH4OH/MeOH 0.1% | 6α-OH-CBD | HQC | 4 | 15600 | 12113 | 16968 | 77.6 | 91.9 | 39608 | 99.6
NH4OH/MeOH 0.1% | 6α-OH-CBD | HQC | 5 | 13862 | 10653 | 16856 | 76.9 | 82.2 | 37301 | 104.7
NH4OH/MeOH 0.1% | 6α-OH-CBD | HQC | 6 | 14027 | 10580 | 16437 | 75.4 | 85.3 | 35854 | 98.1
NH4OH/MeOH 0.1% | CBN | LQC | 1 | 343 | 270 | 323 | 83.7 | 106.2 | 41163 | 103.3
NH4OH/MeOH 0.1% | CBN | LQC | 2 | 289 | 266 | 336 | 79.3 | 86 | 39195 | 105
NH4OH/MeOH 0.1% | CBN | LQC | 3 | 218 | 266 | 342 | 77.8 | 63.7 | 40952 | 91.2
NH4OH/MeOH 0.1% | CBN | LQC | 4 | 275 | 267 | 340 | 78.6 | 80.9 | 39507 | 97.3
NH4OH/MeOH 0.1% | CBN | LQC | 5 | 257 | 259 | 342 | 75.8 | 75.1 | 39325 | 104.7
NH4OH/MeOH 0.1% | CBN | LQC | 6 | 243 | 268 | 337 | 79.5 | 72.1 | 39289 | 99.1
NH4OH/MeOH 0.1% | CBN | HQC | 1 | 14370 | 14021 | 17771 | 78.9 | 80.9 | 40655 | 101.3
NH4OH/MeOH 0.1% | CBN | HQC | 2 | 16501 | 14253 | 17662 | 80.7 | 93.4 | 41918 | 102.6
NH4OH/MeOH 0.1% | CBN | HQC | 3 | 11969 | 13599 | 17279 | 78.7 | 69.3 | 39871 | 106.4
NH4OH/MeOH 0.1% | CBN | HQC | 4 | 10524 | 13475 | 17569 | 76.7 | 59.9 | 42677 | 100.4
NH4OH/MeOH 0.1% | CBN | HQC | 5 | 16006 | 13156 | 17333 | 75.9 | 92.3 | 39798 | 104.8
NH4OH/MeOH 0.1% | CBN | HQC | 6 | 17532 | 14068 | 18683 | 75.3 | 93.8 | 38896 | 93.2
NH4OH/MeOH 0.1% | CBG | LQC | 1 | 161 | 228 | 294 | 77.7 | 54.8 | 34511 | 92.4
NH4OH/MeOH 0.1% | CBG | LQC | 2 | 242 | 215 | 290 | 74.3 | 83.4 | 35102 | 102.1
NH4OH/MeOH 0.1% | CBG | LQC | 3 | 222 | 214 | 280 | 76.4 | 79.3 | 35395 | 96.6
NH4OH/MeOH 0.1% | CBG | LQC | 4 | 143 | 230 | 316 | 72.7 | 45.3 | 34313 | 94.3
NH4OH/MeOH 0.1% | CBG | LQC | 5 | 263 | 222 | 297 | 74.8 | 88.6 | 34942 | 99.3
NH4OH/MeOH 0.1% | CBG | LQC | 6 | 240 | 203 | 271 | 74.9 | 88.6 | 34841 | 96.1
NH4OH/MeOH 0.1% | CBG | HQC | 1 | 14415 | 11634 | 15328 | 75.9 | 94 | 34980 | 95.6
NH4OH/MeOH 0.1% | CBG | HQC | 2 | 14997 | 11256 | 15831 | 71.1 | 94.7 | 33551 | 95.6
NH4OH/MeOH 0.1% | CBG | HQC | 3 | 14728 | 11749 | 16095 | 73 | 91.5 | 34326 | 98.9
NH4OH/MeOH 0.1% | CBG | HQC | 4 | 11457 | 12179 | 15339 | 79.4 | 74.7 | 36123 | 102.6
NH4OH/MeOH 0.1% | CBG | HQC | 5 | 12557 | 11942 | 15631 | 76.4 | 80.3 | 34167 | 105
NH4OH/MeOH 0.1% | CBG | HQC | 6 | 10261 | 11585 | 15263 | 75.9 | 67.2 | 35984 | 97.4
NH4OH/MeOH 0.1% | Δ9-THC | LQC | 1 | 247 | 292 | 378 | 77.2 | 65.3 | 44642 | 92.2
NH4OH/MeOH 0.1% | Δ9-THC | LQC | 2 | 393 | 298 | 389 | 76.5 | 101 | 43233 | 90.4
NH4OH/MeOH 0.1% | Δ9-THC | LQC | 3 | 349 | 310 | 387 | 80.2 | 90.2 | 46089 | 95.7
NH4OH/MeOH 0.1% | Δ9-THC | LQC | 4 | 264 | 309 | 374 | 82.5 | 70.6 | 44312 | 105.5
NH4OH/MeOH 0.1% | Δ9-THC | LQC | 5 | 298 | 285 | 374 | 76.3 | 79.7 | 43889 | 103.4
NH4OH/MeOH 0.1% | Δ9-THC | LQC | 6 | 233 | 277 | 370 | 74.8 | 63 | 43829 | 98.9
NH4OH/MeOH 0.1% | Δ9-THC | HQC | 1 | 15839 | 16181 | 20353 | 79.5 | 77.8 | 45354 | 98.1
NH4OH/MeOH 0.1% | Δ9-THC | HQC | 2 | 20250 | 16375 | 20623 | 79.4 | 98.2 | 44213 | 93.1
NH4OH/MeOH 0.1% | Δ9-THC | HQC | 3 | 15904 | 16291 | 20466 | 79.6 | 77.7 | 45119 | 102.5
NH4OH/MeOH 0.1% | Δ9-THC | HQC | 4 | 11766 | 15332 | 18999 | 80.7 | 61.9 | 44392 | 98.2
NH4OH/MeOH 0.1% | Δ9-THC | HQC | 5 | 19764 | 16460 | 20321 | 81 | 97.3 | 44314 | 100.9
NH4OH/MeOH 0.1% | Δ9-THC | HQC | 6 | 16627 | 15987 | 19616 | 81.5 | 84.8 | 43462 | 103.4
NH4OH/MeOH 0.1% | 11-OH-THC | LQC | 1 | 335 | 249 | 343 | 74.3 | 97.7 | 42474 | 99.9
NH4OH/MeOH 0.1% | 11-OH-THC | LQC | 2 | 274 | 203 | 352 | 74.1 | 77.8 | 42804 | 97.3
NH4OH/MeOH 0.1% | 11-OH-THC | LQC | 3 | 241 | 194 | 331 | 80.5 | 72.8 | 42933 | 99.1
NH4OH/MeOH 0.1% | 11-OH-THC | LQC | 4 | 292 | 228 | 358 | 78.1 | 81.6 | 42988 | 100.7
NH4OH/MeOH 0.1% | 11-OH-THC | LQC | 5 | 303 | 228 | 359 | 75.2 | 84.4 | 42901 | 100.8
NH4OH/MeOH 0.1% | 11-OH-THC | LQC | 6 | 241 | 183 | 340 | 75.9 | 70.9 | 41277 | 95.9
NH4OH/MeOH 0.1% | 11-OH-THC | HQC | 1 | 17530 | 13323 | 18112 | 76 | 96.8 | 42218 | 101.4
NH4OH/MeOH 0.1% | 11-OH-THC | HQC | 2 | 15466 | 11603 | 18976 | 75 | 81.5 | 39929 | 94.7
NH4OH/MeOH 0.1% | 11-OH-THC | HQC | 3 | 16409 | 12659 | 18343 | 77.1 | 89.5 | 41527 | 95.4
NH4OH/MeOH 0.1% | 11-OH-THC | HQC | 4 | 13682 | 10547 | 18274 | 77.1 | 74.9 | 43436 | 108
NH4OH/MeOH 0.1% | 11-OH-THC | HQC | 5 | 16345 | 11954 | 18474 | 73.1 | 88.5 | 43244 | 100.7
NH4OH/MeOH 0.1% | 11-OH-THC | HQC | 6 | 15181 | 11964 | 18699 | 78.8 | 81.2 | 42232 | 106
NH4OH/MeOH 0.1% | THC-COOH | LQC | 1 | 1496 | 1322 | 1671 | 88.4 | 89.5 | 46682 | 109.3
NH4OH/MeOH 0.1% | THC-COOH | LQC | 2 | 1548 | 1426 | 1661 | 92.1 | 93.2 | 50693 | 110.6
NH4OH/MeOH 0.1% | THC-COOH | LQC | 3 | 1494 | 1343 | 1720 | 89.9 | 86.9 | 49037 | 96.3
NH4OH/MeOH 0.1% | THC-COOH | LQC | 4 | 1451 | 1356 | 1689 | 93.5 | 85.9 | 50708 | 103.4
NH4OH/MeOH 0.1% | THC-COOH | LQC | 5 | 1553 | 1505 | 1671 | 96.9 | 92.9 | 51323 | 100.1
NH4OH/MeOH 0.1% | THC-COOH | LQC | 6 | 1611 | 1532 | 1681 | 95.1 | 95.8 | 49629 | 100.9
NH4OH/MeOH 0.1% | THC-COOH | HQC | 1 | 44716 | 39155 | 44284 | 87.6 | 101 | 52129 | 104
NH4OH/MeOH 0.1% | THC-COOH | HQC | 2 | 39662 | 37155 | 43981 | 93.7 | 90.2 | 51338 | 94.3
NH4OH/MeOH 0.1% | THC-COOH | HQC | 3 | 36377 | 32750 | 46877 | 90 | 77.6 | 50158 | 96.1
NH4OH/MeOH 0.1% | THC-COOH | HQC | 4 | 40729 | 37500 | 46758 | 92.1 | 87.1 | 48004 | 96.8
NH4OH/MeOH 0.1% | THC-COOH | HQC | 5 | 36879 | 34336 | 42122 | 93.1 | 87.6 | 48121 | 98.5
NH4OH/MeOH 0.1% | THC-COOH | HQC | 6 | 43231 | 40133 | 46399 | 92.8 | 93.2 | 50606 | 107.1
MeOH-H2O 80:20 | CBD | LQC | 1 | 1207 | 724 | 1557 | 46.5 | 77.5 | 47273 | 98.7
MeOH-H2O 80:20 | CBD | LQC | 2 | 1443 | 793 | 1602 | 49.5 | 90.1 | 46436 | 105.7
MeOH-H2O 80:20 | CBD | LQC | 3 | 1454 | 720 | 1615 | 44.6 | 90 | 49719 | 98.4
MeOH-H2O 80:20 | CBD | LQC | 4 | 1302 | 746 | 1551 | 48.1 | 83.9 | 50048 | 102.2
MeOH-H2O 80:20 | CBD | LQC | 5 | 1012 | 760 | 1625 | 46.8 | 62.3 | 48602 | 102.2
MeOH-H2O 80:20 | CBD | LQC | 6 | 1473 | 730 | 1576 | 46.3 | 93.5 | 48360 | 96.9
MeOH-H2O 80:20 | CBD | HQC | 1 | 39420 | 20741 | 42943 | 48.3 | 91.8 | 47538 | 104.7
MeOH-H2O 80:20 | CBD | HQC | 2 | 39210 | 20541 | 43062 | 47.7 | 91.1 | 47111 | 101.7
MeOH-H2O 80:20 | CBD | HQC | 3 | 34477 | 20456 | 42974 | 47.6 | 80.2 | 47991 | 105.5
MeOH-H2O 80:20 | CBD | HQC | 4 | 41725 | 20438 | 42847 | 47.7 | 97.4 | 47896 | 96.5
MeOH-H2O 80:20 | CBD | HQC | 5 | 38067 | 20097 | 42759 | 47 | 89 | 48010 | 96
MeOH-H2O 80:20 | CBD | HQC | 6 | 40027 | 20871 | 41826 | 49.9 | 95.7 | 49579 | 101.2
MeOH-H2O 80:20 | 7-OH-CBD | LQC | 1 | 1317 | 649 | 1442 | 45 | 91.3 | 44902 | 92.5
MeOH-H2O 80:20 | 7-OH-CBD | LQC | 2 | 1065 | 663 | 1496 | 44.3 | 71.2 | 44539 | 101
MeOH-H2O 80:20 | 7-OH-CBD | LQC | 3 | 1454 | 638 | 1450 | 44 | 100.3 | 43292 | 98.2
MeOH-H2O 80:20 | 7-OH-CBD | LQC | 4 | 1179 | 612 | 1408 | 43.5 | 83.7 | 44662 | 99
MeOH-H2O 80:20 | 7-OH-CBD | LQC | 5 | 1259 | 624 | 1457 | 42.8 | 86.4 | 42727 | 96.1
MeOH-H2O 80:20 | 7-OH-CBD | LQC | 6 | 947 | 633 | 1422 | 44.5 | 66.6 | 42780 | 99.6
MeOH-H2O 80:20 | 7-OH-CBD | HQC | 1 | 30341 | 16396 | 39037 | 42 | 77.7 | 43889 | 98.4
MeOH-H2O 80:20 | 7-OH-CBD | HQC | 2 | 41480 | 16898 | 40137 | 42.1 | 103.3 | 44233 | 99.2
MeOH-H2O 80:20 | 7-OH-CBD | HQC | 3 | 38781 | 16344 | 37746 | 43.3 | 102.7 | 45071 | 94.2
MeOH-H2O 80:20 | 7-OH-CBD | HQC | 4 | 35442 | 16333 | 37548 | 43.5 | 94.4 | 44358 | 100.6
MeOH-H2O 80:20 | 7-OH-CBD | HQC | 5 | 36891 | 17490 | 39391 | 44.4 | 93.7 | 43697 | 104.7
MeOH-H2O 80:20 | 7-OH-CBD | HQC | 6 | 34198 | 15660 | 38666 | 40.5 | 88.4 | 43559 | 104.9
MeOH-H2O 80:20 | 7-COOH-CBD | LQC | 1 | 3531 | 2000 | 4504 | 44.4 | 78.4 | 51794 | 103.8
MeOH-H2O 80:20 | 7-COOH-CBD | LQC | 2 | 3986 | 1832 | 4290 | 42.7 | 92.9 | 53075 | 97.7
MeOH-H2O 80:20 | 7-COOH-CBD | LQC | 3 | 3831 | 1791 | 4389 | 40.8 | 87.3 | 50570 | 97.7
MeOH-H2O 80:20 | 7-COOH-CBD | LQC | 4 | 4096 | 1771 | 4395 | 40.3 | 93.2 | 52373 | 100.1
MeOH-H2O 80:20 | 7-COOH-CBD | LQC | 5 | 4019 | 1892 | 4389 | 43.1 | 91.6 | 53355 | 105.1
MeOH-H2O 80:20 | 7-COOH-CBD | LQC | 6 | 4131 | 1898 | 4393 | 43.2 | 94 | 55687 | 101.9
MeOH-H2O 80:20 | 7-COOH-CBD | HQC | 1 | 128446 | 54006 | 120819 | 44.7 | 106.3 | 49911 | 104
MeOH-H2O 80:20 | 7-COOH-CBD | HQC | 2 | 100596 | 47331 | 109563 | 43.2 | 91.8 | 53340 | 99
MeOH-H2O 80:20 | 7-COOH-CBD | HQC | 3 | 112895 | 48685 | 114283 | 42.6 | 98.8 | 51230 | 108.4
MeOH-H2O 80:20 | 7-COOH-CBD | HQC | 4 | 98055 | 47936 | 113324 | 42.3 | 86.5 | 52530 | 98.8
MeOH-H2O 80:20 | 7-COOH-CBD | HQC | 5 | 93031 | 48685 | 116472 | 41.8 | 79.9 | 56600 | 108.1
MeOH-H2O 80:20 | 7-COOH-CBD | HQC | 6 | 99997 | 45872 | 108701 | 42.2 | 92 | 51602 | 98.7
MeOH-H2O 80:20 | 6α-OH-CBD | LQC | 1 | 265 | 166 | 323 | 51.3 | 82 | 38199 | 107.8
MeOH-H2O 80:20 | 6α-OH-CBD | LQC | 2 | 284 | 156 | 308 | 50.8 | 92.2 | 38691 | 109.9
MeOH-H2O 80:20 | 6α-OH-CBD | LQC | 3 | 297 | 170 | 306 | 55.7 | 97.1 | 37524 | 96.3
MeOH-H2O 80:20 | 6α-OH-CBD | LQC | 4 | 360 | 167 | 319 | 52.5 | 112.9 | 36565 | 100.3
MeOH-H2O 80:20 | 6α-OH-CBD | LQC | 5 | 204 | 173 | 331 | 52.3 | 61.6 | 39164 | 111.6
MeOH-H2O 80:20 | 6α-OH-CBD | LQC | 6 | 283 | 171 | 323 | 52.8 | 87.6 | 38609 | 92.1
MeOH-H2O 80:20 | 6α-OH-CBD | HQC | 1 | 13364 | 8307 | 16450 | 50.5 | 81.2 | 36945 | 103.5
MeOH-H2O 80:20 | 6α-OH-CBD | HQC | 2 | 14466 | 8879 | 16849 | 52.7 | 85.9 | 37732 | 104.7
MeOH-H2O 80:20 | 6α-OH-CBD | HQC | 3 | 18498 | 8973 | 17027 | 52.7 | 108.6 | 39669 | 102.1
MeOH-H2O 80:20 | 6α-OH-CBD | HQC | 4 | 11328 | 8577 | 16784 | 51.1 | 67.5 | 39394 | 105.3
MeOH-H2O 80:20 | 6α-OH-CBD | HQC | 5 | 14173 | 9493 | 17228 | 55.1 | 82.3 | 39410 | 98.8
MeOH-H2O 80:20 | 6α-OH-CBD | HQC | 6 | 10400 | 8960 | 17131 | 52.3 | 60.7 | 38240 | 105.2
MeOH-H2O 80:20 | CBN | LQC | 1 | 207 | 167 | 341 | 49 | 60.7 | 39250 | 101.3
MeOH-H2O 80:20 | CBN | LQC | 2 | 311 | 171 | 339 | 50.5 | 91.7 | 38634 | 99.4
MeOH-H2O 80:20 | CBN | LQC | 3 | 328 | 171 | 334 | 51.2 | 98.2 | 40548 | 102.2
MeOH-H2O 80:20 | CBN | LQC | 4 | 267 | 164 | 329 | 49.8 | 81.2 | 38966 | 102.1
MeOH-H2O 80:20 | CBN | LQC | 5 | 338 | 158 | 335 | 47.3 | 100.9 | 40236 | 96
MeOH-H2O 80:20 | CBN | LQC | 6 | 243 | 174 | 328 | 52.9 | 74.1 | 40477 | 106.1
MeOH-H2O 80:20 | CBN | HQC | 1 | 17412 | 8668 | 17406 | 49.8 | 100 | 39555 | 95.7
MeOH-H2O 80:20 | CBN | HQC | 2 | 13300 | 9358 | 18207 | 51.4 | 73 | 41236 | 95.6
MeOH-H2O 80:20 | CBN | HQC | 3 | 13064 | 9243 | 18017 | 51.3 | 72.5 | 38589 | 95.6
MeOH-H2O 80:20 | CBN | HQC | 4 | 14097 | 9451 | 18070 | 52.3 | 78 | 38520 | 99.8
MeOH-H2O 80:20 | CBN | HQC | 5 | 11905 | 9212 | 17853 | 51.6 | 66.7 | 38789 | 93.1
MeOH-H2O 80:20 | CBN | HQC | 6 | 16340 | 9209 | 17951 | 51.3 | 91 | 39637 | 96.9
MeOH-H2O 80:20 | CBG | LQC | 1 | 229 | 141 | 292 | 48.3 | 78.4 | 35794 | 105
MeOH-H2O 80:20 | CBG | LQC | 2 | 281 | 152 | 302 | 50.3 | 93 | 36502 | 107.5
MeOH-H2O 80:20 | CBG | LQC | 3 | 292 | 144 | 294 | 49.1 | 99.3 | 37112 | 102.3
MeOH-H2O 80:20 | CBG | LQC | 4 | 257 | 143 | 282 | 50.6 | 91.1 | 36641 | 93.3
MeOH-H2O 80:20 | CBG | LQC | 5 | 257 | 157 | 307 | 51.2 | 83.7 | 35656 | 104.5
MeOH-H2O 80:20 | CBG | LQC | 6 | 293 | 159 | 302 | 52.6 | 97 | 36555 | 105.9
MeOH-H2O 80:20 | CBG | HQC | 1 | 11181 | 8499 | 16407 | 51.8 | 68.1 | 36292 | 103.3
MeOH-H2O 80:20 | CBG | HQC | 2 | 14142 | 7811 | 15812 | 49.4 | 89.4 | 35359 | 102
MeOH-H2O 80:20 | CBG | HQC | 3 | 10151 | 7733 | 15781 | 49 | 64.3 | 32343 | 98.4
MeOH-H2O 80:20 | CBG | HQC | 4 | 12664 | 7603 | 15454 | 49.2 | 81.9 | 36651 | 99
MeOH-H2O 80:20 | CBG | HQC | 5 | 11056 | 8469 | 15980 | 53 | 69.2 | 33812 | 100
MeOH-H2O 80:20 | CBG | HQC | 6 | 11560 | 7979 | 15554 | 51.3 | 74.3 | 35861 | 98.6
MeOH-H2O 80:20 | Δ9-THC | LQC | 1 | 299 | 165 | 375 | 44.1 | 79.7 | 44761 | 96.8
MeOH-H2O 80:20 | Δ9-THC | LQC | 2 | 253 | 163 | 368 | 44.2 | 68.8 | 45647 | 104.7
MeOH-H2O 80:20 | Δ9-THC | LQC | 3 | 324 | 167 | 383 | 43.6 | 84.6 | 43960 | 97.4
MeOH-H2O 80:20 | Δ9-THC | LQC | 4 | 323 | 158 | 361 | 43.9 | 89.5 | 43748 | 102.5
MeOH-H2O 80:20 | Δ9-THC | LQC | 5 | 261 | 169 | 377 | 44.7 | 69.2 | 45248 | 107
MeOH-H2O 80:20 | Δ9-THC | LQC | 6 | 315 | 170 | 371 | 45.7 | 84.9 | 43875 | 104.5
MeOH-H2O 80:20 | Δ9-THC | HQC | 1 | 17208 | 8345 | 20354 | 41 | 84.5 | 47473 | 101
MeOH-H2O 80:20 | Δ9-THC | HQC | 2 | 17873 | 9036 | 21015 | 43 | 85 | 46140 | 104.2
MeOH-H2O 80:20 | Δ9-THC | HQC | 3 | 15148 | 9194 | 20943 | 43.9 | 72.3 | 46284 | 104.2
MeOH-H2O 80:20 | Δ9-THC | HQC | 4 | 22152 | 8974 | 20076 | 44.7 | 110.3 | 44595 | 100.8
MeOH-H2O 80:20 | Δ9-THC | HQC | 5 | 13804 | 9195 | 19859 | 46.3 | 69.5 | 46993 | 95.9
MeOH-H2O 80:20 | Δ9-THC | HQC | 6 | 14658 | 9039 | 20267 | 44.6 | 72.3 | 44272 | 96
MeOH-H2O 80:20 | 11-OH-THC | LQC | 1 | 314 | 170 | 349 | 48.6 | 90 | 41095 | 95.9
MeOH-H2O 80:20 | 11-OH-THC | LQC | 2 | 274 | 174 | 345 | 50.4 | 79.4 | 42858 | 102.2
MeOH-H2O 80:20 | 11-OH-THC | LQC | 3 | 281 | 176 | 353 | 49.8 | 79.6 | 41196 | 108.1
MeOH-H2O 80:20 | 11-OH-THC | LQC | 4 | 213 | 171 | 335 | 50.9 | 63.6 | 41158 | 101.4
MeOH-H2O 80:20 | 11-OH-THC | LQC | 5 | 356 | 176 | 350 | 50.2 | 101.7 | 38538 | 100.1
MeOH-H2O 80:20 | 11-OH-THC | LQC | 6 | 343 | 192 | 388 | 49.5 | 88.4 | 39315 | 101.6
MeOH-H2O 80:20 | 11-OH-THC | HQC | 1 | 15524 | 9095 | 18599 | 48.9 | 83.5 | 41637 | 101
MeOH-H2O 80:20 | 11-OH-THC | HQC | 2 | 15914 | 10003 | 19461 | 51.4 | 81.8 | 42945 | 102.2
MeOH-H2O 80:20 | 11-OH-THC | HQC | 3 | 20146 | 9449 | 18785 | 50.3 | 107.2 | 41078 | 100.2
MeOH-H2O 80:20 | 11-OH-THC | HQC | 4 | 16085 | 10097 | 18837 | 53.6 | 85.4 | 43632 | 95.7
MeOH-H2O 80:20 | 11-OH-THC | HQC | 5 | 18554 | 10140 | 19278 | 52.6 | 96.2 | 41853 | 96.1
MeOH-H2O 80:20 | 11-OH-THC | HQC | 6 | 14716 | 9830 | 18759 | 52.4 | 78.4 | 43203 | 98.3
MeOH-H2O 80:20 | THC-COOH | LQC | 1 | 1510 | 729 | 1703 | 42.8 | 88.7 | 51494 | 96.7
MeOH-H2O 80:20 | THC-COOH | LQC | 2 | 1784 | 735 | 1722 | 42.7 | 103.6 | 45513 | 104.6
MeOH-H2O 80:20 | THC-COOH | LQC | 3 | 1535 | 757 | 1737 | 43.6 | 88.4 | 48576 | 102.4
MeOH-H2O 80:20 | THC-COOH | LQC | 4 | 1737 | 722 | 1678 | 43 | 103.5 | 49087 | 101.6
MeOH-H2O 80:20 | THC-COOH | LQC | 5 | 1557 | 711 | 1631 | 43.6 | 95.5 | 47269 | 106.9
MeOH-H2O 80:20 | THC-COOH | LQC | 6 | 1656 | 720 | 1698 | 42.4 | 97.5 | 50158 | 106.5
MeOH-H2O 80:20 | THC-COOH | HQC | 1 | 37891 | 19269 | 44917 | 42.9 | 84.4 | 51098 | 106.1
MeOH-H2O 80:20 | THC-COOH | HQC | 2 | 40892 | 20033 | 44816 | 44.7 | 91.2 | 53617 | 102.7
MeOH-H2O 80:20 | THC-COOH | HQC | 3 | 38404 | 18536 | 42907 | 43.2 | 89.5 | 50386 | 95.3
MeOH-H2O 80:20 | THC-COOH | HQC | 4 | 42724 | 18538 | 43313 | 42.8 | 98.6 | 52831 | 105.9
MeOH-H2O 80:20 | THC-COOH | HQC | 5 | 36468 | 19890 | 44398 | 44.8 | 82.1 | 49859 | 108.2
MeOH-H2O 80:20 | THC-COOH | HQC | 6 | 45338 | 17680 | 44534 | 39.7 | 101.8 | 45751 | 97
ACN | CBD | LQC | 1 | 1372 | 803 | 1563 | 51.4 | 87.8 | 45307 | 98.3
ACN | CBD | LQC | 2 | 1310 | 820 | 1601 | 51.2 | 81.8 | 47052 | 104.4
ACN | CBD | LQC | 3 | 1312 | 794 | 1560 | 50.9 | 84.1 | 46613 | 97.6
ACN | CBD | LQC | 4 | 1386 | 824 | 1631 | 50.5 | 85 | 45048 | 104.5
ACN | CBD | LQC | 5 | 1095 | 809 | 1562 | 51.8 | 70.1 | 47584 | 92.8
ACN | CBD | LQC | 6 | 1217 | 864 | 1563 | 55.3 | 77.9 | 47430 | 102
ACN | CBD | HQC | 1 | 38529 | 22331 | 40899 | 54.6 | 94.2 | 50527 | 99.4
ACN | CBD | HQC | 2 | 39683 | 21926 | 42824 | 51.2 | 92.7 | 50349 | 92.7
ACN | CBD | HQC | 3 | 28035 | 22409 | 42043 | 53.3 | 66.7 | 49728 | 106.7
ACN | CBD | HQC | 4 | 24340 | 22769 | 43042 | 52.9 | 56.5 | 49755 | 93.7
ACN | CBD | HQC | 5 | 38164 | 22567 | 42103 | 53.6 | 90.6 | 46756 | 100.7
ACN | CBD | HQC | 6 | 29723 | 21822 | 42621 | 51.2 | 69.7 | 49232 | 101.7
ACN | 7-OH-CBD | LQC | 1 | 1161 | 655 | 1445 | 45.3 | 80.3 | 44594 | 97.8
ACN | 7-OH-CBD | LQC | 2 | 1449 | 703 | 1516 | 46.4 | 95.6 | 44331 | 101.7
ACN | 7-OH-CBD | LQC | 3 | 1326 | 617 | 1386 | 44.5 | 95.7 | 44648 | 99.3
ACN | 7-OH-CBD | LQC | 4 | 1012 | 692 | 1510 | 45.8 | 67 | 43456 | 99.8
ACN | 7-OH-CBD | LQC | 5 | 1175 | 681 | 1487 | 45.8 | 79 | 45256 | 95.6
ACN | 7-OH-CBD | LQC | 6 | 1111 | 615 | 1430 | 43 | 77.7 | 43204 | 99.9
ACN | 7-OH-CBD | HQC | 1 | 29790 | 17438 | 37827 | 46.1 | 78.8 | 45111 | 108.3
ACN | 7-OH-CBD | HQC | 2 | 30724 | 17243 | 39278 | 43.9 | 78.2 | 44144 | 103
ACN | 7-OH-CBD | HQC | 3 | 33456 | 16698 | 38742 | 43.1 | 86.4 | 42476 | 95.9
ACN | 7-OH-CBD | HQC | 4 | 27642 | 17757 | 38105 | 46.6 | 72.5 | 44802 | 93.9
ACN | 7-OH-CBD | HQC | 5 | 34121 | 16692 | 37764 | 44.2 | 90.4 | 45287 | 98.7
ACN | 7-OH-CBD | HQC | 6 | 38225 | 19047 | 40612 | 46.9 | 94.1 | 45756 | 93.9
ACN | 7-COOH-CBD | LQC | 1 | 3000 | 2020 | 4392 | 46 | 68.3 | 55559 | 102.9
ACN | 7-COOH-CBD | LQC | 2 | 3975 | 2186 | 4453 | 49.1 | 89.3 | 52682 | 99.3
ACN | 7-COOH-CBD | LQC | 3 | 4043 | 1946 | 4443 | 43.8 | 91 | 50794 | 99.8
ACN | 7-COOH-CBD | LQC | 4 | 3068 | 2061 | 4259 | 48.4 | 72 | 54376 | 97
ACN | 7-COOH-CBD | LQC | 5 | 4069 | 1986 | 4346 | 45.7 | 93.6 | 50361 | 98.2
ACN | 7-COOH-CBD | LQC | 6 | 3404 | 1856 | 4097 | 45.3 | 83.1 | 53949 | 102.5
ACN | 7-COOH-CBD | HQC | 1 | 90639 | 54109 | 115617 | 46.8 | 78.4 | 50002 | 97.6
ACN | 7-COOH-CBD | HQC | 2 | 103259 | 54782 | 117307 | 46.7 | 88 | 56474 | 92.1
ACN | 7-COOH-CBD | HQC | 3 | 109533 | 53902 | 116923 | 46.1 | 93.7 | 52357 | 92.7
ACN | 7-COOH-CBD | HQC | 4 | 118884 | 54840 | 119218 | 46 | 99.7 | 52231 | 99.5
ACN | 7-COOH-CBD | HQC | 5 | 119952 | 53240 | 112559 | 47.3 | 106.6 | 50588 | 101.5
ACN | 7-COOH-CBD | HQC | 6 | 111898 | 54011 | 117161 | 46.1 | 95.5 | 53008 | 102
ACN | 6α-OH-CBD | LQC | 1 | 277 | 142 | 309 | 46 | 89.6 | 39334 | 100.2
ACN | 6α-OH-CBD | LQC | 2 | 243 | 153 | 314 | 48.8 | 77.4 | 38999 | 91.2
ACN | 6α-OH-CBD | LQC | 3 | 260 | 141 | 306 | 46 | 85 | 39617 | 89.1
ACN | 6α-OH-CBD | LQC | 4 | 269 | 151 | 308 | 49 | 87.3 | 36922 | 93.4
ACN | 6α-OH-CBD | LQC | 5 | 288 | 154 | 330 | 46.6 | 87.3 | 39254 | 101.6
ACN | 6α-OH-CBD | LQC | 6 | 249 | 145 | 319 | 45.6 | 78.1 | 38512 | 109.6
ACN | 6α-OH-CBD | HQC | 1 | 11619 | 7800 | 16217 | 48.1 | 71.6 | 39278 | 106.3
ACN | 6α-OH-CBD | HQC | 2 | 12080 | 8198 | 17222 | 47.6 | 70.1 | 38541 | 101.8
ACN | 6α-OH-CBD | HQC | 3 | 10673 | 7381 | 16294 | 45.3 | 65.5 | 39812 | 97.6
ACN | 6α-OH-CBD | HQC | 4 | 11890 | 8138 | 17389 | 46.8 | 68.4 | 40047 | 102.9
ACN | 6α-OH-CBD | HQC | 5 | 13404 | 7643 | 16651 | 45.9 | 80.5 | 37675 | 102.7
ACN | 6α-OH-CBD | HQC | 6 | 13507 | 7139 | 15972 | 44.7 | 84.6 | 38087 | 98.7
ACN | CBN | LQC | 1 | 227 | 142 | 328 | 43.4 | 69.2 | 38616 | 100.3
ACN | CBN | LQC | 2 | 274 | 139 | 323 | 42.9 | 84.8 | 40213 | 100.8
ACN | CBN | LQC | 3 | 238 | 136 | 338 | 40.1 | 70.4 | 37980 | 97.7
ACN | CBN | LQC | 4 | 313 | 141 | 329 | 42.8 | 95.1 | 41428 | 97.6
ACN | CBN | LQC | 5 | 244 | 140 | 333 | 41.9 | 73.3 | 41033 | 105.4
ACN | CBN | LQC | 6 | 370 | 134 | 330 | 40.5 | 112.1 | 38336 | 97.2
ACN | CBN | HQC | 1 | 9544 | 7771 | 18114 | 42.9 | 52.7 | 40234 | 98.4
ACN | CBN | HQC | 2 | 18568 | 7726 | 17440 | 44.3 | 106.5 | 42456 | 99.2
ACN | CBN | HQC | 3 | 13470 | 7209 | 17540 | 41.1 | 76.8 | 41008 | 104.7
ACN | CBN | HQC | 4 | 16076 | 7739 | 18427 | 42 | 87.2 | 40501 | 91.7
ACN | CBN | HQC | 5 | 15749 | 8188 | 18823 | 43.5 | 83.7 | 40559 | 102.4
ACN | CBN | HQC | 6 | 14978 | 7471 | 17537 | 42.6 | 85.4 | 38468 | 95.3
ACN | CBG | LQC | 1 | 192 | 135 | 290 | 46.5 | 66.2 | 35730 | 96.9
ACN | CBG | LQC | 2 | 104 | 155 | 308 | 50.4 | 33.8 | 33682 | 101.3
ACN | CBG | LQC | 3 | 328 | 150 | 294 | 50.9 | 111.6 | 35019 | 91.8
ACN | CBG | LQC | 4 | 210 | 152 | 291 | 52.2 | 72.2 | 33596 | 98.4
ACN | CBG | LQC | 5 | 170 | 152 | 300 | 50.6 | 56.7 | 35054 | 98.9
ACN | CBG | LQC | 6 | 211 | 147 | 292 | 50.4 | 72.3 | 36455 | 107.5
ACN | CBG | HQC | 1 | 8386 | 7503 | 15503 | 48.4 | 54.1 | 36606 | 104.9
ACN | CBG | HQC | 2 | 14388 | 7697 | 15612 | 49.3 | 92.2 | 34437 | 96
ACN | CBG | HQC | 3 | 12069 | 7902 | 15434 | 51.2 | 78.2 | 36239 | 98
ACN | CBG | HQC | 4 | 15639 | 7445 | 15288 | 48.7 | 102.3 | 34935 | 103.2
ACN | CBG | HQC | 5 | 13305 | 8082 | 16131 | 50.1 | 82.5 | 35456 | 103.9
ACN | CBG | HQC | 6 | 10638 | 7510 | 15172 | 49.5 | 70.1 | 32382 | 92.6
ACN | Δ9-THC | LQC | 1 | 376 | 193 | 376 | 51.3 | 100 | 42527 | 92.7
ACN | Δ9-THC | LQC | 2 | 303 | 193 | 376 | 51.4 | 80.6 | 47152 | 104.6
ACN | Δ9-THC | LQC | 3 | 333 | 185 | 384 | 48.1 | 86.7 | 46187 | 92.9
ACN | Δ9-THC | LQC | 4 | 327 | 209 | 401 | 52.2 | 81.5 | 44951 | 103.3
ACN | Δ9-THC | LQC | 5 | 245 | 195 | 380 | 51.2 | 64.5 | 45864 | 100.7
ACN | Δ9-THC | LQC | 6 | 325 | 194 | 388 | 50.1 | 83.8 | 46612 | 102.7
ACN | Δ9-THC | HQC | 1 | 16307 | 10285 | 19516 | 52.7 | 83.6 | 44598 | 102.9
ACN | Δ9-THC | HQC | 2 | 22878 | 11330 | 20827 | 54.4 | 109.8 | 44138 | 100.8
ACN | Δ9-THC | HQC | 3 | 14432 | 11014 | 20781 | 53 | 69.4 | 45096 | 98.2
ACN | Δ9-THC | HQC | 4 | 19437 | 10637 | 20734 | 51.3 | 93.7 | 45815 | 111.9
ACN | Δ9-THC | HQC | 5 | 15329 | 11039 | 21189 | 52.1 | 72.3 | 45453 | 101.8
ACN | Δ9-THC | HQC | 6 | 12960 | 10432 | 19947 | 52.3 | 65 | 43932 | 94.6
ACN | 11-OH-THC | LQC | 1 | 315 | 158 | 345 | 45.8 | 91.3 | 41774 | 99.1
ACN | 11-OH-THC | LQC | 2 | 237 | 169 | 354 | 47.8 | 66.9 | 43127 | 94.9
ACN | 11-OH-THC | LQC | 3 | 277 | 173 | 359 | 48.1 | 77.2 | 40948 | 107.3
ACN | 11-OH-THC | LQC | 4 | 289 | 166 | 357 | 46.5 | 81 | 42937 | 100.5
ACN | 11-OH-THC | LQC | 5 | 285 | 156 | 346 | 45.2 | 82.4 | 41373 | 102.5
ACN | 11-OH-THC | LQC | 6 | 320 | 166 | 358 | 46.4 | 89.4 | 42268 | 103.7
ACN | 11-OH-THC | HQC | 1 | 12574 | 8562 | 18943 | 45.2 | 66.4 | 42188 | 101.1
ACN | 11-OH-THC | HQC | 2 | 17915 | 8351 | 18274 | 45.7 | 98 | 41082 | 100.2
ACN | 11-OH-THC | HQC | 3 | 17397 | 8284 | 18328 | 45.2 | 94.9 | 41820 | 98
ACN | 11-OH-THC | HQC | 4 | 17782 | 8209 | 18828 | 43.6 | 94.4 | 42091 | 100.7
ACN | 11-OH-THC | HQC | 5 | 15316 | 9619 | 20040 | 48 | 76.4 | 42225 | 101.8
ACN | 11-OH-THC | HQC | 6 | 19088 | 9199 | 19366 | 47.5 | 98.6 | 41808 | 99.1
ACN | THC-COOH | LQC | 1 | 1568 | 818 | 1734 | 47.2 | 90.4 | 49142 | 98.2
ACN | THC-COOH | LQC | 2 | 1502 | 878 | 1746 | 50.3 | 86 | 48744 | 98.7
ACN | THC-COOH | LQC | 3 | 1235 | 889 | 1732 | 51.3 | 71.3 | 47846 | 95.2
ACN | THC-COOH | LQC | 4 | 1674 | 917 | 1761 | 52.1 | 95.1 | 53461 | 99.8
ACN | THC-COOH | LQC | 5 | 1545 | 826 | 1638 | 50.4 | 94.3 | 49066 | 100.3
ACN | THC-COOH | LQC | 6 | 1611 | 837 | 1684 | 49.7 | 95.7 | 48757 | 93.5
ACN | THC-COOH | HQC | 1 | 42951 | 23708 | 45418 | 52.2 | 94.6 | 49444 | 98.1
ACN | THC-COOH | HQC | 2 | 43745 | 22435 | 43395 | 51.7 | 100.8 | 49727 | 93
ACN | THC-COOH | HQC | 3 | 34315 | 22089 | 45172 | 48.9 | 76 | 51879 | 100.7
ACN | THC-COOH | HQC | 4 | 38272 | 23773 | 45983 | 51.7 | 83.2 | 49966 | 101
ACN | THC-COOH | HQC | 5 | 38901 | 24108 | 46184 | 52.2 | 84.2 | 48536 | 101.9
ACN | THC-COOH | HQC | 6 | 31043 | 22185 | 43672 | 50.8 | 71.1 | 48860 | 93.4
ACN-H2O 80:20 | CBD | LQC | 1 | 1061 | 719 | 1602 | 44.9 | 66.2 | 47783 | 102.2
ACN-H2O 80:20 | CBD | LQC | 2 | 1435 | 663 | 1567 | 42.3 | 91.6 | 46725 | 96.8
ACN-H2O 80:20 | CBD | LQC | 3 | 1301 | 675 | 1544 | 43.7 | 84.3 | 48011 | 99.5
ACN-H2O 80:20 | CBD | LQC | 4 | 1137 | 682 | 1604 | 42.5 | 70.9 | 45695 | 95.9
ACN-H2O 80:20 | CBD | LQC | 5 | 1380 | 642 | 1550 | 41.4 | 89 | 48943 | 105.6
ACN-H2O 80:20 | CBD | LQC | 6 | 1183 | 644 | 1537 | 41.9 | 77 | 45417 | 99.2
ACN-H2O 80:20 | CBD | HQC | 1 | 41055 | 16891 | 40997 | 41.2 | 100.1 | 47935 | 102.6
ACN-H2O 80:20 | CBD | HQC | 2 | 38874 | 16784 | 40737 | 41.2 | 95.4 | 48692 | 101
ACN-H2O 80:20 | CBD | HQC | 3 | 32325 | 17676 | 41591 | 42.5 | 77.7 | 49782 | 104.2
ACN-H2O 80:20 | CBD | HQC | 4 | 40821 | 17076 | 41751 | 40.9 | 97.8 | 46024 | 102.3
ACN-H2O 80:20 | CBD | HQC | 5 | 28336 | 16951 | 41854 | 40.5 | 67.7 | 47244 | 97.7
ACN-H2O 80:20 | CBD | HQC | 6 | 31836 | 17308 | 41406 | 41.8 | 76.9 | 48851 | 98.4
ACN-H2O 80:20 | 7-OH-CBD | LQC | 1 | 1561 | 844 | 1461 | 57.8 | 106.8 | 43509 | 96.9
ACN-H2O 80:20 | 7-OH-CBD | LQC | 2 | 1211 | 839 | 1499 | 56 | 80.8 | 44208 | 101.2
ACN-H2O 80:20 | 7-OH-CBD | LQC | 3 | 1215 | 753 | 1389 | 54.2 | 87.5 | 41976 | 98.5
ACN-H2O 80:20 | 7-OH-CBD | LQC | 4 | 1337 | 871 | 1487 | 58.6 | 89.9 | 44707 | 93.9
ACN-H2O 80:20 | 7-OH-CBD | LQC | 5 | 1495 | 833 | 1507 | 55.3 | 99.2 | 42944 | 95.4
ACN-H2O 80:20 | 7-OH-CBD | LQC | 6 | 1540 | 781 | 1422 | 54.9 | 108.3 | 44501 | 98.7
ACN-H2O 80:20 | 7-OH-CBD | HQC | 1 | 29860 | 21428 | 37462 | 57.2 | 79.7 | 45749 | 94.1
ACN-H2O 80:20 | 7-OH-CBD | HQC | 2 | 34607 | 21574 | 39368 | 54.8 | 87.9 | 42884 | 95.8
ACN-H2O 80:20 | 7-OH-CBD | HQC | 3 | 26845 | 22338 | 38581 | 57.9 | 69.6 | 46449 | 101.7
ACN-H2O 80:20 | 7-OH-CBD | HQC | 4 | 38714 | 23949 | 40319 | 59.4 | 96 | 44610 | 96
ACN-H2O 80:20 | 7-OH-CBD | HQC | 5 | 27675 | 21673 | 38496 | 56.3 | 71.9 | 44079 | 97.5
ACN-H2O 80:20 | 7-OH-CBD | HQC | 6 | 33249 | 23198 | 39723 | 58.4 | 83.7 | 45623 | 105.8
ACN-H2O 80:20 | 7-COOH-CBD | LQC | 1 | 3795 | 2621 | 4551 | 57.6 | 83.4 | 51563 | 94.7
ACN-H2O 80:20 | 7-COOH-CBD | LQC | 2 | 4315 | 2458 | 4335 | 56.7 | 99.5 | 51046 | 100.4
ACN-H2O 80:20 | 7-COOH-CBD | LQC | 3 | 3927 | 2488 | 4349 | 57.2 | 90.3 | 52076 | 104.2
ACN-H2O 80:20 | 7-COOH-CBD | LQC | 4 | 4081 | 2483 | 4325 | 57.4 | 94.4 | 54013 | 102.2
ACN-H2O 80:20 | 7-COOH-CBD | LQC | 5 | 3705 | 2583 | 4462 | 57.9 | 83 | 52789 | 102.6
ACN-H2O 80:20 | 7-COOH-CBD | LQC | 6 | 4228 | 2362 | 4256 | 55.5 | 99.3 | 50247 | 107.4
ACN-H2O 80:20 | 7-COOH-CBD | HQC | 1 | 106531 | 63413 | 109144 | 58.1 | 97.6 | 49364 | 104.8
ACN-H2O 80:20 | 7-COOH-CBD | HQC | 2 | 94537 | 65663 | 114996 | 57.1 | 82.2 | 51500 | 103.8
ACN-H2O 80:20 | 7-COOH-CBD | HQC | 3 | 97166 | 62070 | 113681 | 54.6 | 85.5 | 49867 | 110.8
ACN-H2O 80:20 | 7-COOH-CBD | HQC | 4 | 100700 | 67712 | 113420 | 59.7 | 88.8 | 49807 | 96.5
ACN-H2O 80:20 | 7-COOH-CBD | HQC | 5 | 85667 | 58391 | 109758 | 53.2 | 78.1 | 51192 | 93.6
ACN-H2O 80:20 | 7-COOH-CBD | HQC | 6 | 126127 | 66445 | 117394 | 56.6 | 107.4 | 51503 | 98.4
ACN-H2O 80:20 | 6α-OH-CBD | LQC | 1 | 307 | 149 | 320 | 46.5 | 95.9 | 36687 | 93.8
ACN-H2O 80:20 | 6α-OH-CBD | LQC | 2 | 206 | 156 | 318 | 48.9 | 64.8 | 37214 | 98.4
ACN-H2O 80:20 | 6α-OH-CBD | LQC | 3 | 277 | 155 | 310 | 50 | 89.4 | 38798 | 108.8
ACN-H2O 80:20 | 6α-OH-CBD | LQC | 4 | 334 | 159 | 319 | 49.8 | 104.7 | 38222 | 96.3
ACN-H2O 80:20 | 6α-OH-CBD | LQC | 5 | 268 | 138 | 308 | 44.9 | 87 | 38854 | 103.8
ACN-H2O 80:20 | 6α-OH-CBD | LQC | 6 | 274 | 148 | 310 | 47.9 | 88.4 | 38338 | 93.5
ACN-H2O 80:20 | 6α-OH-CBD | HQC | 1 | 9734 | 8186 | 16706 | 49 | 58.3 | 37491 | 101.4
ACN-H2O 80:20 | 6α-OH-CBD | HQC | 2 | 17217 | 7968 | 16775 | 47.5 | 102.6 | 37204 | 111.4
ACN-H2O 80:20 | 6α-OH-CBD | HQC | 3 | 17284 | 8108 | 16822 | 48.2 | 102.7 | 39299 | 102.2
ACN-H2O 80:20 | 6α-OH-CBD | HQC | 4 | 13526 | 7909 | 16900 | 46.8 | 80 | 36823 | 103.6
ACN-H2O 80:20 | 6α-OH-CBD | HQC | 5 | 15015 | 8088 | 16746 | 48.3 | 89.7 | 39540 | 102.9
ACN-H2O 80:20 | 6α-OH-CBD | HQC | 6 | 17432 | 7755 | 16224 | 47.8 | 107.4 | 38264 | 105.9
ACN-H2O 80:20 | CBN | LQC | 1 | 235 | 190 | 328 | 57.8 | 71.6 | 39708 | 95.6
ACN-H2O 80:20 | CBN | LQC | 2 | 253 | 194 | 348 | 55.8 | 72.7 | 40249 | 104.4
ACN-H2O 80:20 | CBN | LQC | 3 | 331 | 195 | 333 | 58.7 | 99.4 | 39708 | 96.9
ACN-H2O 80:20 | CBN | LQC | 4 | 322 | 207 | 347 | 59.7 | 92.8 | 40405 | 102.7
ACN-H2O 80:20 | CBN | LQC | 5 | 276 | 181 | 329 | 55 | 83.9 | 38177 | 91
ACN-H2O 80:20 | CBN | LQC | 6 | 324 | 187 | 332 | 56.2 | 97.6 | 41353 | 102.6
ACN-H2O 80:20 | CBN | HQC | 1 | 15446 | 9925 | 17691 | 56.1 | 87.3 | 39220 | 100.2
ACN-H2O 80:20 | CBN | HQC | 2 | 16414 | 9919 | 17494 | 56.7 | 93.8 | 40952 | 99.8
ACN-H2O 80:20 | CBN | HQC | 3 | 16265 | 9918 | 17807 | 55.7 | 91.3 | 39418 | 97.6
ACN-H2O 80:20 | CBN | HQC | 4 | 12948 | 9512 | 17550 | 54.2 | 73.8 | 39588 | 102.4
ACN-H2O 80:20 | CBN | HQC | 5 | 15194 | 10576 | 17987 | 58.8 | 84.5 | 42818 | 99.1
ACN-H2O 80:20 | CBN | HQC | 6 | 15789 | 9909 | 17114 | 57.9 | 92.3 | 37088 | 94.2
ACN-H2O 80:20 | CBG | LQC | 1 | 190 | 140 | 293 | 47.9 | 64.8 | 35255 | 99.7
ACN-H2O 80:20 | CBG | LQC | 2 | 226 | 144 | 289 | 49.7 | 78.2 | 32002 | 100.8
ACN-H2O 80:20 | CBG | LQC | 3 | 209 | 148 | 295 | 50.1 | 70.8 | 35047 | 96.3
ACN-H2O 80:20 | CBG | LQC | 4 | 182 | 144 | 296 | 48.5 | 61.5 | 35887 | 96.6
ACN-H2O 80:20 | CBG | LQC | 5 | 298 | 145 | 288 | 50.3 | 103.5 | 35290 | 105
ACN-H2O 80:20 | CBG | LQC | 6 | 293 | 134 | 292 | 46 | 100.3 | 33761 | 105.4
ACN-H2O 80:20 | CBG | HQC | 1 | 9320 | 7502 | 15695 | 47.8 | 59.4 | 34380 | 102
ACN-H2O 80:20 | CBG | HQC | 2 | 12925 | 7115 | 15204 | 46.8 | 85 | 34939 | 102.6
ACN-H2O 80:20 | CBG | HQC | 3 | 10712 | 7722 | 16020 | 48.2 | 66.9 | 34122 | 104.6
ACN-H2O 80:20 | CBG | HQC | 4 | 12471 | 7440 | 14851 | 50.1 | 84 | 34693 | 98.4
ACN-H2O 80:20 | CBG | HQC | 5 | 9510 | 7168 | 14840 | 48.3 | 64.1 | 34861 | 102.6
ACN-H2O 80:20 | CBG | HQC | 6 | 17606 | 7580 | 15040 | 50.4 | 117.1 | 33985 | 98.7
ACN-H2O 80:20 | Δ9-THC | LQC | 1 | 335 | 199 | 392 | 50.7 | 85.5 | 43948 | 102
ACN-H2O 80:20 | Δ9-THC | LQC | 2 | 343 | 186 | 363 | 51.2 | 94.5 | 44419 | 96.1
ACN-H2O 80:20 | Δ9-THC | LQC | 3 | 247 | 195 | 368 | 52.9 | 67.1 | 43670 | 98.9
ACN-H2O 80:20 | Δ9-THC | LQC | 4 | 298 | 185 | 373 | 49.7 | 79.9 | 42842 | 96
ACN-H2O 80:20 | Δ9-THC | LQC | 5 | 353 | 195 | 367 | 53.1 | 96.2 | 45667 | 102.4
ACN-H2O 80:20 | Δ9-THC | LQC | 6 | 349 | 185 | 367 | 50.3 | 95.1 | 44545 | 100.3
ACN-H2O 80:20 | Δ9-THC | HQC | 1 | 15555 | 10721 | 20499 | 52.3 | 75.9 | 44196 | 97.7
ACN-H2O 80:20 | Δ9-THC | HQC | 2 | 12300 | 10572 | 20409 | 51.8 | 60.3 | 45269 | 103.4
ACN-H2O 80:20 | Δ9-THC | HQC | 3 | 18736 | 9814 | 19826 | 49.5 | 94.5 | 46270 | 95
ACN-H2O 80:20 | Δ9-THC | HQC | 4 | 11737 | 10110 | 20180 | 50.1 | 58.2 | 42864 | 98.8
ACN-H2O 80:20 | Δ9-THC | HQC | 5 | 16293 | 11127 | 20454 | 54.4 | 79.7 | 45545 | 101.5
ACN-H2O 80:20 | Δ9-THC | HQC | 6 | 15823 | 11137 | 21213 | 52.5 | 74.6 | 43886 | 97.3
ACN-H2O 80:20 | 11-OH-THC | LQC | 1 | 278 | 151 | 344 | 44 | 80.8 | 41152 | 106.2
ACN-H2O 80:20 | 11-OH-THC | LQC | 2 | 309 | 151 | 346 | 43.6 | 89.3 | 42813 | 87.5
ACN-H2O 80:20 | 11-OH-THC | LQC | 3 | 276 | 147 | 347 | 42.3 | 79.5 | 41159 | 102.5
ACN-H2O 80:20 | 11-OH-THC | LQC | 4 | 317 | 147 | 349 | 42.2 | 90.8 | 42737 | 99.1
ACN-H2O 80:20 | 11-OH-THC | LQC | 5 | 349 | 151 | 358 | 42.2 | 97.5 | 40865 | 94.1
ACN-H2O 80:20 | 11-OH-THC | LQC | 6 | 300 | 157 | 350 | 44.8 | 85.7 | 42439 | 97.2
ACN-H2O 80:20 | 11-OH-THC | HQC | 1 | 11692 | 8324 | 18962 | 43.9 | 61.7 | 41157 | 103.1
ACN-H2O 80:20 | 11-OH-THC | HQC | 2 | 18798 | 8517 | 18355 | 46.4 | 102.4 | 39543 | 106
ACN-H2O 80:20 | 11-OH-THC | HQC | 3 | 14535 | 8342 | 18916 | 44.1 | 76.8 | 40865 | 95.6
ACN-H2O 80:20 | 11-OH-THC | HQC | 4 | 15820 | 8024 | 19196 | 41.8 | 82.4 | 39189 | 102.4
ACN-H2O 80:20 | 11-OH-THC | HQC | 5 | 17129 | 8282 | 19128 | 43.3 | 89.5 | 41570 | 105.8
ACN-H2O 80:20 | 11-OH-THC | HQC | 6 | 17625 | 8613 | 19574 | 44 | 90 | 41040 | 97.1
ACN-H2O 80:20 | THC-COOH | LQC | 1 | 1495 | 823 | 1640 | 50.2 | 91.2 | 51724 | 97.2
ACN-H2O 80:20 | THC-COOH | LQC | 2 | 1690 | 874 | 1738 | 50.3 | 97.2 | 50993 | 98.1
ACN-H2O 80:20 | THC-COOH | LQC | 3 | 1650 | 872 | 1703 | 51.2 | 96.9 | 51275 | 100.9
ACN-H2O 80:20 | THC-COOH | LQC | 4 | 1459 | 912 | 1721 | 53 | 84.8 | 50816 | 101
ACN-H2O 80:20 | THC-COOH | LQC | 5 | 1567 | 918 | 1688 | 54.4 | 92.8 | 48208 | 103.5
ACN-H2O 80:20 | THC-COOH | LQC | 6 | 1541 | 857 | 1649 | 52 | 93.5 | 49133 | 102.6
ACN-H2O 80:20 | THC-COOH | HQC | 1 | 41759 | 22116 | 42612 | 51.9 | 98 | 51285 | 107.6
ACN-H2O 80:20 | THC-COOH | HQC | 2 | 41210 | 23651 | 45483 | 52 | 90.6 | 47971 | 94.1
ACN-H2O 80:20 | THC-COOH | HQC | 3 | 43663 | 24194 | 44392 | 54.5 | 98.4 | 50986 | 103.5
ACN-H2O 80:20 | THC-COOH | HQC | 4 | 41073 | 23035 | 44044 | 52.3 | 93.3 | 48630 | 97.7
ACN-H2O 80:20 | THC-COOH | HQC | 5 | 44585 | 22442 | 43492 | 51.6 | 102.5 | 50744 | 102.6
ACN-H2O 80:20 | THC-COOH | HQC | 6 | 42697 | 21769 | 44066 | 49.4 | 96.9 | 49173 | 97.7
ACN-MeOH 50:50 | CBD | LQC | 1 | 1494 | 830 | 1542 | 53.8 | 96.9 | 48312 | 99.7
ACN-MeOH 50:50 | CBD | LQC | 2 | 1297 | 824 | 1581 | 52.1 | 82 | 47139 | 99.4
ACN-MeOH 50:50 | CBD | LQC | 3 | 1092 | 886 | 1580 | 56.1 | 69.1 | 50006 | 99.9
ACN-MeOH 50:50 | CBD | LQC | 4 | 1105 | 862 | 1652 | 52.2 | 66.9 | 48518 | 96.5
ACN-MeOH 50:50 | CBD | LQC | 5 | 1082 | 889 | 1641 | 54.2 | 65.9 | 49140 | 95.8
ACN-MeOH 50:50 | CBD | LQC | 6 | 1115 | 880 | 1575 | 55.9 | 70.8 | 47569 | 104.3
ACN-MeOH 50:50 | CBD | HQC | 1 | 29723 | 23862 | 42233 | 56.5 | 70.4 | 47184 | 101.9
ACN-MeOH 50:50 | CBD | HQC | 2 | 31052 | 21741 | 41099 | 52.9 | 75.6 | 49718 | 100.3
ACN-MeOH 50:50 | CBD | HQC | 3 | 36561 | 22903 | 41946 | 54.6 | 87.2 | 48785 | 93.5
ACN-MeOH 50:50 | CBD | HQC | 4 | 44905 | 23034 | 43625 | 52.8 | 102.9 | 49886 | 96.6
ACN-MeOH 50:50 | CBD | HQC | 5 | 28945 | 22383 | 42311 | 52.9 | 68.4 | 47502 | 99.4
ACN-MeOH 50:50 | CBD | HQC | 6 | 38650 | 22015 | 39243 | 56.1 | 98.5 | 47262 | 101.1
ACN-MeOH 50:50 | 7-OH-CBD | LQC | 1 | 1427 | 717 | 1469 | 48.8 | 97.1 | 43401 | 108.1
ACN-MeOH 50:50 | 7-OH-CBD | LQC | 2 | 1015 | 722 | 1436 | 50.3 | 70.7 | 43032 | 102
ACN-MeOH 50:50 | 7-OH-CBD | LQC | 3 | 959 | 720 | 1382 | 52.1 | 69.4 | 43103 | 89.8
ACN-MeOH 50:50 | 7-OH-CBD | LQC | 4 | 1375 | 764 | 1515 | 50.4 | 90.8 | 44169 | 95.4
ACN-MeOH 50:50 | 7-OH-CBD | LQC | 5 | 1376 | 761 | 1458 | 52.2 | 94.4 | 41865 | 100.1
ACN-MeOH 50:50 | 7-OH-CBD | LQC | 6 | 1310 | 739 | 1427 | 51.8 | 91.8 | 46844 | 101.9
ACN-MeOH 50:50 | 7-OH-CBD | HQC | 1 | 33531 | 18801 | 38213 | 49.2 | 87.7 | 43253 | 96.5
ACN-MeOH 50:50 | 7-OH-CBD | HQC | 2 | 35247 | 19813 | 37811 | 52.4 | 93.2 | 44046 | 107.2
ACN-MeOH 50:50 | 7-OH-CBD | HQC | 3 | 34072 | 18456 | 38611 | 47.8 | 88.2 | 44914 | 100.7
ACN-MeOH 50:50 | 7-OH-CBD | HQC | 4 | 28765 | 18539 | 37913 | 48.9 | 75.9 | 43076 | 95
ACN-MeOH 50:50 | 7-OH-CBD | HQC | 5 | 35496 | 19441 | 38650 | 50.3 | 91.8 | 43890 | 97.9
ACN-MeOH 50:50 | 7-OH-CBD | HQC | 6 | 28841 | 19016 | 37433 | 50.8 | 77 | 45679 | 101.8
ACN-MeOH 50:50 | 7-COOH-CBD | LQC | 1 | 3774 | 2355 | 4418 | 53.3 | 85.4 | 52900 | 95.6
ACN-MeOH 50:50 | 7-COOH-CBD | LQC | 2 | 4104 | 2255 | 4311 | 52.3 | 95.2 | 52924 | 100.2
ACN-MeOH 50:50 | 7-COOH-CBD | LQC | 3 | 4078 | 2168 | 4318 | 50.2 | 94.4 | 53127 | 97.4
ACN-MeOH 50:50 | 7-COOH-CBD | LQC | 4 | 4436 | 2275 | 4317 | 52.7 | 102.8 | 47054 | 102.1
ACN-MeOH 50:50 | 7-COOH-CBD | LQC | 5 | 4199 | 2233 | 4230 | 52.8 | 99.3 | 50368 | 100.5
ACN-MeOH 50:50 | 7-COOH-CBD | LQC | 6 | 4055 | 2424 | 4360 | 55.6 | 93 | 53094 | 100.1
ACN-MeOH 50:50 | 7-COOH-CBD | HQC | 1 | 107103 | 63861 | 118262 | 54 | 90.6 | 52216 | 99.3
ACN-MeOH 50:50 | 7-COOH-CBD | HQC | 2 | 88884 | 66512 | 118771 | 56 | 74.8 | 51365 | 97.5
ACN-MeOH 50:50 | 7-COOH-CBD | HQC | 3 | 105744 | 59367 | 112225 | 52.9 | 94.2 | 50018 | 100
ACN-MeOH 50:50 | 7-COOH-CBD | HQC | 4 | 102635 | 59284 | 112281 | 52.8 | 91.4 | 51605 | 100.7
ACN-MeOH 50:50 | 7-COOH-CBD | HQC | 5 | 106445 | 61676 | 111935 | 55.1 | 95.1 | 48380 | 102.4
ACN-MeOH 50:50 | 7-COOH-CBD | HQC | 6 | 134798 | 60073 | 114207 | 52.6 | 118 | 52454 | 97.8
ACN-MeOH 50:50 | 6α-OH-CBD | LQC | 1 | 248 | 158 | 315 | 50.2 | 78.7 | 35617 | 100.3
ACN-MeOH 50:50 | 6α-OH-CBD | LQC | 2 | 186 | 164 | 322 | 50.9 | 57.8 | 36559 | 107.8
ACN-MeOH 50:50 | 6α-OH-CBD | LQC | 3 | 198 | 158 | 310 | 51 | 63.9 | 36409 | 98.7
ACN-MeOH 50:50 | 6α-OH-CBD | LQC | 4 | 286 | 163 | 309 | 52.8 | 92.6 | 37521 | 108.5
ACN-MeOH 50:50 | 6α-OH-CBD | LQC | 5 | 251 | 158 | 315 | 50 | 79.7 | 39505 | 95.6
ACN-MeOH 50:50 | 6α-OH-CBD | LQC | 6 | 277 | 168 | 322 | 52.1 | 86 | 37983 | 99.1
ACN-MeOH 50:50 | 6α-OH-CBD | HQC | 1 | 14725 | 8626 | 16064 | 53.7 | 91.7 | 38990 | 94.8
ACN-MeOH 50:50 | 6α-OH-CBD | HQC | 2 | 12563 | 8620 | 16902 | 51 | 74.3 | 37135 | 106.1
ACN-MeOH 50:50 | 6α-OH-CBD | HQC | 3 | 12905 | 8117 | 16633 | 48.8 | 77.6 | 39329 | 108.7
ACN-MeOH 50:50 | 6α-OH-CBD | HQC | 4 | 17924 | 9393 | 17858 | 52.6 | 100.4 | 40905 | 101.1
ACN-MeOH 50:50 | 6α-OH-CBD | HQC | 5 | 11788 | 8605 | 16580 | 51.9 | 71.1 | 38148 | 101
ACN-MeOH 50:50 | 6α-OH-CBD | HQC | 6 | 15106 | 8125 | 16650 | 48.8 | 90.7 | 36079 | 100.7
ACN-MeOH 50:50 | CBN | LQC | 1 | 233 | 186 | 322 | 57.9 | 72.4 | 38506 | 102.6
ACN-MeOH 50:50 | CBN | LQC | 2 | 252 | 190 | 337 | 56.4 | 74.8 | 41161 | 98.9
ACN-MeOH 50:50 | CBN | LQC | 3 | 293 | 175 | 327 | 53.4 | 89.6 | 39073 | 104.1
ACN-MeOH 50:50 | CBN | LQC | 4 | 209 | 185 | 324 | 57.2 | 64.5 | 39851 | 96.8
ACN-MeOH 50:50 | CBN | LQC | 5 | 275 | 190 | 339 | 56 | 81.1 | 39529 | 100
ACN-MeOH 50:50 | CBN | LQC | 6 | 273 | 189 | 345 | 54.8 | 79.1 | 38615 | 103.4
ACN-MeOH 50:50 | CBN | HQC | 1 | 15058 | 10789 | 18412 | 58.6 | 81.8 | 38473 | 97.5
ACN-MeOH 50:50 | CBN | HQC | 2 | 15643 | 9972 | 17588 | 56.7 | 88.9 | 40366 | 97.6
ACN-MeOH 50:50 | CBN | HQC | 3 | 10893 | 10431 | 18494 | 56.4 | 58.9 | 38451 | 96.4
ACN-MeOH 50:50 | CBN | HQC | 4 | 16740 | 10261 | 17187 | 59.7 | 97.4 | 40652 | 96.5
ACN-MeOH 50:50 | CBN | HQC | 5 | 16498 | 10052 | 17481 | 57.5 | 94.4 | 40099 | 92.9
ACN-MeOH 50:50 | CBN | HQC | 6 | 17245 | 10672 | 18242 | 58.5 | 94.5 | 40521 | 97.1
ACN-MeOH 50:50 | CBG | LQC | 1 | 252 | 145 | 290 | 50.1 | 86.9 | 34794 | 103.3
ACN-MeOH 50:50 | CBG | LQC | 2 | 197 | 156 | 301 | 51.7 | 65.4 | 35324 | 98.7
ACN-MeOH 50:50 | CBG | LQC | 3 | 131 | 142 | 282 | 50.4 | 46.5 | 36605 | 102.9
ACN-MeOH 50:50 | CBG | LQC | 4 | 185 | 155 | 299 | 51.9 | 61.9 | 35121 | 108.4
ACN-MeOH 50:50 | CBG | LQC | 5 | 252 | 157 | 294 | 53.4 | 85.7 | 35593 | 105.6
ACN-MeOH 50:50 | CBG | LQC | 6 | 208 | 151 | 292 | 51.7 | 71.2 | 35717 | 98
ACN-MeOH 50:50 | CBG | HQC | 1 | 10919 | 7509 | 14929 | 50.3 | 73.1 | 35848 | 98.7
ACN-MeOH 50:50 | CBG | HQC | 2 | 11119 | 7865 | 15575 | 50.5 | 71.4 | 33662 | 101.7
ACN-MeOH 50:50 | CBG | HQC | 3 | 11205 | 7345 | 14808 | 49.6 | 75.7 | 35009 | 99.1
ACN-MeOH 50:50 | CBG | HQC | 4 | 12214 | 7653 | 15005 | 51 | 81.4 | 35418 | 102.1
ACN-MeOH 50:50 | CBG | HQC | 5 | 12581 | 7149 | 15147 | 47.2 | 83.1 | 33916 | 95
ACN-MeOH 50:50 | CBG | HQC | 6 | 18298 | 7804 | 15362 | 50.8 | 119.1 | 35565 | 95.3
ACN-MeOH 50:50 | Δ9-THC | LQC | 1 | 285 | 229 | 385 | 59.5 | 74 | 45673 | 95.4
ACN-MeOH 50:50 | Δ9-THC | LQC | 2 | 347 | 217 | 381 | 57 | 91.1 | 44970 | 103.9
ACN-MeOH 50:50 | Δ9-THC | LQC | 3 | 306 | 218 | 376 | 57.9 | 81.4 | 43357 | 94.6
ACN-MeOH 50:50 | Δ9-THC | LQC | 4 | 340 | 223 | 375 | 59.5 | 90.7 | 46413 | 103.1
ACN-MeOH 50:50 | Δ9-THC | LQC | 5 | 295 | 212 | 383 | 55.4 | 77 | 44972 | 104.6
ACN-MeOH 50:50 | Δ9-THC | LQC | 6 | 329 | 229 | 391 | 58.6 | 84.1 | 45136 | 97.1
ACN-MeOH 50:50 | Δ9-THC | HQC | 1 | 22271 | 11524 | 20147 | 57.2 | 110.5 | 45114 | 94.1
ACN-MeOH 50:50 | Δ9-THC | HQC | 2 | 13781 | 11789 | 21015 | 56.1 | 65.6 | 44036 | 96.6
ACN-MeOH 50:50 | Δ9-THC | HQC | 3 | 17310 | 11040 | 20674 | 53.4 | 83.7 | 46200 | 98.9
ACN-MeOH 50:50 | Δ9-THC | HQC | 4 | 13819 | 11663 | 20179 | 57.8 | 68.5 | 45105 | 100.6
ACN-MeOH 50:50 | Δ9-THC | HQC | 5 | 13311 | 11490 | 20088 | 57.2 | 66.3 | 44351 | 107.7
ACN-MeOH 50:50 | Δ9-THC | HQC | 6 | 13063 | 11483 | 19936 | 57.6 | 65.5 | 44204 | 93
ACN-MeOH 50:50 | 11-OH-THC | LQC | 1 | 304 | 161 | 355 | 45.3 | 85.6 | 42868 | 98.8
ACN-MeOH 50:50 | 11-OH-THC | LQC | 2 | 249 | 157 | 354 | 44.3 | 70.3 | 43063 | 99.8
ACN-MeOH 50:50 | 11-OH-THC | LQC | 3 | 327 | 164 | 362 | 45.3 | 90.3 | 41127 | 99.7
ACN-MeOH 50:50 | 11-OH-THC | LQC | 4 | 347 | 155 | 356 | 43.5 | 97.5 | 42155 | 100.3
ACN-MeOH 50:50 | 11-OH-THC | LQC | 5 | 321 | 143 | 354 | 40.5 | 90.7 | 42652 | 101
ACN-MeOH 50:50 | 11-OH-THC | LQC | 6 | 304 | 151 | 353 | 42.9 | 86.1 | 40575 | 107.5
ACN-MeOH 50:50 | 11-OH-THC | HQC | 1 | 13791 | 8176 | 19014 | 43 | 72.5 | 41365 | 101.9
ACN-MeOH 50:50 | 11-OH-THC | HQC | 2 | 13045 | 7791 | 17994 | 43.3 | 72.5 | 40746 | 96.4
ACN-MeOH 50:50 | 11-OH-THC | HQC | 3 | 16648 | 8208 | 19358 | 42.4 | 86 | 42275 | 99.3
ACN-MeOH 50:50 | 11-OH-THC | HQC | 4 | 15130 | 8161 | 18848 | 43.3 | 80.3 | 42269 | 92.4
ACN-MeOH 50:50 | 11-OH-THC | HQC | 5 | 15500 | 8163 | 18468 | 44.2 | 83.9 | 39673 | 95.8
ACN-MeOH 50:50 | 11-OH-THC | HQC | 6 | 21637 | 8242 | 19258 | 42.8 | 112.4 | 41357 | 98.2
ACN-MeOH 50:50 | THC-COOH | LQC | 1 | 1383 | 765 | 1682 | 45.5 | 82.2 | 48461 | 99.9
ACN-MeOH 50:50 | THC-COOH | LQC | 2 | 1465 | 768 | 1638 | 46.9 | 89.4 | 50756 | 102.4
ACN-MeOH 50:50 | THC-COOH | LQC | 3 | 1375 | 732 | 1671 | 43.8 | 82.3 | 49662 | 89.3
ACN-MeOH 50:50 | THC-COOH | LQC | 4 | 1380 | 723 | 1678 | 43.1 | 82.2 | 51436 | 103.8
ACN-MeOH 50:50 | THC-COOH | LQC | 5 | 1622 | 745 | 1685 | 44.2 | 96.3 | 47903 | 103.1
ACN-MeOH 50:50 | THC-COOH | LQC | 6 | 1413 | 751 | 1623 | 46.3 | 87.1 | 50981 | 103
ACN-MeOH 50:50 | THC-COOH | HQC | 1 | 33702 | 21096 | 44134 | 47.8 | 76.4 | 48551 | 99.8
ACN-MeOH 50:50 | THC-COOH | HQC | 2 | 45083 | 20366 | 44178 | 46.1 | 102 | 48891 | 89.7
ACN-MeOH 50:50 | THC-COOH | HQC | 3 | 38049 | 19729 | 43551 | 45.3 | 87.4 | 51786 | 94.4
ACN-MeOH 50:50 | THC-COOH | HQC | 4 | 38924 | 20259 | 45938 | 44.1 | 84.7 | 53881 | 94.4
ACN-MeOH 50:50 | THC-COOH | HQC | 5 | 45095 | 20749 | 46212 | 44.9 | 97.6 | 48836 | 102.2
ACN-MeOH 50:50 | THC-COOH | HQC | 6 | 40702 | 20921 | 46081 | 45.4 | 88.3 | 50623 | 92.6

## EY_ME_Optimised
Extraction Yield and Matrix Effect — Final Method (100 µL FA/MeOH 0.1% V/V, 6 Donors, LQC and HQC)
Analyte | QC Level | Donor | Set A Area (cts) | Set B Area (cts) | Set C Area (cts) | EY% | ME% | IS Area (cts) | IS Set A (cts) | IS Set C (cts) | IS-ME%
CBD | LQC | 1 | 1322 | 907 | 1608 | 68.6 | 82.2 | 49587 | 39720 | 48095 | 99.5
CBD | LQC | 2 | 1439 | 984 | 1618 | 68.4 | 88.9 | 49029 | 43189 | 47691 | 98.2
CBD | LQC | 3 | 1347 | 914 | 1589 | 67.9 | 84.8 | 47673 | 54899 | 47033 | 72.6
CBD | LQC | 4 | 1278 | 893 | 1546 | 69.9 | 82.7 | 47767 | 43515 | 48267 | 91.7
CBD | LQC | 5 | 1388 | 968 | 1528 | 69.7 | 90.8 | 47378 | 52814 | 48570 | 83.5
CBD | LQC | 6 | 1425 | 1046 | 1558 | 73.4 | 91.5 | 47073 | 48972 | 47384 | 88.5
CBD | HQC | 1 | 36469 | 24845 | 43654 | 68.1 | 83.5 | 45957 | 46647 | 47604 | 85.3
CBD | HQC | 2 | 33423 | 23829 | 43221 | 71.3 | 77.3 | 48736 | 46342 | 48201 | 80.4
CBD | HQC | 3 | 34062 | 22883 | 42336 | 67.2 | 80.5 | 45831 | 37937 | 47944 | 101.7
CBD | HQC | 4 | 38056 | 25040 | 43582 | 65.8 | 87.3 | 49635 | 44114 | 47413 | 93.9
CBD | HQC | 5 | 34131 | 23640 | 40339 | 69.3 | 84.6 | 48517 | 43339 | 51086 | 99.7
CBD | HQC | 6 | 37598 | 25193 | 42259 | 67 | 89 | 46909 | 47949 | 49010 | 90.9
7-OH-CBD | LQC | 1 | 1257 | 908 | 1460 | 72.2 | 86.1 | 43230 | 42468 | 44113 | 89.4
7-OH-CBD | LQC | 2 | 1404 | 1089 | 1444 | 77.6 | 97.2 | 41917 | 47690 | 42693 | 87
7-OH-CBD | LQC | 3 | 1338 | 1006 | 1488 | 75.2 | 89.9 | 42315 | 39961 | 43774 | 98.5
7-OH-CBD | LQC | 4 | 1262 | 960 | 1398 | 76.1 | 90.3 | 43219 | 37290 | 43579 | 105.5
7-OH-CBD | LQC | 5 | 1429 | 1048 | 1475 | 73.3 | 96.9 | 45018 | 49030 | 43804 | 86.6
7-OH-CBD | LQC | 6 | 1335 | 952 | 1446 | 71.3 | 92.3 | 43245 | 49674 | 44117 | 82
7-OH-CBD | HQC | 1 | 35797 | 27797 | 38344 | 77.7 | 93.4 | 44376 | 42007 | 43337 | 96.3
7-OH-CBD | HQC | 2 | 37039 | 27993 | 37743 | 75.6 | 98.1 | 42515 | 45296 | 44351 | 96.1
7-OH-CBD | HQC | 3 | 39311 | 28805 | 38620 | 73.3 | 101.8 | 45073 | 45584 | 43308 | 96.7
7-OH-CBD | HQC | 4 | 37039 | 28760 | 38819 | 77.6 | 95.4 | 45199 | 39933 | 41835 | 100
7-OH-CBD | HQC | 5 | 35817 | 28083 | 38447 | 78.4 | 93.2 | 45666 | 40277 | 43589 | 100.8
7-OH-CBD | HQC | 6 | 36957 | 28053 | 39186 | 75.9 | 94.3 | 43469 | 42552 | 46155 | 102.3
7-COOH-CBD | LQC | 1 | 3651 | 3036 | 4282 | 83.2 | 85.3 | 52174 | 46641 | 51075 | 93.4
7-COOH-CBD | LQC | 2 | 3892 | 3390 | 4259 | 87.1 | 91.4 | 52560 | 51305 | 51832 | 92.3
7-COOH-CBD | LQC | 3 | 4311 | 3340 | 4388 | 77.5 | 98.2 | 52215 | 48301 | 51371 | 104.5
7-COOH-CBD | LQC | 4 | 4247 | 3291 | 4338 | 77.5 | 97.9 | 52176 | 51647 | 52828 | 100.1
7-COOH-CBD | LQC | 5 | 4282 | 3447 | 4258 | 80.5 | 100.6 | 50388 | 50905 | 51030 | 100.8
7-COOH-CBD | LQC | 6 | 4095 | 3176 | 4306 | 77.6 | 95.1 | 53454 | 59746 | 51766 | 82.4
7-COOH-CBD | HQC | 1 | 115053 | 94172 | 114272 | 81.9 | 100.7 | 51791 | 52449 | 50627 | 97.2
7-COOH-CBD | HQC | 2 | 120725 | 101949 | 111239 | 84.4 | 108.5 | 53591 | 55400 | 51525 | 100.9
7-COOH-CBD | HQC | 3 | 110171 | 89026 | 116556 | 80.8 | 94.5 | 50735 | 52701 | 51775 | 92.9
7-COOH-CBD | HQC | 4 | 110602 | 92304 | 115025 | 83.5 | 96.2 | 51962 | 51428 | 50640 | 94.7
7-COOH-CBD | HQC | 5 | 109449 | 88394 | 112788 | 80.8 | 97 | 53414 | 47264 | 52355 | 107.5
7-COOH-CBD | HQC | 6 | 107443 | 85077 | 115531 | 79.2 | 93 | 50726 | 49475 | 51131 | 96.1
6α-OH-CBD | LQC | 1 | 319 | 226 | 311 | 70.8 | 102.6 | 39561 | 44013 | 39109 | 91.1
6α-OH-CBD | LQC | 2 | 323 | 237 | 318 | 73.4 | 101.6 | 37474 | 34977 | 37283 | 108.3
6α-OH-CBD | LQC | 3 | 298 | 218 | 315 | 73.2 | 94.6 | 38107 | 37528 | 37821 | 95.3
6α-OH-CBD | LQC | 4 | 304 | 229 | 316 | 75.3 | 96.2 | 36792 | 38332 | 38853 | 97.5
6α-OH-CBD | LQC | 5 | 312 | 226 | 314 | 72.4 | 99.4 | 37514 | 44706 | 37840 | 84.1
6α-OH-CBD | LQC | 6 | 288 | 202 | 315 | 70.1 | 91.4 | 37107 | 37232 | 39630 | 97.3
6α-OH-CBD | HQC | 1 | 14725 | 10952 | 16971 | 74.4 | 86.8 | 37296 | 36011 | 38143 | 91.9
6α-OH-CBD | HQC | 2 | 16641 | 12044 | 16681 | 72.4 | 99.8 | 37677 | 38073 | 38747 | 101.5
6α-OH-CBD | HQC | 3 | 16669 | 11544 | 17187 | 69.3 | 97 | 37422 | 47062 | 39375 | 81.1
6α-OH-CBD | HQC | 4 | 15895 | 11776 | 16855 | 74.1 | 94.3 | 37167 | 38232 | 38959 | 96.1
6α-OH-CBD | HQC | 5 | 15341 | 11004 | 16385 | 71.7 | 93.6 | 38006 | 44414 | 37536 | 79.1
6α-OH-CBD | HQC | 6 | 14367 | 10838 | 17823 | 75.4 | 80.6 | 40372 | 33445 | 37390 | 90.1
CBN | LQC | 1 | 302 | 223 | 331 | 73.8 | 91.2 | 40359 | 34920 | 39048 | 102
CBN | LQC | 2 | 295 | 195 | 332 | 66.1 | 88.9 | 41255 | 36593 | 38731 | 94
CBN | LQC | 3 | 307 | 214 | 328 | 69.7 | 93.6 | 41711 | 40616 | 39600 | 91.3
CBN | LQC | 4 | 299 | 213 | 320 | 71.2 | 93.4 | 40058 | 43259 | 39656 | 85.7
CBN | LQC | 5 | 296 | 220 | 336 | 74.3 | 88.1 | 40087 | 40974 | 40856 | 87.8
CBN | LQC | 6 | 321 | 224 | 345 | 69.8 | 93 | 40555 | 42524 | 39566 | 86.6
CBN | HQC | 1 | 15880 | 10945 | 17792 | 68.9 | 89.3 | 37791 | 37823 | 40135 | 94.7
CBN | HQC | 2 | 17696 | 12572 | 18314 | 71 | 96.6 | 39425 | 44253 | 40579 | 88.6
CBN | HQC | 3 | 16729 | 11638 | 17941 | 69.6 | 93.2 | 40020 | 32633 | 38230 | 109.2
CBN | HQC | 4 | 15868 | 10798 | 17328 | 68 | 91.6 | 39954 | 38755 | 41390 | 97.8
CBN | HQC | 5 | 16827 | 11585 | 17995 | 68.8 | 93.5 | 39481 | 45936 | 40029 | 81.5
CBN | HQC | 6 | 16499 | 11624 | 17232 | 70.5 | 95.7 | 40626 | 45410 | 40205 | 84.8
CBG | LQC | 1 | 274 | 183 | 290 | 66.8 | 94.5 | 35269 | 29577 | 35395 | 113.1
CBG | LQC | 2 | 238 | 163 | 296 | 68.5 | 80.4 | 35400 | 35877 | 34879 | 78.2
CBG | LQC | 3 | 225 | 152 | 293 | 67.6 | 76.8 | 33577 | 30413 | 35260 | 89
CBG | LQC | 4 | 322 | 223 | 291 | 69.3 | 110.7 | 35446 | 38923 | 35258 | 100.2
CBG | LQC | 5 | 290 | 202 | 291 | 69.7 | 99.7 | 32983 | 40397 | 37365 | 92.2
CBG | LQC | 6 | 261 | 171 | 294 | 65.5 | 88.8 | 32869 | 32333 | 34287 | 94.1
CBG | HQC | 1 | 13545 | 9258 | 15107 | 68.3 | 89.7 | 34024 | 32801 | 34477 | 94.2
CBG | HQC | 2 | 14463 | 10102 | 15887 | 69.8 | 91 | 34793 | 37007 | 34034 | 83.7
CBG | HQC | 3 | 11297 | 7664 | 15256 | 67.8 | 74 | 33925 | 32712 | 35313 | 79.9
CBG | HQC | 4 | 14700 | 9731 | 15454 | 66.2 | 95.1 | 37561 | 36698 | 34193 | 88.6
CBG | HQC | 5 | 14452 | 9421 | 15469 | 65.2 | 93.4 | 33791 | 38882 | 33054 | 79.4
CBG | HQC | 6 | 15018 | 10400 | 15443 | 69.3 | 97.2 | 34707 | 40708 | 34504 | 82.4
Δ9-THC | LQC | 1 | 342 | 244 | 371 | 71.3 | 92.2 | 44285 | 43066 | 45405 | 97.2
Δ9-THC | LQC | 2 | 333 | 241 | 388 | 72.4 | 85.8 | 44080 | 37903 | 42295 | 95.8
Δ9-THC | LQC | 3 | 320 | 249 | 375 | 77.8 | 85.3 | 44170 | 37985 | 42450 | 95.4
Δ9-THC | LQC | 4 | 332 | 238 | 386 | 71.7 | 86 | 45507 | 43939 | 43232 | 84.6
Δ9-THC | LQC | 5 | 349 | 226 | 375 | 64.8 | 93.1 | 44829 | 46849 | 43994 | 87.4
Δ9-THC | LQC | 6 | 391 | 276 | 381 | 70.6 | 102.6 | 45568 | 46631 | 46414 | 102.1
Δ9-THC | HQC | 1 | 20108 | 14043 | 20221 | 69.8 | 99.4 | 45726 | 54313 | 46408 | 85
Δ9-THC | HQC | 2 | 16814 | 11903 | 20318 | 70.8 | 82.8 | 45756 | 44100 | 46328 | 86.9
Δ9-THC | HQC | 3 | 18852 | 13497 | 19519 | 71.6 | 96.6 | 46882 | 43057 | 46380 | 104
Δ9-THC | HQC | 4 | 17555 | 13476 | 20492 | 76.8 | 85.7 | 45001 | 42914 | 45554 | 90.9
Δ9-THC | HQC | 5 | 17513 | 12540 | 20482 | 71.6 | 85.5 | 45153 | 42416 | 44193 | 89.1
Δ9-THC | HQC | 6 | 19593 | 14300 | 21005 | 73 | 93.3 | 44498 | 37955 | 44085 | 108.3
11-OH-THC | LQC | 1 | 362 | 281 | 349 | 77.6 | 103.7 | 42893 | 45206 | 42354 | 97.2
11-OH-THC | LQC | 2 | 355 | 261 | 358 | 73.5 | 99.2 | 42469 | 46435 | 43735 | 93.4
11-OH-THC | LQC | 3 | 311 | 252 | 349 | 81 | 89.1 | 44090 | 45286 | 43669 | 85.9
11-OH-THC | LQC | 4 | 353 | 271 | 357 | 76.8 | 98.9 | 42564 | 45295 | 40951 | 89.4
11-OH-THC | LQC | 5 | 346 | 256 | 349 | 74 | 99.1 | 42491 | 38940 | 40830 | 104
11-OH-THC | LQC | 6 | 327 | 261 | 347 | 79.8 | 94.2 | 41662 | 46961 | 43197 | 86.7
11-OH-THC | HQC | 1 | 18222 | 14826 | 18941 | 81.4 | 96.2 | 42502 | 44651 | 44287 | 95.4
11-OH-THC | HQC | 2 | 18640 | 14229 | 18912 | 76.3 | 98.6 | 44616 | 44593 | 41666 | 92.1
11-OH-THC | HQC | 3 | 15534 | 11871 | 18531 | 76.4 | 83.8 | 40284 | 32767 | 42527 | 108.8
11-OH-THC | HQC | 4 | 20023 | 14705 | 19327 | 73.4 | 103.6 | 42861 | 46973 | 43001 | 94.8
11-OH-THC | HQC | 5 | 18247 | 14251 | 18543 | 78.1 | 98.4 | 42002 | 39059 | 40587 | 102.3
11-OH-THC | HQC | 6 | 18910 | 15186 | 19255 | 80.3 | 98.2 | 42358 | 43021 | 42078 | 96.1
THC-COOH | LQC | 1 | 1634 | 1388 | 1680 | 84.9 | 97.3 | 49543 | 52556 | 50225 | 92.9
THC-COOH | LQC | 2 | 1795 | 1426 | 1704 | 79.4 | 105.3 | 50720 | 54353 | 50231 | 97.4
THC-COOH | LQC | 3 | 1619 | 1292 | 1700 | 79.8 | 95.2 | 48329 | 46629 | 51935 | 106.1
THC-COOH | LQC | 4 | 1679 | 1361 | 1688 | 81.1 | 99.5 | 49612 | 49240 | 50330 | 101.7
THC-COOH | LQC | 5 | 1663 | 1308 | 1692 | 78.7 | 98.3 | 50090 | 47459 | 48570 | 100.6
THC-COOH | LQC | 6 | 1727 | 1398 | 1725 | 80.9 | 100.1 | 48583 | 45791 | 48880 | 106.9
THC-COOH | HQC | 1 | 44478 | 34938 | 44269 | 78.6 | 100.5 | 51301 | 51351 | 50600 | 99
THC-COOH | HQC | 2 | 45488 | 36430 | 45189 | 80.1 | 100.7 | 50689 | 48765 | 49856 | 102.9
THC-COOH | HQC | 3 | 46199 | 36558 | 43744 | 79.1 | 105.6 | 51464 | 60472 | 48979 | 85.5
THC-COOH | HQC | 4 | 47458 | 39478 | 44541 | 83.2 | 106.5 | 48997 | 58389 | 51362 | 93.7
THC-COOH | HQC | 5 | 43364 | 35146 | 44857 | 81 | 96.7 | 49979 | 51274 | 50738 | 95.7
THC-COOH | HQC | 6 | 41430 | 34367 | 44271 | 83 | 93.6 | 52365 | 45495 | 49033 | 100.9

## Calibration_Curves
Calibration curves — analyte/IS area ratio vs nominal, 1/x² weighted, 5 days.
Analyte | Protocol | Day | Level | Nominal (ng/mL) | Analyte Area (cts) | IS Area (cts) | Area Ratio | Back-calc (ng/mL) | Accuracy (%)
Δ9-THC | 4-disc | 1 | 1 | 0.5 | 0 | 47216 | 0 | 0.523 | 104.7
Δ9-THC | 4-disc | 1 | 2 | 1.209 | 1 | 48939 | 0 | 1.307 | 108.1
Δ9-THC | 4-disc | 1 | 3 | 2.924 | 3 | 48913 | 0.0001 | 3.092 | 105.7
Δ9-THC | 4-disc | 1 | 4 | 7.071 | 6 | 47580 | 0.0001 | 7.165 | 101.3
Δ9-THC | 4-disc | 1 | 5 | 17.1 | 15 | 48854 | 0.0003 | 15.987 | 93.5
Δ9-THC | 4-disc | 1 | 6 | 41.352 | 41 | 48059 | 0.0008 | 44.696 | 108.1
Δ9-THC | 4-disc | 1 | 7 | 100 | 94 | 48783 | 0.0019 | 101.548 | 101.5
Δ9-THC | 4-disc | 2 | 1 | 0.5 | 0 | 48792 | 0 | 0.531 | 106.3
Δ9-THC | 4-disc | 2 | 2 | 1.209 | 1 | 47605 | 0 | 1.109 | 91.7
Δ9-THC | 4-disc | 2 | 3 | 2.924 | 3 | 47865 | 0.0001 | 2.799 | 95.7
Δ9-THC | 4-disc | 2 | 4 | 7.071 | 7 | 48746 | 0.0001 | 7.329 | 103.6
Δ9-THC | 4-disc | 2 | 5 | 17.1 | 17 | 48460 | 0.0003 | 18.158 | 106.2
Δ9-THC | 4-disc | 2 | 6 | 41.352 | 39 | 48598 | 0.0008 | 42.426 | 102.6
Δ9-THC | 4-disc | 2 | 7 | 100 | 86 | 48152 | 0.0018 | 94.066 | 94.1
Δ9-THC | 4-disc | 3 | 1 | 0.5 | 0 | 47333 | 0 | 0.509 | 101.7
Δ9-THC | 4-disc | 3 | 2 | 1.209 | 1 | 48133 | 0 | 1.315 | 108.8
Δ9-THC | 4-disc | 3 | 3 | 2.924 | 3 | 48126 | 0.0001 | 3.003 | 102.7
Δ9-THC | 4-disc | 3 | 4 | 7.071 | 6 | 47452 | 0.0001 | 6.836 | 96.7
Δ9-THC | 4-disc | 3 | 5 | 17.1 | 16 | 47154 | 0.0003 | 17.511 | 102.4
Δ9-THC | 4-disc | 3 | 6 | 41.352 | 35 | 47647 | 0.0007 | 38.465 | 93
Δ9-THC | 4-disc | 3 | 7 | 100 | 99 | 48756 | 0.002 | 107.452 | 107.5
Δ9-THC | 4-disc | 4 | 1 | 0.5 | 0 | 48351 | 0 | 0.479 | 95.8
Δ9-THC | 4-disc | 4 | 2 | 1.209 | 1 | 48545 | 0 | 1.181 | 97.7
Δ9-THC | 4-disc | 4 | 3 | 2.924 | 3 | 48898 | 0.0001 | 2.721 | 93
Δ9-THC | 4-disc | 4 | 4 | 7.071 | 6 | 48082 | 0.0001 | 6.571 | 92.9
Δ9-THC | 4-disc | 4 | 5 | 17.1 | 17 | 48739 | 0.0003 | 18.397 | 107.6
Δ9-THC | 4-disc | 4 | 6 | 41.352 | 36 | 48679 | 0.0007 | 38.725 | 93.6
Δ9-THC | 4-disc | 4 | 7 | 100 | 86 | 47331 | 0.0018 | 96.172 | 96.2
Δ9-THC | 4-disc | 5 | 1 | 0.5 | 0 | 47109 | 0 | 0.52 | 103.9
Δ9-THC | 4-disc | 5 | 2 | 1.209 | 1 | 47226 | 0 | 1.321 | 109.3
Δ9-THC | 4-disc | 5 | 3 | 2.924 | 3 | 47656 | 0.0001 | 3.175 | 108.6
Δ9-THC | 4-disc | 5 | 4 | 7.071 | 7 | 48950 | 0.0001 | 7.051 | 99.7
Δ9-THC | 4-disc | 5 | 5 | 17.1 | 16 | 47964 | 0.0003 | 17.302 | 101.2
Δ9-THC | 4-disc | 5 | 6 | 41.352 | 35 | 48904 | 0.0007 | 38.015 | 91.9
Δ9-THC | 4-disc | 5 | 7 | 100 | 85 | 47876 | 0.0018 | 93.705 | 93.7
11-OH-THC | 4-disc | 1 | 1 | 0.5 | 1 | 47749 | 0 | 0.486 | 97.2
11-OH-THC | 4-disc | 1 | 2 | 1.209 | 1 | 48389 | 0 | 1.118 | 92.4
11-OH-THC | 4-disc | 1 | 3 | 2.924 | 3 | 47653 | 0.0001 | 2.75 | 94.1
11-OH-THC | 4-disc | 1 | 4 | 7.071 | 8 | 47047 | 0.0002 | 6.78 | 95.9
11-OH-THC | 4-disc | 1 | 5 | 17.1 | 22 | 47093 | 0.0005 | 18.761 | 109.7
11-OH-THC | 4-disc | 1 | 6 | 41.352 | 54 | 48708 | 0.0011 | 44.407 | 107.4
11-OH-THC | 4-disc | 1 | 7 | 100 | 112 | 47110 | 0.0024 | 95.115 | 95.1
11-OH-THC | 4-disc | 2 | 1 | 0.5 | 1 | 47426 | 0 | 0.521 | 104.2
11-OH-THC | 4-disc | 2 | 2 | 1.209 | 2 | 47978 | 0 | 1.285 | 106.3
11-OH-THC | 4-disc | 2 | 3 | 2.924 | 4 | 48116 | 0.0001 | 3.053 | 104.4
11-OH-THC | 4-disc | 2 | 4 | 7.071 | 9 | 48313 | 0.0002 | 7.475 | 105.7
11-OH-THC | 4-disc | 2 | 5 | 17.1 | 21 | 47482 | 0.0004 | 17.91 | 104.7
11-OH-THC | 4-disc | 2 | 6 | 41.352 | 51 | 47954 | 0.0011 | 42.317 | 102.3
11-OH-THC | 4-disc | 2 | 7 | 100 | 118 | 47330 | 0.0025 | 100.148 | 100.1
11-OH-THC | 4-disc | 3 | 1 | 0.5 | 1 | 47072 | 0 | 0.551 | 110.2
11-OH-THC | 4-disc | 3 | 2 | 1.209 | 2 | 48794 | 0 | 1.246 | 103.1
11-OH-THC | 4-disc | 3 | 3 | 2.924 | 3 | 47240 | 0.0001 | 2.937 | 100.4
11-OH-THC | 4-disc | 3 | 4 | 7.071 | 9 | 48468 | 0.0002 | 7.28 | 103
11-OH-THC | 4-disc | 3 | 5 | 17.1 | 19 | 48005 | 0.0004 | 16.004 | 93.6
11-OH-THC | 4-disc | 3 | 6 | 41.352 | 46 | 47319 | 0.001 | 39.19 | 94.8
11-OH-THC | 4-disc | 3 | 7 | 100 | 134 | 48643 | 0.0028 | 110.189 | 110.2
11-OH-THC | 4-disc | 4 | 1 | 0.5 | 1 | 48016 | 0 | 0.454 | 90.8
11-OH-THC | 4-disc | 4 | 2 | 1.209 | 1 | 48699 | 0 | 1.203 | 99.5
11-OH-THC | 4-disc | 4 | 3 | 2.924 | 4 | 48285 | 0.0001 | 3.062 | 104.7
11-OH-THC | 4-disc | 4 | 4 | 7.071 | 8 | 47216 | 0.0002 | 7.11 | 100.6
11-OH-THC | 4-disc | 4 | 5 | 17.1 | 20 | 48936 | 0.0004 | 16.708 | 97.7
11-OH-THC | 4-disc | 4 | 6 | 41.352 | 46 | 47601 | 0.001 | 38.342 | 92.7
11-OH-THC | 4-disc | 4 | 7 | 100 | 118 | 47175 | 0.0025 | 99.93 | 99.9
11-OH-THC | 4-disc | 5 | 1 | 0.5 | 1 | 47602 | 0 | 0.479 | 95.9
11-OH-THC | 4-disc | 5 | 2 | 1.209 | 2 | 47620 | 0 | 1.285 | 106.3
11-OH-THC | 4-disc | 5 | 3 | 2.924 | 4 | 47898 | 0.0001 | 3.063 | 104.8
11-OH-THC | 4-disc | 5 | 4 | 7.071 | 9 | 48205 | 0.0002 | 7.369 | 104.2
11-OH-THC | 4-disc | 5 | 5 | 17.1 | 19 | 47644 | 0.0004 | 15.838 | 92.6
11-OH-THC | 4-disc | 5 | 6 | 41.352 | 47 | 48752 | 0.001 | 38.616 | 93.4
11-OH-THC | 4-disc | 5 | 7 | 100 | 107 | 47326 | 0.0023 | 90.34 | 90.3
THC-COOH | 1-disc | 1 | 1 | 2 | 2 | 47195 | 0.0001 | 2.182 | 109.1
THC-COOH | 1-disc | 1 | 2 | 4.309 | 5 | 48937 | 0.0001 | 4.219 | 97.9
THC-COOH | 1-disc | 1 | 3 | 9.283 | 9 | 47564 | 0.0002 | 8.555 | 92.2
THC-COOH | 1-disc | 1 | 4 | 20 | 20 | 47164 | 0.0004 | 18.017 | 90.1
THC-COOH | 1-disc | 1 | 5 | 43.089 | 53 | 47919 | 0.0011 | 47.668 | 110.6
THC-COOH | 1-disc | 1 | 6 | 92.832 | 108 | 48624 | 0.0022 | 94.859 | 102.2
THC-COOH | 1-disc | 1 | 7 | 200 | 204 | 47542 | 0.0043 | 183.541 | 91.8
THC-COOH | 1-disc | 2 | 1 | 2 | 2 | 48371 | 0 | 2.108 | 105.4
THC-COOH | 1-disc | 2 | 2 | 4.309 | 5 | 48809 | 0.0001 | 4.777 | 110.9
THC-COOH | 1-disc | 2 | 3 | 9.283 | 10 | 47358 | 0.0002 | 8.666 | 93.3
THC-COOH | 1-disc | 2 | 4 | 20 | 23 | 48918 | 0.0005 | 19.95 | 99.8
THC-COOH | 1-disc | 2 | 5 | 43.089 | 47 | 48857 | 0.001 | 41 | 95.2
THC-COOH | 1-disc | 2 | 6 | 92.832 | 93 | 47113 | 0.002 | 84.597 | 91.1
THC-COOH | 1-disc | 2 | 7 | 200 | 213 | 48645 | 0.0044 | 187.716 | 93.9
THC-COOH | 1-disc | 3 | 1 | 2 | 2 | 47889 | 0 | 1.789 | 89.5
THC-COOH | 1-disc | 3 | 2 | 4.309 | 5 | 48680 | 0.0001 | 4.251 | 98.6
THC-COOH | 1-disc | 3 | 3 | 9.283 | 10 | 47377 | 0.0002 | 9.289 | 100.1
THC-COOH | 1-disc | 3 | 4 | 20 | 24 | 47849 | 0.0005 | 21.51 | 107.6
THC-COOH | 1-disc | 3 | 5 | 43.089 | 45 | 48084 | 0.0009 | 40.198 | 93.3
THC-COOH | 1-disc | 3 | 6 | 92.832 | 108 | 48956 | 0.0022 | 94.119 | 101.4
THC-COOH | 1-disc | 3 | 7 | 200 | 224 | 47802 | 0.0047 | 200.405 | 100.2
THC-COOH | 1-disc | 4 | 1 | 2 | 2 | 47353 | 0.0001 | 2.167 | 108.3
THC-COOH | 1-disc | 4 | 2 | 4.309 | 5 | 48712 | 0.0001 | 4.36 | 101.2
THC-COOH | 1-disc | 4 | 3 | 9.283 | 10 | 47867 | 0.0002 | 9.331 | 100.5
THC-COOH | 1-disc | 4 | 4 | 20 | 21 | 47139 | 0.0004 | 19.139 | 95.7
THC-COOH | 1-disc | 4 | 5 | 43.089 | 54 | 48287 | 0.0011 | 48.14 | 111.7
THC-COOH | 1-disc | 4 | 6 | 92.832 | 115 | 47750 | 0.0024 | 103.336 | 111.3
THC-COOH | 1-disc | 4 | 7 | 200 | 212 | 47474 | 0.0045 | 191.203 | 95.6
THC-COOH | 1-disc | 5 | 1 | 2 | 2 | 48754 | 0 | 1.837 | 91.9
THC-COOH | 1-disc | 5 | 2 | 4.309 | 5 | 47695 | 0.0001 | 4.4 | 102.1
THC-COOH | 1-disc | 5 | 3 | 9.283 | 11 | 47832 | 0.0002 | 10.24 | 110.3
THC-COOH | 1-disc | 5 | 4 | 20 | 25 | 47782 | 0.0005 | 22.207 | 111
THC-COOH | 1-disc | 5 | 5 | 43.089 | 47 | 48522 | 0.001 | 41.379 | 96
THC-COOH | 1-disc | 5 | 6 | 92.832 | 114 | 47344 | 0.0024 | 103.242 | 111.2
THC-COOH | 1-disc | 5 | 7 | 200 | 197 | 47619 | 0.0041 | 177.489 | 88.7
CBD | 1-disc | 1 | 1 | 2 | 2 | 48418 | 0 | 2.04 | 102
CBD | 1-disc | 1 | 2 | 4.309 | 5 | 47117 | 0.0001 | 4.272 | 99.1
CBD | 1-disc | 1 | 3 | 9.283 | 10 | 47199 | 0.0002 | 9.104 | 98.1
CBD | 1-disc | 1 | 4 | 20 | 21 | 47172 | 0.0005 | 20.078 | 100.4
CBD | 1-disc | 1 | 5 | 43.089 | 45 | 48371 | 0.0009 | 40.892 | 94.9
CBD | 1-disc | 1 | 6 | 92.832 | 99 | 48184 | 0.0021 | 90.672 | 97.7
CBD | 1-disc | 1 | 7 | 200 | 223 | 47201 | 0.0047 | 208.852 | 104.4
CBD | 1-disc | 2 | 1 | 2 | 2 | 48881 | 0 | 2.039 | 101.9
CBD | 1-disc | 2 | 2 | 4.309 | 4 | 47698 | 0.0001 | 4.052 | 94
CBD | 1-disc | 2 | 3 | 9.283 | 10 | 48305 | 0.0002 | 9.279 | 100
CBD | 1-disc | 2 | 4 | 20 | 22 | 48262 | 0.0004 | 19.728 | 98.6
CBD | 1-disc | 2 | 5 | 43.089 | 44 | 47669 | 0.0009 | 40.711 | 94.5
CBD | 1-disc | 2 | 6 | 92.832 | 100 | 47428 | 0.0021 | 92.991 | 100.2
CBD | 1-disc | 2 | 7 | 200 | 225 | 47097 | 0.0048 | 211.318 | 105.7
CBD | 1-disc | 3 | 1 | 2 | 2 | 48258 | 0 | 2.115 | 105.7
CBD | 1-disc | 3 | 2 | 4.309 | 5 | 47623 | 0.0001 | 4.184 | 97.1
CBD | 1-disc | 3 | 3 | 9.283 | 10 | 47709 | 0.0002 | 9.542 | 102.8
CBD | 1-disc | 3 | 4 | 20 | 22 | 47156 | 0.0005 | 20.221 | 101.1
CBD | 1-disc | 3 | 5 | 43.089 | 48 | 47131 | 0.001 | 45.324 | 105.2
CBD | 1-disc | 3 | 6 | 92.832 | 94 | 47798 | 0.002 | 87.003 | 93.7
CBD | 1-disc | 3 | 7 | 200 | 225 | 47134 | 0.0048 | 211.121 | 105.6
CBD | 1-disc | 4 | 1 | 2 | 2 | 48650 | 0 | 2.033 | 101.6
CBD | 1-disc | 4 | 2 | 4.309 | 5 | 48878 | 0.0001 | 4.225 | 98.1
CBD | 1-disc | 4 | 3 | 9.283 | 10 | 48094 | 0.0002 | 8.875 | 95.6
CBD | 1-disc | 4 | 4 | 20 | 22 | 48650 | 0.0004 | 19.689 | 98.4
CBD | 1-disc | 4 | 5 | 43.089 | 46 | 48293 | 0.001 | 42.264 | 98.1
CBD | 1-disc | 4 | 6 | 92.832 | 99 | 47111 | 0.0021 | 92.867 | 100
CBD | 1-disc | 4 | 7 | 200 | 201 | 47750 | 0.0042 | 186.341 | 93.2
CBD | 1-disc | 5 | 1 | 2 | 2 | 48189 | 0 | 1.882 | 94.1
CBD | 1-disc | 5 | 2 | 4.309 | 5 | 47775 | 0.0001 | 4.344 | 100.8
CBD | 1-disc | 5 | 3 | 9.283 | 10 | 47809 | 0.0002 | 9.347 | 100.7
CBD | 1-disc | 5 | 4 | 20 | 22 | 48705 | 0.0004 | 19.849 | 99.2
CBD | 1-disc | 5 | 5 | 43.089 | 48 | 48743 | 0.001 | 43.912 | 101.9
CBD | 1-disc | 5 | 6 | 92.832 | 102 | 47912 | 0.0021 | 93.721 | 101
CBD | 1-disc | 5 | 7 | 200 | 210 | 48274 | 0.0044 | 192.313 | 96.2
7-OH-CBD | 1-disc | 1 | 1 | 2 | 2 | 47664 | 0 | 1.789 | 89.5
7-OH-CBD | 1-disc | 1 | 2 | 4.309 | 4 | 48834 | 0.0001 | 3.984 | 92.4
7-OH-CBD | 1-disc | 1 | 3 | 9.283 | 10 | 48175 | 0.0002 | 9.201 | 99.1
7-OH-CBD | 1-disc | 1 | 4 | 20 | 20 | 47466 | 0.0004 | 18.232 | 91.2
7-OH-CBD | 1-disc | 1 | 5 | 43.089 | 46 | 48029 | 0.001 | 42.15 | 97.8
7-OH-CBD | 1-disc | 1 | 6 | 92.832 | 113 | 48837 | 0.0023 | 101.042 | 108.8
7-OH-CBD | 1-disc | 1 | 7 | 200 | 223 | 48026 | 0.0046 | 202.541 | 101.3
7-OH-CBD | 1-disc | 2 | 1 | 2 | 2 | 48720 | 0 | 1.936 | 96.8
7-OH-CBD | 1-disc | 2 | 2 | 4.309 | 5 | 48289 | 0.0001 | 4.687 | 108.8
7-OH-CBD | 1-disc | 2 | 3 | 9.283 | 10 | 47784 | 0.0002 | 9.021 | 97.2
7-OH-CBD | 1-disc | 2 | 4 | 20 | 24 | 47788 | 0.0005 | 21.56 | 107.8
7-OH-CBD | 1-disc | 2 | 5 | 43.089 | 47 | 48353 | 0.001 | 42.299 | 98.2
7-OH-CBD | 1-disc | 2 | 6 | 92.832 | 107 | 47823 | 0.0022 | 97.487 | 105
7-OH-CBD | 1-disc | 2 | 7 | 200 | 234 | 47708 | 0.0049 | 213.308 | 106.7
7-OH-CBD | 1-disc | 3 | 1 | 2 | 2 | 47360 | 0 | 2.102 | 105.1
7-OH-CBD | 1-disc | 3 | 2 | 4.309 | 5 | 47729 | 0.0001 | 4.19 | 97.2
7-OH-CBD | 1-disc | 3 | 3 | 9.283 | 10 | 48896 | 0.0002 | 8.676 | 93.5
7-OH-CBD | 1-disc | 3 | 4 | 20 | 22 | 48392 | 0.0004 | 19.471 | 97.4
7-OH-CBD | 1-disc | 3 | 5 | 43.089 | 50 | 48209 | 0.001 | 45.477 | 105.5
7-OH-CBD | 1-disc | 3 | 6 | 92.832 | 105 | 47865 | 0.0022 | 95.691 | 103.1
7-OH-CBD | 1-disc | 3 | 7 | 200 | 217 | 48621 | 0.0045 | 194.504 | 97.3
7-OH-CBD | 1-disc | 4 | 1 | 2 | 2 | 48913 | 0 | 1.853 | 92.6
7-OH-CBD | 1-disc | 4 | 2 | 4.309 | 5 | 47095 | 0.0001 | 4.25 | 98.6
7-OH-CBD | 1-disc | 4 | 3 | 9.283 | 10 | 47204 | 0.0002 | 9.626 | 103.7
7-OH-CBD | 1-disc | 4 | 4 | 20 | 24 | 48146 | 0.0005 | 21.298 | 106.5
7-OH-CBD | 1-disc | 4 | 5 | 43.089 | 46 | 47303 | 0.001 | 42.562 | 98.8
7-OH-CBD | 1-disc | 4 | 6 | 92.832 | 100 | 48180 | 0.0021 | 90.191 | 97.2
7-OH-CBD | 1-disc | 4 | 7 | 200 | 209 | 47193 | 0.0044 | 192.674 | 96.3
7-OH-CBD | 1-disc | 5 | 1 | 2 | 2 | 48105 | 0 | 2.045 | 102.3
7-OH-CBD | 1-disc | 5 | 2 | 4.309 | 4 | 47644 | 0.0001 | 3.983 | 92.4
7-OH-CBD | 1-disc | 5 | 3 | 9.283 | 10 | 48093 | 0.0002 | 8.916 | 96
7-OH-CBD | 1-disc | 5 | 4 | 20 | 23 | 47658 | 0.0005 | 20.587 | 102.9
7-OH-CBD | 1-disc | 5 | 5 | 43.089 | 53 | 48841 | 0.0011 | 47.585 | 110.4
7-OH-CBD | 1-disc | 5 | 6 | 92.832 | 106 | 48703 | 0.0022 | 95.231 | 102.6
7-OH-CBD | 1-disc | 5 | 7 | 200 | 200 | 48738 | 0.0041 | 178.886 | 89.4
7-COOH-CBD | 1-disc | 1 | 1 | 5 | 6 | 48024 | 0.0001 | 5.184 | 103.7
7-COOH-CBD | 1-disc | 1 | 2 | 10.772 | 13 | 47505 | 0.0003 | 11.66 | 108.2
7-COOH-CBD | 1-disc | 1 | 3 | 23.208 | 24 | 48300 | 0.0005 | 21.628 | 93.2
7-COOH-CBD | 1-disc | 1 | 4 | 50 | 50 | 47720 | 0.0011 | 45.178 | 90.4
7-COOH-CBD | 1-disc | 1 | 5 | 107.722 | 114 | 47258 | 0.0024 | 103.397 | 96
7-COOH-CBD | 1-disc | 1 | 6 | 232.079 | 280 | 47500 | 0.0059 | 252.881 | 109
7-COOH-CBD | 1-disc | 1 | 7 | 500 | 519 | 47844 | 0.0109 | 465.98 | 93.2
7-COOH-CBD | 1-disc | 2 | 1 | 5 | 6 | 47839 | 0.0001 | 5.416 | 108.3
7-COOH-CBD | 1-disc | 2 | 2 | 10.772 | 13 | 47264 | 0.0003 | 11.398 | 105.8
7-COOH-CBD | 1-disc | 2 | 3 | 23.208 | 28 | 48215 | 0.0006 | 24.717 | 106.5
7-COOH-CBD | 1-disc | 2 | 4 | 50 | 51 | 47685 | 0.0011 | 45.543 | 91.1
7-COOH-CBD | 1-disc | 2 | 5 | 107.722 | 107 | 47837 | 0.0022 | 96.314 | 89.4
7-COOH-CBD | 1-disc | 2 | 6 | 232.079 | 251 | 48109 | 0.0052 | 224.113 | 96.6
7-COOH-CBD | 1-disc | 2 | 7 | 500 | 497 | 48174 | 0.0103 | 442.651 | 88.5
7-COOH-CBD | 1-disc | 3 | 1 | 5 | 6 | 47079 | 0.0001 | 5.091 | 101.8
7-COOH-CBD | 1-disc | 3 | 2 | 10.772 | 12 | 47872 | 0.0003 | 10.869 | 100.9
7-COOH-CBD | 1-disc | 3 | 3 | 23.208 | 24 | 47272 | 0.0005 | 22.035 | 94.9
7-COOH-CBD | 1-disc | 3 | 4 | 50 | 50 | 47270 | 0.0011 | 45.643 | 91.3
7-COOH-CBD | 1-disc | 3 | 5 | 107.722 | 108 | 47632 | 0.0023 | 96.916 | 90
7-COOH-CBD | 1-disc | 3 | 6 | 232.079 | 258 | 48451 | 0.0053 | 228.261 | 98.4
7-COOH-CBD | 1-disc | 3 | 7 | 500 | 562 | 48621 | 0.0116 | 496.362 | 99.3
7-COOH-CBD | 1-disc | 4 | 1 | 5 | 5 | 47424 | 0.0001 | 4.827 | 96.5
7-COOH-CBD | 1-disc | 4 | 2 | 10.772 | 13 | 48423 | 0.0003 | 11.696 | 108.6
7-COOH-CBD | 1-disc | 4 | 3 | 23.208 | 27 | 48826 | 0.0006 | 23.959 | 103.2
7-COOH-CBD | 1-disc | 4 | 4 | 50 | 52 | 47840 | 0.0011 | 46.811 | 93.6
7-COOH-CBD | 1-disc | 4 | 5 | 107.722 | 131 | 48692 | 0.0027 | 115.622 | 107.3
7-COOH-CBD | 1-disc | 4 | 6 | 232.079 | 240 | 47048 | 0.0051 | 219.331 | 94.5
7-COOH-CBD | 1-disc | 4 | 7 | 500 | 540 | 48750 | 0.0111 | 475.734 | 95.1
7-COOH-CBD | 1-disc | 5 | 1 | 5 | 5 | 47914 | 0.0001 | 4.517 | 90.3
7-COOH-CBD | 1-disc | 5 | 2 | 10.772 | 11 | 47551 | 0.0002 | 10.292 | 95.5
7-COOH-CBD | 1-disc | 5 | 3 | 23.208 | 23 | 48238 | 0.0005 | 20.689 | 89.1
7-COOH-CBD | 1-disc | 5 | 4 | 50 | 57 | 48574 | 0.0012 | 50.038 | 100.1
7-COOH-CBD | 1-disc | 5 | 5 | 107.722 | 130 | 48136 | 0.0027 | 116.313 | 108
7-COOH-CBD | 1-disc | 5 | 6 | 232.079 | 273 | 47455 | 0.0058 | 247.159 | 106.5
7-COOH-CBD | 1-disc | 5 | 7 | 500 | 611 | 48338 | 0.0126 | 542.547 | 108.5
6α-OH-CBD | 4-disc | 1 | 1 | 0.5 | 1 | 48682 | 0 | 0.491 | 98.1
6α-OH-CBD | 4-disc | 1 | 2 | 1.209 | 1 | 48667 | 0 | 1.165 | 96.4
6α-OH-CBD | 4-disc | 1 | 3 | 2.924 | 3 | 48863 | 0.0001 | 3.057 | 104.5
6α-OH-CBD | 4-disc | 1 | 4 | 7.071 | 7 | 48255 | 0.0002 | 6.856 | 97
6α-OH-CBD | 4-disc | 1 | 5 | 17.1 | 18 | 47081 | 0.0004 | 17.622 | 103.1
6α-OH-CBD | 4-disc | 1 | 6 | 41.352 | 42 | 47525 | 0.0009 | 40.296 | 97.4
6α-OH-CBD | 4-disc | 1 | 7 | 100 | 107 | 47519 | 0.0023 | 103.105 | 103.1
6α-OH-CBD | 4-disc | 2 | 1 | 0.5 | 0 | 48280 | 0 | 0.463 | 92.7
6α-OH-CBD | 4-disc | 2 | 2 | 1.209 | 1 | 48802 | 0 | 1.295 | 107.1
6α-OH-CBD | 4-disc | 2 | 3 | 2.924 | 3 | 47071 | 0.0001 | 3.139 | 107.3
6α-OH-CBD | 4-disc | 2 | 4 | 7.071 | 8 | 47651 | 0.0002 | 7.625 | 107.8
6α-OH-CBD | 4-disc | 2 | 5 | 17.1 | 17 | 48632 | 0.0004 | 16.144 | 94.4
6α-OH-CBD | 4-disc | 2 | 6 | 41.352 | 48 | 47783 | 0.001 | 45.571 | 110.2
6α-OH-CBD | 4-disc | 2 | 7 | 100 | 114 | 47319 | 0.0024 | 109.563 | 109.6
6α-OH-CBD | 4-disc | 3 | 1 | 0.5 | 1 | 48896 | 0 | 0.551 | 110.2
6α-OH-CBD | 4-disc | 3 | 2 | 1.209 | 1 | 47967 | 0 | 1.289 | 106.6
6α-OH-CBD | 4-disc | 3 | 3 | 2.924 | 3 | 47876 | 0.0001 | 3.079 | 105.3
6α-OH-CBD | 4-disc | 3 | 4 | 7.071 | 7 | 48488 | 0.0001 | 6.668 | 94.3
6α-OH-CBD | 4-disc | 3 | 5 | 17.1 | 18 | 47635 | 0.0004 | 17.37 | 101.6
6α-OH-CBD | 4-disc | 3 | 6 | 41.352 | 44 | 48185 | 0.0009 | 41.678 | 100.8
6α-OH-CBD | 4-disc | 3 | 7 | 100 | 95 | 48167 | 0.002 | 90.223 | 90.2
6α-OH-CBD | 4-disc | 4 | 1 | 0.5 | 1 | 48189 | 0 | 0.533 | 106.5
6α-OH-CBD | 4-disc | 4 | 2 | 1.209 | 1 | 47654 | 0 | 1.092 | 90.4
6α-OH-CBD | 4-disc | 4 | 3 | 2.924 | 3 | 48378 | 0.0001 | 3.002 | 102.7
6α-OH-CBD | 4-disc | 4 | 4 | 7.071 | 7 | 47592 | 0.0001 | 6.629 | 93.8
6α-OH-CBD | 4-disc | 4 | 5 | 17.1 | 17 | 48652 | 0.0004 | 16.149 | 94.4
6α-OH-CBD | 4-disc | 4 | 6 | 41.352 | 42 | 48189 | 0.0009 | 39.524 | 95.6
6α-OH-CBD | 4-disc | 4 | 7 | 100 | 114 | 47734 | 0.0024 | 108.547 | 108.5
6α-OH-CBD | 4-disc | 5 | 1 | 0.5 | 1 | 48587 | 0 | 0.475 | 95
6α-OH-CBD | 4-disc | 5 | 2 | 1.209 | 1 | 47765 | 0 | 1.086 | 89.8
6α-OH-CBD | 4-disc | 5 | 3 | 2.924 | 3 | 47803 | 0.0001 | 3.053 | 104.4
6α-OH-CBD | 4-disc | 5 | 4 | 7.071 | 8 | 47840 | 0.0002 | 7.281 | 103
6α-OH-CBD | 4-disc | 5 | 5 | 17.1 | 16 | 47416 | 0.0003 | 15.382 | 90
6α-OH-CBD | 4-disc | 5 | 6 | 41.352 | 40 | 47573 | 0.0008 | 37.941 | 91.8
6α-OH-CBD | 4-disc | 5 | 7 | 100 | 112 | 47672 | 0.0024 | 107.532 | 107.5
CBN | 4-disc | 1 | 1 | 0.5 | 0 | 47898 | 0 | 0.474 | 94.7
CBN | 4-disc | 1 | 2 | 1.209 | 1 | 48816 | 0 | 1.103 | 91.2
CBN | 4-disc | 1 | 3 | 2.924 | 3 | 47413 | 0.0001 | 3.204 | 109.6
CBN | 4-disc | 1 | 4 | 7.071 | 7 | 47238 | 0.0001 | 6.736 | 95.3
CBN | 4-disc | 1 | 5 | 17.1 | 19 | 48148 | 0.0004 | 18.205 | 106.5
CBN | 4-disc | 1 | 6 | 41.352 | 42 | 48078 | 0.0009 | 40.581 | 98.1
CBN | 4-disc | 1 | 7 | 100 | 115 | 48280 | 0.0024 | 110.157 | 110.2
CBN | 4-disc | 2 | 1 | 0.5 | 0 | 48405 | 0 | 0.47 | 94
CBN | 4-disc | 2 | 2 | 1.209 | 1 | 48309 | 0 | 1.204 | 99.6
CBN | 4-disc | 2 | 3 | 2.924 | 3 | 48042 | 0.0001 | 2.714 | 92.8
CBN | 4-disc | 2 | 4 | 7.071 | 7 | 47075 | 0.0002 | 6.943 | 98.2
CBN | 4-disc | 2 | 5 | 17.1 | 16 | 48668 | 0.0003 | 15.509 | 90.7
CBN | 4-disc | 2 | 6 | 41.352 | 39 | 47810 | 0.0008 | 37.44 | 90.5
CBN | 4-disc | 2 | 7 | 100 | 112 | 48049 | 0.0023 | 107.715 | 107.7
CBN | 4-disc | 3 | 1 | 0.5 | 0 | 48229 | 0 | 0.459 | 91.9
CBN | 4-disc | 3 | 2 | 1.209 | 1 | 48673 | 0 | 1.318 | 109
CBN | 4-disc | 3 | 3 | 2.924 | 3 | 48874 | 0.0001 | 3.003 | 102.7
CBN | 4-disc | 3 | 4 | 7.071 | 8 | 47374 | 0.0002 | 7.557 | 106.9
CBN | 4-disc | 3 | 5 | 17.1 | 17 | 47092 | 0.0004 | 16.372 | 95.7
CBN | 4-disc | 3 | 6 | 41.352 | 40 | 48943 | 0.0008 | 37.379 | 90.4
CBN | 4-disc | 3 | 7 | 100 | 115 | 48278 | 0.0024 | 109.863 | 109.9
CBN | 4-disc | 4 | 1 | 0.5 | 1 | 48171 | 0 | 0.527 | 105.4
CBN | 4-disc | 4 | 2 | 1.209 | 1 | 48820 | 0 | 1.28 | 105.9
CBN | 4-disc | 4 | 3 | 2.924 | 3 | 47096 | 0.0001 | 2.883 | 98.6
CBN | 4-disc | 4 | 4 | 7.071 | 7 | 47353 | 0.0002 | 7.288 | 103.1
CBN | 4-disc | 4 | 5 | 17.1 | 19 | 47297 | 0.0004 | 18.883 | 110.4
CBN | 4-disc | 4 | 6 | 41.352 | 45 | 47713 | 0.0009 | 43.535 | 105.3
CBN | 4-disc | 4 | 7 | 100 | 110 | 48290 | 0.0023 | 105.182 | 105.2
CBN | 4-disc | 5 | 1 | 0.5 | 1 | 48247 | 0 | 0.537 | 107.5
CBN | 4-disc | 5 | 2 | 1.209 | 1 | 48327 | 0 | 1.234 | 102.1
CBN | 4-disc | 5 | 3 | 2.924 | 3 | 47349 | 0.0001 | 2.951 | 100.9
CBN | 4-disc | 5 | 4 | 7.071 | 7 | 47754 | 0.0002 | 7.034 | 99.5
CBN | 4-disc | 5 | 5 | 17.1 | 18 | 48598 | 0.0004 | 16.954 | 99.1
CBN | 4-disc | 5 | 6 | 41.352 | 45 | 47801 | 0.0009 | 43.055 | 104.1
CBN | 4-disc | 5 | 7 | 100 | 98 | 48431 | 0.002 | 93.792 | 93.8
CBG | 4-disc | 1 | 1 | 0.5 | 1 | 48953 | 0 | 0.543 | 108.6
CBG | 4-disc | 1 | 2 | 1.209 | 1 | 47228 | 0 | 1.066 | 88.1
CBG | 4-disc | 1 | 3 | 2.924 | 3 | 48727 | 0.0001 | 3.158 | 108
CBG | 4-disc | 1 | 4 | 7.071 | 7 | 47321 | 0.0001 | 6.707 | 94.9
CBG | 4-disc | 1 | 5 | 17.1 | 17 | 48816 | 0.0003 | 16.682 | 97.6
CBG | 4-disc | 1 | 6 | 41.352 | 37 | 48895 | 0.0008 | 36.937 | 89.3
CBG | 4-disc | 1 | 7 | 100 | 105 | 48248 | 0.0022 | 105.312 | 105.3
CBG | 4-disc | 2 | 1 | 0.5 | 1 | 48390 | 0 | 0.549 | 109.7
CBG | 4-disc | 2 | 2 | 1.209 | 1 | 47174 | 0 | 1.269 | 105
CBG | 4-disc | 2 | 3 | 2.924 | 3 | 47726 | 0.0001 | 2.646 | 90.5
CBG | 4-disc | 2 | 4 | 7.071 | 7 | 48296 | 0.0001 | 7.108 | 100.5
CBG | 4-disc | 2 | 5 | 17.1 | 15 | 47542 | 0.0003 | 15.463 | 90.4
CBG | 4-disc | 2 | 6 | 41.352 | 45 | 47424 | 0.001 | 46.059 | 111.4
CBG | 4-disc | 2 | 7 | 100 | 108 | 48396 | 0.0022 | 107.554 | 107.6
CBG | 4-disc | 3 | 1 | 0.5 | 1 | 48101 | 0 | 0.504 | 100.9
CBG | 4-disc | 3 | 2 | 1.209 | 1 | 48661 | 0 | 1.135 | 93.9
CBG | 4-disc | 3 | 3 | 2.924 | 3 | 47984 | 0.0001 | 3.035 | 103.8
CBG | 4-disc | 3 | 4 | 7.071 | 7 | 47971 | 0.0001 | 6.923 | 97.9
CBG | 4-disc | 3 | 5 | 17.1 | 17 | 47438 | 0.0004 | 17.294 | 101.1
CBG | 4-disc | 3 | 6 | 41.352 | 42 | 48256 | 0.0009 | 42.085 | 101.8
CBG | 4-disc | 3 | 7 | 100 | 92 | 48726 | 0.0019 | 91.649 | 91.6
CBG | 4-disc | 4 | 1 | 0.5 | 1 | 48060 | 0 | 0.514 | 102.8
CBG | 4-disc | 4 | 2 | 1.209 | 1 | 48677 | 0 | 1.27 | 105
CBG | 4-disc | 4 | 3 | 2.924 | 3 | 47477 | 0.0001 | 3.021 | 103.3
CBG | 4-disc | 4 | 4 | 7.071 | 7 | 47220 | 0.0002 | 7.349 | 103.9
CBG | 4-disc | 4 | 5 | 17.1 | 17 | 47072 | 0.0004 | 17.848 | 104.4
CBG | 4-disc | 4 | 6 | 41.352 | 45 | 47684 | 0.0009 | 45.888 | 111
CBG | 4-disc | 4 | 7 | 100 | 97 | 48677 | 0.002 | 96.596 | 96.6
CBG | 4-disc | 5 | 1 | 0.5 | 1 | 47649 | 0 | 0.526 | 105.3
CBG | 4-disc | 5 | 2 | 1.209 | 1 | 48392 | 0 | 1.259 | 104.1
CBG | 4-disc | 5 | 3 | 2.924 | 3 | 47108 | 0.0001 | 2.786 | 95.3
CBG | 4-disc | 5 | 4 | 7.071 | 7 | 48061 | 0.0001 | 6.591 | 93.2
CBG | 4-disc | 5 | 5 | 17.1 | 17 | 47837 | 0.0003 | 16.824 | 98.4
CBG | 4-disc | 5 | 6 | 41.352 | 43 | 48399 | 0.0009 | 42.729 | 103.3
CBG | 4-disc | 5 | 7 | 100 | 93 | 47650 | 0.002 | 94.625 | 94.6

## Precision_Accuracy
Intra-/inter-day precision (%RSD) & accuracy (%) — 4 QC × 6 rep × 5 days.
Analyte | Protocol | QC Level | Nominal (ng/mL) | Day | Replicate | Measured (ng/mL) | Accuracy (%)
Δ9-THC | 4-disc | LLOQ | 0.5 | 1 | 1 | 0.703 | 140.6
Δ9-THC | 4-disc | LLOQ | 0.5 | 1 | 2 | 0.572 | 114.4
Δ9-THC | 4-disc | LLOQ | 0.5 | 1 | 3 | 0.649 | 129.8
Δ9-THC | 4-disc | LLOQ | 0.5 | 1 | 4 | 0.568 | 113.6
Δ9-THC | 4-disc | LLOQ | 0.5 | 1 | 5 | 0.563 | 112.6
Δ9-THC | 4-disc | LLOQ | 0.5 | 1 | 6 | 0.597 | 119.4
Δ9-THC | 4-disc | LLOQ | 0.5 | 2 | 1 | 0.426 | 85.2
Δ9-THC | 4-disc | LLOQ | 0.5 | 2 | 2 | 0.542 | 108.4
Δ9-THC | 4-disc | LLOQ | 0.5 | 2 | 3 | 0.474 | 94.8
Δ9-THC | 4-disc | LLOQ | 0.5 | 2 | 4 | 0.454 | 90.8
Δ9-THC | 4-disc | LLOQ | 0.5 | 2 | 5 | 0.504 | 100.8
Δ9-THC | 4-disc | LLOQ | 0.5 | 2 | 6 | 0.568 | 113.6
Δ9-THC | 4-disc | LLOQ | 0.5 | 3 | 1 | 0.617 | 123.4
Δ9-THC | 4-disc | LLOQ | 0.5 | 3 | 2 | 0.677 | 135.4
Δ9-THC | 4-disc | LLOQ | 0.5 | 3 | 3 | 0.566 | 113.2
Δ9-THC | 4-disc | LLOQ | 0.5 | 3 | 4 | 0.584 | 116.8
Δ9-THC | 4-disc | LLOQ | 0.5 | 3 | 5 | 0.604 | 120.8
Δ9-THC | 4-disc | LLOQ | 0.5 | 3 | 6 | 0.619 | 123.8
Δ9-THC | 4-disc | LLOQ | 0.5 | 4 | 1 | 0.754 | 150.8
Δ9-THC | 4-disc | LLOQ | 0.5 | 4 | 2 | 0.586 | 117.2
Δ9-THC | 4-disc | LLOQ | 0.5 | 4 | 3 | 0.392 | 78.4
Δ9-THC | 4-disc | LLOQ | 0.5 | 4 | 4 | 0.566 | 113.2
Δ9-THC | 4-disc | LLOQ | 0.5 | 4 | 5 | 0.587 | 117.4
Δ9-THC | 4-disc | LLOQ | 0.5 | 4 | 6 | 0.387 | 77.4
Δ9-THC | 4-disc | LLOQ | 0.5 | 5 | 1 | 0.635 | 127
Δ9-THC | 4-disc | LLOQ | 0.5 | 5 | 2 | 0.55 | 110
Δ9-THC | 4-disc | LLOQ | 0.5 | 5 | 3 | 0.544 | 108.8
Δ9-THC | 4-disc | LLOQ | 0.5 | 5 | 4 | 0.694 | 138.8
Δ9-THC | 4-disc | LLOQ | 0.5 | 5 | 5 | 0.494 | 98.8
Δ9-THC | 4-disc | LLOQ | 0.5 | 5 | 6 | 0.506 | 101.2
Δ9-THC | 4-disc | LQC | 1.5 | 1 | 1 | 1.373 | 91.5
Δ9-THC | 4-disc | LQC | 1.5 | 1 | 2 | 1.734 | 115.6
Δ9-THC | 4-disc | LQC | 1.5 | 1 | 3 | 1.78 | 118.7
Δ9-THC | 4-disc | LQC | 1.5 | 1 | 4 | 1.462 | 97.5
Δ9-THC | 4-disc | LQC | 1.5 | 1 | 5 | 1.346 | 89.7
Δ9-THC | 4-disc | LQC | 1.5 | 1 | 6 | 1.766 | 117.7
Δ9-THC | 4-disc | LQC | 1.5 | 2 | 1 | 1.343 | 89.5
Δ9-THC | 4-disc | LQC | 1.5 | 2 | 2 | 1.382 | 92.1
Δ9-THC | 4-disc | LQC | 1.5 | 2 | 3 | 1.794 | 119.6
Δ9-THC | 4-disc | LQC | 1.5 | 2 | 4 | 1.464 | 97.6
Δ9-THC | 4-disc | LQC | 1.5 | 2 | 5 | 1.675 | 111.7
Δ9-THC | 4-disc | LQC | 1.5 | 2 | 6 | 1.297 | 86.5
Δ9-THC | 4-disc | LQC | 1.5 | 3 | 1 | 1.627 | 108.5
Δ9-THC | 4-disc | LQC | 1.5 | 3 | 2 | 1.822 | 121.5
Δ9-THC | 4-disc | LQC | 1.5 | 3 | 3 | 1.559 | 103.9
Δ9-THC | 4-disc | LQC | 1.5 | 3 | 4 | 1.361 | 90.7
Δ9-THC | 4-disc | LQC | 1.5 | 3 | 5 | 1.376 | 91.7
Δ9-THC | 4-disc | LQC | 1.5 | 3 | 6 | 1.531 | 102.1
Δ9-THC | 4-disc | LQC | 1.5 | 4 | 1 | 1.716 | 114.4
Δ9-THC | 4-disc | LQC | 1.5 | 4 | 2 | 1.404 | 93.6
Δ9-THC | 4-disc | LQC | 1.5 | 4 | 3 | 1.976 | 131.7
Δ9-THC | 4-disc | LQC | 1.5 | 4 | 4 | 1.568 | 104.5
Δ9-THC | 4-disc | LQC | 1.5 | 4 | 5 | 1.515 | 101
Δ9-THC | 4-disc | LQC | 1.5 | 4 | 6 | 1.542 | 102.8
Δ9-THC | 4-disc | LQC | 1.5 | 5 | 1 | 1.358 | 90.5
Δ9-THC | 4-disc | LQC | 1.5 | 5 | 2 | 1.35 | 90
Δ9-THC | 4-disc | LQC | 1.5 | 5 | 3 | 1.511 | 100.7
Δ9-THC | 4-disc | LQC | 1.5 | 5 | 4 | 1.463 | 97.5
Δ9-THC | 4-disc | LQC | 1.5 | 5 | 5 | 1.43 | 95.3
Δ9-THC | 4-disc | LQC | 1.5 | 5 | 6 | 1.553 | 103.5
Δ9-THC | 4-disc | MQC | 50.2 | 1 | 1 | 41.85 | 83.4
Δ9-THC | 4-disc | MQC | 50.2 | 1 | 2 | 44.138 | 87.9
Δ9-THC | 4-disc | MQC | 50.2 | 1 | 3 | 55.846 | 111.2
Δ9-THC | 4-disc | MQC | 50.2 | 1 | 4 | 52.732 | 105
Δ9-THC | 4-disc | MQC | 50.2 | 1 | 5 | 46.837 | 93.3
Δ9-THC | 4-disc | MQC | 50.2 | 1 | 6 | 49.172 | 98
Δ9-THC | 4-disc | MQC | 50.2 | 2 | 1 | 58.432 | 116.4
Δ9-THC | 4-disc | MQC | 50.2 | 2 | 2 | 58.195 | 115.9
Δ9-THC | 4-disc | MQC | 50.2 | 2 | 3 | 48.916 | 97.4
Δ9-THC | 4-disc | MQC | 50.2 | 2 | 4 | 51.381 | 102.4
Δ9-THC | 4-disc | MQC | 50.2 | 2 | 5 | 53.939 | 107.4
Δ9-THC | 4-disc | MQC | 50.2 | 2 | 6 | 58.521 | 116.6
Δ9-THC | 4-disc | MQC | 50.2 | 3 | 1 | 39.829 | 79.3
Δ9-THC | 4-disc | MQC | 50.2 | 3 | 2 | 52.64 | 104.9
Δ9-THC | 4-disc | MQC | 50.2 | 3 | 3 | 44.77 | 89.2
Δ9-THC | 4-disc | MQC | 50.2 | 3 | 4 | 57.077 | 113.7
Δ9-THC | 4-disc | MQC | 50.2 | 3 | 5 | 56.377 | 112.3
Δ9-THC | 4-disc | MQC | 50.2 | 3 | 6 | 54.145 | 107.9
Δ9-THC | 4-disc | MQC | 50.2 | 4 | 1 | 46.148 | 91.9
Δ9-THC | 4-disc | MQC | 50.2 | 4 | 2 | 62.621 | 124.7
Δ9-THC | 4-disc | MQC | 50.2 | 4 | 3 | 45.382 | 90.4
Δ9-THC | 4-disc | MQC | 50.2 | 4 | 4 | 45.452 | 90.5
Δ9-THC | 4-disc | MQC | 50.2 | 4 | 5 | 45.124 | 89.9
Δ9-THC | 4-disc | MQC | 50.2 | 4 | 6 | 54.689 | 108.9
Δ9-THC | 4-disc | MQC | 50.2 | 5 | 1 | 51.839 | 103.3
Δ9-THC | 4-disc | MQC | 50.2 | 5 | 2 | 50.011 | 99.6
Δ9-THC | 4-disc | MQC | 50.2 | 5 | 3 | 59.441 | 118.4
Δ9-THC | 4-disc | MQC | 50.2 | 5 | 4 | 57.994 | 115.5
Δ9-THC | 4-disc | MQC | 50.2 | 5 | 5 | 52.87 | 105.3
Δ9-THC | 4-disc | MQC | 50.2 | 5 | 6 | 45.777 | 91.2
Δ9-THC | 4-disc | HQC | 75 | 1 | 1 | 64.305 | 85.7
Δ9-THC | 4-disc | HQC | 75 | 1 | 2 | 61.05 | 81.4
Δ9-THC | 4-disc | HQC | 75 | 1 | 3 | 73.483 | 98
Δ9-THC | 4-disc | HQC | 75 | 1 | 4 | 64.754 | 86.3
Δ9-THC | 4-disc | HQC | 75 | 1 | 5 | 69.523 | 92.7
Δ9-THC | 4-disc | HQC | 75 | 1 | 6 | 59.406 | 79.2
Δ9-THC | 4-disc | HQC | 75 | 2 | 1 | 80.286 | 107
Δ9-THC | 4-disc | HQC | 75 | 2 | 2 | 64.479 | 86
Δ9-THC | 4-disc | HQC | 75 | 2 | 3 | 61.137 | 81.5
Δ9-THC | 4-disc | HQC | 75 | 2 | 4 | 68.892 | 91.9
Δ9-THC | 4-disc | HQC | 75 | 2 | 5 | 68.502 | 91.3
Δ9-THC | 4-disc | HQC | 75 | 2 | 6 | 68.699 | 91.6
Δ9-THC | 4-disc | HQC | 75 | 3 | 1 | 68.483 | 91.3
Δ9-THC | 4-disc | HQC | 75 | 3 | 2 | 65.155 | 86.9
Δ9-THC | 4-disc | HQC | 75 | 3 | 3 | 61.572 | 82.1
Δ9-THC | 4-disc | HQC | 75 | 3 | 4 | 74.392 | 99.2
Δ9-THC | 4-disc | HQC | 75 | 3 | 5 | 72.091 | 96.1
Δ9-THC | 4-disc | HQC | 75 | 3 | 6 | 69.086 | 92.1
Δ9-THC | 4-disc | HQC | 75 | 4 | 1 | 76.081 | 101.4
Δ9-THC | 4-disc | HQC | 75 | 4 | 2 | 63.328 | 84.4
Δ9-THC | 4-disc | HQC | 75 | 4 | 3 | 76.73 | 102.3
Δ9-THC | 4-disc | HQC | 75 | 4 | 4 | 74.845 | 99.8
Δ9-THC | 4-disc | HQC | 75 | 4 | 5 | 66.307 | 88.4
Δ9-THC | 4-disc | HQC | 75 | 4 | 6 | 65.183 | 86.9
Δ9-THC | 4-disc | HQC | 75 | 5 | 1 | 72.359 | 96.5
Δ9-THC | 4-disc | HQC | 75 | 5 | 2 | 68.33 | 91.1
Δ9-THC | 4-disc | HQC | 75 | 5 | 3 | 67.954 | 90.6
Δ9-THC | 4-disc | HQC | 75 | 5 | 4 | 69.217 | 92.3
Δ9-THC | 4-disc | HQC | 75 | 5 | 5 | 69.473 | 92.6
Δ9-THC | 4-disc | HQC | 75 | 5 | 6 | 75.897 | 101.2
11-OH-THC | 4-disc | LLOQ | 0.5 | 1 | 1 | 0.549 | 109.8
11-OH-THC | 4-disc | LLOQ | 0.5 | 1 | 2 | 0.579 | 115.8
11-OH-THC | 4-disc | LLOQ | 0.5 | 1 | 3 | 0.635 | 127
11-OH-THC | 4-disc | LLOQ | 0.5 | 1 | 4 | 0.558 | 111.6
11-OH-THC | 4-disc | LLOQ | 0.5 | 1 | 5 | 0.531 | 106.2
11-OH-THC | 4-disc | LLOQ | 0.5 | 1 | 6 | 0.75 | 150
11-OH-THC | 4-disc | LLOQ | 0.5 | 2 | 1 | 0.673 | 134.6
11-OH-THC | 4-disc | LLOQ | 0.5 | 2 | 2 | 0.537 | 107.4
11-OH-THC | 4-disc | LLOQ | 0.5 | 2 | 3 | 0.542 | 108.4
11-OH-THC | 4-disc | LLOQ | 0.5 | 2 | 4 | 0.736 | 147.2
11-OH-THC | 4-disc | LLOQ | 0.5 | 2 | 5 | 0.539 | 107.8
11-OH-THC | 4-disc | LLOQ | 0.5 | 2 | 6 | 0.275 | 55
11-OH-THC | 4-disc | LLOQ | 0.5 | 3 | 1 | 0.601 | 120.2
11-OH-THC | 4-disc | LLOQ | 0.5 | 3 | 2 | 0.586 | 117.2
11-OH-THC | 4-disc | LLOQ | 0.5 | 3 | 3 | 0.586 | 117.2
11-OH-THC | 4-disc | LLOQ | 0.5 | 3 | 4 | 0.657 | 131.4
11-OH-THC | 4-disc | LLOQ | 0.5 | 3 | 5 | 0.419 | 83.8
11-OH-THC | 4-disc | LLOQ | 0.5 | 3 | 6 | 0.563 | 112.6
11-OH-THC | 4-disc | LLOQ | 0.5 | 4 | 1 | 0.593 | 118.6
11-OH-THC | 4-disc | LLOQ | 0.5 | 4 | 2 | 0.663 | 132.6
11-OH-THC | 4-disc | LLOQ | 0.5 | 4 | 3 | 0.444 | 88.8
11-OH-THC | 4-disc | LLOQ | 0.5 | 4 | 4 | 0.522 | 104.4
11-OH-THC | 4-disc | LLOQ | 0.5 | 4 | 5 | 0.545 | 109
11-OH-THC | 4-disc | LLOQ | 0.5 | 4 | 6 | 0.58 | 116
11-OH-THC | 4-disc | LLOQ | 0.5 | 5 | 1 | 0.535 | 107
11-OH-THC | 4-disc | LLOQ | 0.5 | 5 | 2 | 0.642 | 128.4
11-OH-THC | 4-disc | LLOQ | 0.5 | 5 | 3 | 0.632 | 126.4
11-OH-THC | 4-disc | LLOQ | 0.5 | 5 | 4 | 0.648 | 129.6
11-OH-THC | 4-disc | LLOQ | 0.5 | 5 | 5 | 0.656 | 131.2
11-OH-THC | 4-disc | LLOQ | 0.5 | 5 | 6 | 0.443 | 88.6
11-OH-THC | 4-disc | LQC | 1.5 | 1 | 1 | 1.65 | 110
11-OH-THC | 4-disc | LQC | 1.5 | 1 | 2 | 1.529 | 101.9
11-OH-THC | 4-disc | LQC | 1.5 | 1 | 3 | 1.529 | 101.9
11-OH-THC | 4-disc | LQC | 1.5 | 1 | 4 | 1.134 | 75.6
11-OH-THC | 4-disc | LQC | 1.5 | 1 | 5 | 1.65 | 110
11-OH-THC | 4-disc | LQC | 1.5 | 1 | 6 | 1.655 | 110.3
11-OH-THC | 4-disc | LQC | 1.5 | 2 | 1 | 1.571 | 104.7
11-OH-THC | 4-disc | LQC | 1.5 | 2 | 2 | 1.847 | 123.1
11-OH-THC | 4-disc | LQC | 1.5 | 2 | 3 | 1.575 | 105
11-OH-THC | 4-disc | LQC | 1.5 | 2 | 4 | 1.259 | 83.9
11-OH-THC | 4-disc | LQC | 1.5 | 2 | 5 | 1.459 | 97.3
11-OH-THC | 4-disc | LQC | 1.5 | 2 | 6 | 1.489 | 99.3
11-OH-THC | 4-disc | LQC | 1.5 | 3 | 1 | 1.411 | 94.1
11-OH-THC | 4-disc | LQC | 1.5 | 3 | 2 | 1.595 | 106.3
11-OH-THC | 4-disc | LQC | 1.5 | 3 | 3 | 1.565 | 104.3
11-OH-THC | 4-disc | LQC | 1.5 | 3 | 4 | 1.431 | 95.4
11-OH-THC | 4-disc | LQC | 1.5 | 3 | 5 | 1.579 | 105.3
11-OH-THC | 4-disc | LQC | 1.5 | 3 | 6 | 1.945 | 129.7
11-OH-THC | 4-disc | LQC | 1.5 | 4 | 1 | 1.271 | 84.7
11-OH-THC | 4-disc | LQC | 1.5 | 4 | 2 | 1.439 | 95.9
11-OH-THC | 4-disc | LQC | 1.5 | 4 | 3 | 1.729 | 115.3
11-OH-THC | 4-disc | LQC | 1.5 | 4 | 4 | 1.607 | 107.1
11-OH-THC | 4-disc | LQC | 1.5 | 4 | 5 | 1.211 | 80.7
11-OH-THC | 4-disc | LQC | 1.5 | 4 | 6 | 1.388 | 92.5
11-OH-THC | 4-disc | LQC | 1.5 | 5 | 1 | 1.927 | 128.5
11-OH-THC | 4-disc | LQC | 1.5 | 5 | 2 | 1.684 | 112.3
11-OH-THC | 4-disc | LQC | 1.5 | 5 | 3 | 1.587 | 105.8
11-OH-THC | 4-disc | LQC | 1.5 | 5 | 4 | 1.326 | 88.4
11-OH-THC | 4-disc | LQC | 1.5 | 5 | 5 | 1.624 | 108.3
11-OH-THC | 4-disc | LQC | 1.5 | 5 | 6 | 1.617 | 107.8
11-OH-THC | 4-disc | MQC | 50.2 | 1 | 1 | 59.642 | 118.8
11-OH-THC | 4-disc | MQC | 50.2 | 1 | 2 | 46.014 | 91.7
11-OH-THC | 4-disc | MQC | 50.2 | 1 | 3 | 48.566 | 96.7
11-OH-THC | 4-disc | MQC | 50.2 | 1 | 4 | 57.581 | 114.7
11-OH-THC | 4-disc | MQC | 50.2 | 1 | 5 | 51.713 | 103
11-OH-THC | 4-disc | MQC | 50.2 | 1 | 6 | 53.601 | 106.8
11-OH-THC | 4-disc | MQC | 50.2 | 2 | 1 | 52.318 | 104.2
11-OH-THC | 4-disc | MQC | 50.2 | 2 | 2 | 54.652 | 108.9
11-OH-THC | 4-disc | MQC | 50.2 | 2 | 3 | 49.101 | 97.8
11-OH-THC | 4-disc | MQC | 50.2 | 2 | 4 | 36.521 | 72.8
11-OH-THC | 4-disc | MQC | 50.2 | 2 | 5 | 49.503 | 98.6
11-OH-THC | 4-disc | MQC | 50.2 | 2 | 6 | 44.926 | 89.5
11-OH-THC | 4-disc | MQC | 50.2 | 3 | 1 | 54.015 | 107.6
11-OH-THC | 4-disc | MQC | 50.2 | 3 | 2 | 52.178 | 103.9
11-OH-THC | 4-disc | MQC | 50.2 | 3 | 3 | 51.007 | 101.6
11-OH-THC | 4-disc | MQC | 50.2 | 3 | 4 | 47.297 | 94.2
11-OH-THC | 4-disc | MQC | 50.2 | 3 | 5 | 54.806 | 109.2
11-OH-THC | 4-disc | MQC | 50.2 | 3 | 6 | 42.725 | 85.1
11-OH-THC | 4-disc | MQC | 50.2 | 4 | 1 | 56.194 | 111.9
11-OH-THC | 4-disc | MQC | 50.2 | 4 | 2 | 47.344 | 94.3
11-OH-THC | 4-disc | MQC | 50.2 | 4 | 3 | 50.169 | 99.9
11-OH-THC | 4-disc | MQC | 50.2 | 4 | 4 | 53.749 | 107.1
11-OH-THC | 4-disc | MQC | 50.2 | 4 | 5 | 40.974 | 81.6
11-OH-THC | 4-disc | MQC | 50.2 | 4 | 6 | 62.198 | 123.9
11-OH-THC | 4-disc | MQC | 50.2 | 5 | 1 | 44.707 | 89.1
11-OH-THC | 4-disc | MQC | 50.2 | 5 | 2 | 54.216 | 108
11-OH-THC | 4-disc | MQC | 50.2 | 5 | 3 | 63.063 | 125.6
11-OH-THC | 4-disc | MQC | 50.2 | 5 | 4 | 56.081 | 111.7
11-OH-THC | 4-disc | MQC | 50.2 | 5 | 5 | 64.481 | 128.4
11-OH-THC | 4-disc | MQC | 50.2 | 5 | 6 | 49.58 | 98.8
11-OH-THC | 4-disc | HQC | 75 | 1 | 1 | 79.395 | 105.9
11-OH-THC | 4-disc | HQC | 75 | 1 | 2 | 59.753 | 79.7
11-OH-THC | 4-disc | HQC | 75 | 1 | 3 | 65.203 | 86.9
11-OH-THC | 4-disc | HQC | 75 | 1 | 4 | 70.821 | 94.4
11-OH-THC | 4-disc | HQC | 75 | 1 | 5 | 66.774 | 89
11-OH-THC | 4-disc | HQC | 75 | 1 | 6 | 74.303 | 99.1
11-OH-THC | 4-disc | HQC | 75 | 2 | 1 | 62.455 | 83.3
11-OH-THC | 4-disc | HQC | 75 | 2 | 2 | 73.36 | 97.8
11-OH-THC | 4-disc | HQC | 75 | 2 | 3 | 73.725 | 98.3
11-OH-THC | 4-disc | HQC | 75 | 2 | 4 | 73.691 | 98.3
11-OH-THC | 4-disc | HQC | 75 | 2 | 5 | 60.815 | 81.1
11-OH-THC | 4-disc | HQC | 75 | 2 | 6 | 74.506 | 99.3
11-OH-THC | 4-disc | HQC | 75 | 3 | 1 | 70.168 | 93.6
11-OH-THC | 4-disc | HQC | 75 | 3 | 2 | 73.507 | 98
11-OH-THC | 4-disc | HQC | 75 | 3 | 3 | 69.551 | 92.7
11-OH-THC | 4-disc | HQC | 75 | 3 | 4 | 74.18 | 98.9
11-OH-THC | 4-disc | HQC | 75 | 3 | 5 | 67.021 | 89.4
11-OH-THC | 4-disc | HQC | 75 | 3 | 6 | 68.975 | 92
11-OH-THC | 4-disc | HQC | 75 | 4 | 1 | 65.224 | 87
11-OH-THC | 4-disc | HQC | 75 | 4 | 2 | 62.033 | 82.7
11-OH-THC | 4-disc | HQC | 75 | 4 | 3 | 65.613 | 87.5
11-OH-THC | 4-disc | HQC | 75 | 4 | 4 | 52.445 | 69.9
11-OH-THC | 4-disc | HQC | 75 | 4 | 5 | 71.217 | 95
11-OH-THC | 4-disc | HQC | 75 | 4 | 6 | 65.693 | 87.6
11-OH-THC | 4-disc | HQC | 75 | 5 | 1 | 64.699 | 86.3
11-OH-THC | 4-disc | HQC | 75 | 5 | 2 | 71.888 | 95.9
11-OH-THC | 4-disc | HQC | 75 | 5 | 3 | 69.713 | 93
11-OH-THC | 4-disc | HQC | 75 | 5 | 4 | 67.659 | 90.2
11-OH-THC | 4-disc | HQC | 75 | 5 | 5 | 63.515 | 84.7
11-OH-THC | 4-disc | HQC | 75 | 5 | 6 | 67.348 | 89.8
THC-COOH | 1-disc | LLOQ | 2 | 1 | 1 | 1.953 | 97.7
THC-COOH | 1-disc | LLOQ | 2 | 1 | 2 | 1.866 | 93.3
THC-COOH | 1-disc | LLOQ | 2 | 1 | 3 | 1.633 | 81.7
THC-COOH | 1-disc | LLOQ | 2 | 1 | 4 | 2.336 | 116.8
THC-COOH | 1-disc | LLOQ | 2 | 1 | 5 | 1.921 | 96
THC-COOH | 1-disc | LLOQ | 2 | 1 | 6 | 2.304 | 115.2
THC-COOH | 1-disc | LLOQ | 2 | 2 | 1 | 2.16 | 108
THC-COOH | 1-disc | LLOQ | 2 | 2 | 2 | 2.697 | 134.8
THC-COOH | 1-disc | LLOQ | 2 | 2 | 3 | 2.33 | 116.5
THC-COOH | 1-disc | LLOQ | 2 | 2 | 4 | 1.929 | 96.5
THC-COOH | 1-disc | LLOQ | 2 | 2 | 5 | 2.188 | 109.4
THC-COOH | 1-disc | LLOQ | 2 | 2 | 6 | 2.391 | 119.5
THC-COOH | 1-disc | LLOQ | 2 | 3 | 1 | 3.156 | 157.8
THC-COOH | 1-disc | LLOQ | 2 | 3 | 2 | 2.074 | 103.7
THC-COOH | 1-disc | LLOQ | 2 | 3 | 3 | 2.511 | 125.6
THC-COOH | 1-disc | LLOQ | 2 | 3 | 4 | 2.299 | 115
THC-COOH | 1-disc | LLOQ | 2 | 3 | 5 | 2.183 | 109.1
THC-COOH | 1-disc | LLOQ | 2 | 3 | 6 | 2.536 | 126.8
THC-COOH | 1-disc | LLOQ | 2 | 4 | 1 | 2.951 | 147.6
THC-COOH | 1-disc | LLOQ | 2 | 4 | 2 | 2.611 | 130.6
THC-COOH | 1-disc | LLOQ | 2 | 4 | 3 | 2.299 | 115
THC-COOH | 1-disc | LLOQ | 2 | 4 | 4 | 1.907 | 95.3
THC-COOH | 1-disc | LLOQ | 2 | 4 | 5 | 1.615 | 80.8
THC-COOH | 1-disc | LLOQ | 2 | 4 | 6 | 2.369 | 118.5
THC-COOH | 1-disc | LLOQ | 2 | 5 | 1 | 2.527 | 126.4
THC-COOH | 1-disc | LLOQ | 2 | 5 | 2 | 2.813 | 140.7
THC-COOH | 1-disc | LLOQ | 2 | 5 | 3 | 2.284 | 114.2
THC-COOH | 1-disc | LLOQ | 2 | 5 | 4 | 1.85 | 92.5
THC-COOH | 1-disc | LLOQ | 2 | 5 | 5 | 2.666 | 133.3
THC-COOH | 1-disc | LLOQ | 2 | 5 | 6 | 2.102 | 105.1
THC-COOH | 1-disc | LQC | 6 | 1 | 1 | 5.895 | 98.2
THC-COOH | 1-disc | LQC | 6 | 1 | 2 | 7.845 | 130.8
THC-COOH | 1-disc | LQC | 6 | 1 | 3 | 7.547 | 125.8
THC-COOH | 1-disc | LQC | 6 | 1 | 4 | 5.484 | 91.4
THC-COOH | 1-disc | LQC | 6 | 1 | 5 | 5.508 | 91.8
THC-COOH | 1-disc | LQC | 6 | 1 | 6 | 5.104 | 85.1
THC-COOH | 1-disc | LQC | 6 | 2 | 1 | 5.592 | 93.2
THC-COOH | 1-disc | LQC | 6 | 2 | 2 | 6.658 | 111
THC-COOH | 1-disc | LQC | 6 | 2 | 3 | 5.584 | 93.1
THC-COOH | 1-disc | LQC | 6 | 2 | 4 | 7.005 | 116.8
THC-COOH | 1-disc | LQC | 6 | 2 | 5 | 5.599 | 93.3
THC-COOH | 1-disc | LQC | 6 | 2 | 6 | 6.329 | 105.5
THC-COOH | 1-disc | LQC | 6 | 3 | 1 | 6.306 | 105.1
THC-COOH | 1-disc | LQC | 6 | 3 | 2 | 6.46 | 107.7
THC-COOH | 1-disc | LQC | 6 | 3 | 3 | 5.97 | 99.5
THC-COOH | 1-disc | LQC | 6 | 3 | 4 | 4.749 | 79.1
THC-COOH | 1-disc | LQC | 6 | 3 | 5 | 6.329 | 105.5
THC-COOH | 1-disc | LQC | 6 | 3 | 6 | 5.837 | 97.3
THC-COOH | 1-disc | LQC | 6 | 4 | 1 | 6.371 | 106.2
THC-COOH | 1-disc | LQC | 6 | 4 | 2 | 5.59 | 93.2
THC-COOH | 1-disc | LQC | 6 | 4 | 3 | 5.42 | 90.3
THC-COOH | 1-disc | LQC | 6 | 4 | 4 | 6.624 | 110.4
THC-COOH | 1-disc | LQC | 6 | 4 | 5 | 6.412 | 106.9
THC-COOH | 1-disc | LQC | 6 | 4 | 6 | 6.541 | 109
THC-COOH | 1-disc | LQC | 6 | 5 | 1 | 6.424 | 107.1
THC-COOH | 1-disc | LQC | 6 | 5 | 2 | 6.869 | 114.5
THC-COOH | 1-disc | LQC | 6 | 5 | 3 | 6.571 | 109.5
THC-COOH | 1-disc | LQC | 6 | 5 | 4 | 4.763 | 79.4
THC-COOH | 1-disc | LQC | 6 | 5 | 5 | 6.936 | 115.6
THC-COOH | 1-disc | LQC | 6 | 5 | 6 | 5.908 | 98.5
THC-COOH | 1-disc | MQC | 101 | 1 | 1 | 99.31 | 98.3
THC-COOH | 1-disc | MQC | 101 | 1 | 2 | 118.552 | 117.4
THC-COOH | 1-disc | MQC | 101 | 1 | 3 | 113.322 | 112.2
THC-COOH | 1-disc | MQC | 101 | 1 | 4 | 105.033 | 104
THC-COOH | 1-disc | MQC | 101 | 1 | 5 | 120.793 | 119.6
THC-COOH | 1-disc | MQC | 101 | 1 | 6 | 94.192 | 93.3
THC-COOH | 1-disc | MQC | 101 | 2 | 1 | 124.371 | 123.1
THC-COOH | 1-disc | MQC | 101 | 2 | 2 | 88.442 | 87.6
THC-COOH | 1-disc | MQC | 101 | 2 | 3 | 102.948 | 101.9
THC-COOH | 1-disc | MQC | 101 | 2 | 4 | 117.254 | 116.1
THC-COOH | 1-disc | MQC | 101 | 2 | 5 | 117.893 | 116.7
THC-COOH | 1-disc | MQC | 101 | 2 | 6 | 93.488 | 92.6
THC-COOH | 1-disc | MQC | 101 | 3 | 1 | 98.075 | 97.1
THC-COOH | 1-disc | MQC | 101 | 3 | 2 | 97.529 | 96.6
THC-COOH | 1-disc | MQC | 101 | 3 | 3 | 112.756 | 111.6
THC-COOH | 1-disc | MQC | 101 | 3 | 4 | 119.279 | 118.1
THC-COOH | 1-disc | MQC | 101 | 3 | 5 | 100.4 | 99.4
THC-COOH | 1-disc | MQC | 101 | 3 | 6 | 107.495 | 106.4
THC-COOH | 1-disc | MQC | 101 | 4 | 1 | 90.447 | 89.6
THC-COOH | 1-disc | MQC | 101 | 4 | 2 | 108.926 | 107.8
THC-COOH | 1-disc | MQC | 101 | 4 | 3 | 110.997 | 109.9
THC-COOH | 1-disc | MQC | 101 | 4 | 4 | 86.941 | 86.1
THC-COOH | 1-disc | MQC | 101 | 4 | 5 | 115.072 | 113.9
THC-COOH | 1-disc | MQC | 101 | 4 | 6 | 105.24 | 104.2
THC-COOH | 1-disc | MQC | 101 | 5 | 1 | 92.56 | 91.6
THC-COOH | 1-disc | MQC | 101 | 5 | 2 | 96.031 | 95.1
THC-COOH | 1-disc | MQC | 101 | 5 | 3 | 111.463 | 110.4
THC-COOH | 1-disc | MQC | 101 | 5 | 4 | 91.151 | 90.2
THC-COOH | 1-disc | MQC | 101 | 5 | 5 | 75.907 | 75.2
THC-COOH | 1-disc | MQC | 101 | 5 | 6 | 85.336 | 84.5
THC-COOH | 1-disc | HQC | 150 | 1 | 1 | 130.622 | 87.1
THC-COOH | 1-disc | HQC | 150 | 1 | 2 | 144.501 | 96.3
THC-COOH | 1-disc | HQC | 150 | 1 | 3 | 146.493 | 97.7
THC-COOH | 1-disc | HQC | 150 | 1 | 4 | 119.4 | 79.6
THC-COOH | 1-disc | HQC | 150 | 1 | 5 | 152.948 | 102
THC-COOH | 1-disc | HQC | 150 | 1 | 6 | 137.506 | 91.7
THC-COOH | 1-disc | HQC | 150 | 2 | 1 | 147.584 | 98.4
THC-COOH | 1-disc | HQC | 150 | 2 | 2 | 141.523 | 94.3
THC-COOH | 1-disc | HQC | 150 | 2 | 3 | 137.178 | 91.5
THC-COOH | 1-disc | HQC | 150 | 2 | 4 | 122.65 | 81.8
THC-COOH | 1-disc | HQC | 150 | 2 | 5 | 153.522 | 102.3
THC-COOH | 1-disc | HQC | 150 | 2 | 6 | 121.419 | 80.9
THC-COOH | 1-disc | HQC | 150 | 3 | 1 | 141.341 | 94.2
THC-COOH | 1-disc | HQC | 150 | 3 | 2 | 118.952 | 79.3
THC-COOH | 1-disc | HQC | 150 | 3 | 3 | 126.987 | 84.7
THC-COOH | 1-disc | HQC | 150 | 3 | 4 | 140.078 | 93.4
THC-COOH | 1-disc | HQC | 150 | 3 | 5 | 129.785 | 86.5
THC-COOH | 1-disc | HQC | 150 | 3 | 6 | 143.896 | 95.9
THC-COOH | 1-disc | HQC | 150 | 4 | 1 | 147.039 | 98
THC-COOH | 1-disc | HQC | 150 | 4 | 2 | 118.251 | 78.8
THC-COOH | 1-disc | HQC | 150 | 4 | 3 | 124.118 | 82.7
THC-COOH | 1-disc | HQC | 150 | 4 | 4 | 137.265 | 91.5
THC-COOH | 1-disc | HQC | 150 | 4 | 5 | 124.145 | 82.8
THC-COOH | 1-disc | HQC | 150 | 4 | 6 | 128.946 | 86
THC-COOH | 1-disc | HQC | 150 | 5 | 1 | 143.998 | 96
THC-COOH | 1-disc | HQC | 150 | 5 | 2 | 139.496 | 93
THC-COOH | 1-disc | HQC | 150 | 5 | 3 | 146.539 | 97.7
THC-COOH | 1-disc | HQC | 150 | 5 | 4 | 138.061 | 92
THC-COOH | 1-disc | HQC | 150 | 5 | 5 | 144.711 | 96.5
THC-COOH | 1-disc | HQC | 150 | 5 | 6 | 128.048 | 85.4
CBD | 1-disc | LLOQ | 2 | 1 | 1 | 1.858 | 92.9
CBD | 1-disc | LLOQ | 2 | 1 | 2 | 2.003 | 100.2
CBD | 1-disc | LLOQ | 2 | 1 | 3 | 1.767 | 88.3
CBD | 1-disc | LLOQ | 2 | 1 | 4 | 2.118 | 105.9
CBD | 1-disc | LLOQ | 2 | 1 | 5 | 2.273 | 113.7
CBD | 1-disc | LLOQ | 2 | 1 | 6 | 2.298 | 114.9
CBD | 1-disc | LLOQ | 2 | 2 | 1 | 2.104 | 105.2
CBD | 1-disc | LLOQ | 2 | 2 | 2 | 2.611 | 130.6
CBD | 1-disc | LLOQ | 2 | 2 | 3 | 1.708 | 85.4
CBD | 1-disc | LLOQ | 2 | 2 | 4 | 1.973 | 98.7
CBD | 1-disc | LLOQ | 2 | 2 | 5 | 2.529 | 126.4
CBD | 1-disc | LLOQ | 2 | 2 | 6 | 2.395 | 119.8
CBD | 1-disc | LLOQ | 2 | 3 | 1 | 2.138 | 106.9
CBD | 1-disc | LLOQ | 2 | 3 | 2 | 2.224 | 111.2
CBD | 1-disc | LLOQ | 2 | 3 | 3 | 1.964 | 98.2
CBD | 1-disc | LLOQ | 2 | 3 | 4 | 1.974 | 98.7
CBD | 1-disc | LLOQ | 2 | 3 | 5 | 2.371 | 118.5
CBD | 1-disc | LLOQ | 2 | 3 | 6 | 2.72 | 136
CBD | 1-disc | LLOQ | 2 | 4 | 1 | 2.065 | 103.2
CBD | 1-disc | LLOQ | 2 | 4 | 2 | 2.223 | 111.1
CBD | 1-disc | LLOQ | 2 | 4 | 3 | 3.001 | 150
CBD | 1-disc | LLOQ | 2 | 4 | 4 | 2.411 | 120.5
CBD | 1-disc | LLOQ | 2 | 4 | 5 | 2.477 | 123.8
CBD | 1-disc | LLOQ | 2 | 4 | 6 | 2.095 | 104.8
CBD | 1-disc | LLOQ | 2 | 5 | 1 | 3.095 | 154.8
CBD | 1-disc | LLOQ | 2 | 5 | 2 | 2.252 | 112.6
CBD | 1-disc | LLOQ | 2 | 5 | 3 | 2.194 | 109.7
CBD | 1-disc | LLOQ | 2 | 5 | 4 | 1.917 | 95.9
CBD | 1-disc | LLOQ | 2 | 5 | 5 | 1.793 | 89.6
CBD | 1-disc | LLOQ | 2 | 5 | 6 | 2.532 | 126.6
CBD | 1-disc | LQC | 6 | 1 | 1 | 4.614 | 76.9
CBD | 1-disc | LQC | 6 | 1 | 2 | 5.496 | 91.6
CBD | 1-disc | LQC | 6 | 1 | 3 | 6.143 | 102.4
CBD | 1-disc | LQC | 6 | 1 | 4 | 6.748 | 112.5
CBD | 1-disc | LQC | 6 | 1 | 5 | 6.651 | 110.9
CBD | 1-disc | LQC | 6 | 1 | 6 | 5.684 | 94.7
CBD | 1-disc | LQC | 6 | 2 | 1 | 6.65 | 110.8
CBD | 1-disc | LQC | 6 | 2 | 2 | 6.295 | 104.9
CBD | 1-disc | LQC | 6 | 2 | 3 | 6.831 | 113.9
CBD | 1-disc | LQC | 6 | 2 | 4 | 5.595 | 93.2
CBD | 1-disc | LQC | 6 | 2 | 5 | 6.392 | 106.5
CBD | 1-disc | LQC | 6 | 2 | 6 | 7.649 | 127.5
CBD | 1-disc | LQC | 6 | 3 | 1 | 5.8 | 96.7
CBD | 1-disc | LQC | 6 | 3 | 2 | 6.577 | 109.6
CBD | 1-disc | LQC | 6 | 3 | 3 | 5.487 | 91.5
CBD | 1-disc | LQC | 6 | 3 | 4 | 6.369 | 106.1
CBD | 1-disc | LQC | 6 | 3 | 5 | 5.693 | 94.9
CBD | 1-disc | LQC | 6 | 3 | 6 | 6.293 | 104.9
CBD | 1-disc | LQC | 6 | 4 | 1 | 6.193 | 103.2
CBD | 1-disc | LQC | 6 | 4 | 2 | 6.018 | 100.3
CBD | 1-disc | LQC | 6 | 4 | 3 | 4.748 | 79.1
CBD | 1-disc | LQC | 6 | 4 | 4 | 6.058 | 101
CBD | 1-disc | LQC | 6 | 4 | 5 | 6.062 | 101
CBD | 1-disc | LQC | 6 | 4 | 6 | 6.672 | 111.2
CBD | 1-disc | LQC | 6 | 5 | 1 | 4.952 | 82.5
CBD | 1-disc | LQC | 6 | 5 | 2 | 6.291 | 104.8
CBD | 1-disc | LQC | 6 | 5 | 3 | 6.05 | 100.8
CBD | 1-disc | LQC | 6 | 5 | 4 | 6.313 | 105.2
CBD | 1-disc | LQC | 6 | 5 | 5 | 6.014 | 100.2
CBD | 1-disc | LQC | 6 | 5 | 6 | 7.439 | 124
CBD | 1-disc | MQC | 101 | 1 | 1 | 101.707 | 100.7
CBD | 1-disc | MQC | 101 | 1 | 2 | 97.092 | 96.1
CBD | 1-disc | MQC | 101 | 1 | 3 | 98.74 | 97.8
CBD | 1-disc | MQC | 101 | 1 | 4 | 112.596 | 111.5
CBD | 1-disc | MQC | 101 | 1 | 5 | 101.306 | 100.3
CBD | 1-disc | MQC | 101 | 1 | 6 | 115.364 | 114.2
CBD | 1-disc | MQC | 101 | 2 | 1 | 86.924 | 86.1
CBD | 1-disc | MQC | 101 | 2 | 2 | 107.305 | 106.2
CBD | 1-disc | MQC | 101 | 2 | 3 | 128.837 | 127.6
CBD | 1-disc | MQC | 101 | 2 | 4 | 119.547 | 118.4
CBD | 1-disc | MQC | 101 | 2 | 5 | 103.935 | 102.9
CBD | 1-disc | MQC | 101 | 2 | 6 | 126.374 | 125.1
CBD | 1-disc | MQC | 101 | 3 | 1 | 119.049 | 117.9
CBD | 1-disc | MQC | 101 | 3 | 2 | 89.399 | 88.5
CBD | 1-disc | MQC | 101 | 3 | 3 | 98.483 | 97.5
CBD | 1-disc | MQC | 101 | 3 | 4 | 97.982 | 97
CBD | 1-disc | MQC | 101 | 3 | 5 | 83.931 | 83.1
CBD | 1-disc | MQC | 101 | 3 | 6 | 99.416 | 98.4
CBD | 1-disc | MQC | 101 | 4 | 1 | 89.963 | 89.1
CBD | 1-disc | MQC | 101 | 4 | 2 | 105.663 | 104.6
CBD | 1-disc | MQC | 101 | 4 | 3 | 93.232 | 92.3
CBD | 1-disc | MQC | 101 | 4 | 4 | 110.246 | 109.2
CBD | 1-disc | MQC | 101 | 4 | 5 | 116.964 | 115.8
CBD | 1-disc | MQC | 101 | 4 | 6 | 102.489 | 101.5
CBD | 1-disc | MQC | 101 | 5 | 1 | 105.076 | 104
CBD | 1-disc | MQC | 101 | 5 | 2 | 99.113 | 98.1
CBD | 1-disc | MQC | 101 | 5 | 3 | 94.671 | 93.7
CBD | 1-disc | MQC | 101 | 5 | 4 | 103.376 | 102.4
CBD | 1-disc | MQC | 101 | 5 | 5 | 88.264 | 87.4
CBD | 1-disc | MQC | 101 | 5 | 6 | 96.589 | 95.6
CBD | 1-disc | HQC | 150 | 1 | 1 | 129.74 | 86.5
CBD | 1-disc | HQC | 150 | 1 | 2 | 154.349 | 102.9
CBD | 1-disc | HQC | 150 | 1 | 3 | 129.948 | 86.6
CBD | 1-disc | HQC | 150 | 1 | 4 | 147.815 | 98.5
CBD | 1-disc | HQC | 150 | 1 | 5 | 128.67 | 85.8
CBD | 1-disc | HQC | 150 | 1 | 6 | 143.192 | 95.5
CBD | 1-disc | HQC | 150 | 2 | 1 | 150.986 | 100.7
CBD | 1-disc | HQC | 150 | 2 | 2 | 150.541 | 100.4
CBD | 1-disc | HQC | 150 | 2 | 3 | 131.511 | 87.7
CBD | 1-disc | HQC | 150 | 2 | 4 | 142.319 | 94.9
CBD | 1-disc | HQC | 150 | 2 | 5 | 137.524 | 91.7
CBD | 1-disc | HQC | 150 | 2 | 6 | 127.283 | 84.9
CBD | 1-disc | HQC | 150 | 3 | 1 | 146.684 | 97.8
CBD | 1-disc | HQC | 150 | 3 | 2 | 127.913 | 85.3
CBD | 1-disc | HQC | 150 | 3 | 3 | 121.831 | 81.2
CBD | 1-disc | HQC | 150 | 3 | 4 | 123.884 | 82.6
CBD | 1-disc | HQC | 150 | 3 | 5 | 135.876 | 90.6
CBD | 1-disc | HQC | 150 | 3 | 6 | 147.851 | 98.6
CBD | 1-disc | HQC | 150 | 4 | 1 | 120.069 | 80
CBD | 1-disc | HQC | 150 | 4 | 2 | 134.646 | 89.8
CBD | 1-disc | HQC | 150 | 4 | 3 | 133.301 | 88.9
CBD | 1-disc | HQC | 150 | 4 | 4 | 142.833 | 95.2
CBD | 1-disc | HQC | 150 | 4 | 5 | 144.076 | 96.1
CBD | 1-disc | HQC | 150 | 4 | 6 | 141.711 | 94.5
CBD | 1-disc | HQC | 150 | 5 | 1 | 143.939 | 96
CBD | 1-disc | HQC | 150 | 5 | 2 | 148.113 | 98.7
CBD | 1-disc | HQC | 150 | 5 | 3 | 156.703 | 104.5
CBD | 1-disc | HQC | 150 | 5 | 4 | 139.944 | 93.3
CBD | 1-disc | HQC | 150 | 5 | 5 | 132.559 | 88.4
CBD | 1-disc | HQC | 150 | 5 | 6 | 142.187 | 94.8
7-OH-CBD | 1-disc | LLOQ | 2 | 1 | 1 | 2.403 | 120.2
7-OH-CBD | 1-disc | LLOQ | 2 | 1 | 2 | 2.936 | 146.8
7-OH-CBD | 1-disc | LLOQ | 2 | 1 | 3 | 2.286 | 114.3
7-OH-CBD | 1-disc | LLOQ | 2 | 1 | 4 | 1.529 | 76.4
7-OH-CBD | 1-disc | LLOQ | 2 | 1 | 5 | 1.884 | 94.2
7-OH-CBD | 1-disc | LLOQ | 2 | 1 | 6 | 2.276 | 113.8
7-OH-CBD | 1-disc | LLOQ | 2 | 2 | 1 | 2.343 | 117.2
7-OH-CBD | 1-disc | LLOQ | 2 | 2 | 2 | 2.472 | 123.6
7-OH-CBD | 1-disc | LLOQ | 2 | 2 | 3 | 2.075 | 103.8
7-OH-CBD | 1-disc | LLOQ | 2 | 2 | 4 | 2.378 | 118.9
7-OH-CBD | 1-disc | LLOQ | 2 | 2 | 5 | 1.941 | 97
7-OH-CBD | 1-disc | LLOQ | 2 | 2 | 6 | 2.287 | 114.3
7-OH-CBD | 1-disc | LLOQ | 2 | 3 | 1 | 1.78 | 89
7-OH-CBD | 1-disc | LLOQ | 2 | 3 | 2 | 2.553 | 127.6
7-OH-CBD | 1-disc | LLOQ | 2 | 3 | 3 | 2.561 | 128.1
7-OH-CBD | 1-disc | LLOQ | 2 | 3 | 4 | 2.569 | 128.4
7-OH-CBD | 1-disc | LLOQ | 2 | 3 | 5 | 1.5 | 75
7-OH-CBD | 1-disc | LLOQ | 2 | 3 | 6 | 2.496 | 124.8
7-OH-CBD | 1-disc | LLOQ | 2 | 4 | 1 | 2.246 | 112.3
7-OH-CBD | 1-disc | LLOQ | 2 | 4 | 2 | 1.838 | 91.9
7-OH-CBD | 1-disc | LLOQ | 2 | 4 | 3 | 2.741 | 137.1
7-OH-CBD | 1-disc | LLOQ | 2 | 4 | 4 | 1.816 | 90.8
7-OH-CBD | 1-disc | LLOQ | 2 | 4 | 5 | 2.193 | 109.7
7-OH-CBD | 1-disc | LLOQ | 2 | 4 | 6 | 2.008 | 100.4
7-OH-CBD | 1-disc | LLOQ | 2 | 5 | 1 | 2.732 | 136.6
7-OH-CBD | 1-disc | LLOQ | 2 | 5 | 2 | 2.825 | 141.2
7-OH-CBD | 1-disc | LLOQ | 2 | 5 | 3 | 2.64 | 132
7-OH-CBD | 1-disc | LLOQ | 2 | 5 | 4 | 1.872 | 93.6
7-OH-CBD | 1-disc | LLOQ | 2 | 5 | 5 | 2.298 | 114.9
7-OH-CBD | 1-disc | LLOQ | 2 | 5 | 6 | 1.902 | 95.1
7-OH-CBD | 1-disc | LQC | 6 | 1 | 1 | 6.467 | 107.8
7-OH-CBD | 1-disc | LQC | 6 | 1 | 2 | 6.367 | 106.1
7-OH-CBD | 1-disc | LQC | 6 | 1 | 3 | 7.004 | 116.7
7-OH-CBD | 1-disc | LQC | 6 | 1 | 4 | 6.249 | 104.1
7-OH-CBD | 1-disc | LQC | 6 | 1 | 5 | 5.524 | 92.1
7-OH-CBD | 1-disc | LQC | 6 | 1 | 6 | 6.699 | 111.7
7-OH-CBD | 1-disc | LQC | 6 | 2 | 1 | 6.896 | 114.9
7-OH-CBD | 1-disc | LQC | 6 | 2 | 2 | 5.991 | 99.8
7-OH-CBD | 1-disc | LQC | 6 | 2 | 3 | 5.382 | 89.7
7-OH-CBD | 1-disc | LQC | 6 | 2 | 4 | 5.933 | 98.9
7-OH-CBD | 1-disc | LQC | 6 | 2 | 5 | 6.401 | 106.7
7-OH-CBD | 1-disc | LQC | 6 | 2 | 6 | 7.081 | 118
7-OH-CBD | 1-disc | LQC | 6 | 3 | 1 | 5.175 | 86.2
7-OH-CBD | 1-disc | LQC | 6 | 3 | 2 | 6.621 | 110.4
7-OH-CBD | 1-disc | LQC | 6 | 3 | 3 | 6.396 | 106.6
7-OH-CBD | 1-disc | LQC | 6 | 3 | 4 | 6.003 | 100
7-OH-CBD | 1-disc | LQC | 6 | 3 | 5 | 5.299 | 88.3
7-OH-CBD | 1-disc | LQC | 6 | 3 | 6 | 4.656 | 77.6
7-OH-CBD | 1-disc | LQC | 6 | 4 | 1 | 4.92 | 82
7-OH-CBD | 1-disc | LQC | 6 | 4 | 2 | 7.676 | 127.9
7-OH-CBD | 1-disc | LQC | 6 | 4 | 3 | 5.677 | 94.6
7-OH-CBD | 1-disc | LQC | 6 | 4 | 4 | 6.127 | 102.1
7-OH-CBD | 1-disc | LQC | 6 | 4 | 5 | 5.313 | 88.5
7-OH-CBD | 1-disc | LQC | 6 | 4 | 6 | 7.956 | 132.6
7-OH-CBD | 1-disc | LQC | 6 | 5 | 1 | 6.933 | 115.5
7-OH-CBD | 1-disc | LQC | 6 | 5 | 2 | 5.967 | 99.4
7-OH-CBD | 1-disc | LQC | 6 | 5 | 3 | 5.667 | 94.5
7-OH-CBD | 1-disc | LQC | 6 | 5 | 4 | 6.03 | 100.5
7-OH-CBD | 1-disc | LQC | 6 | 5 | 5 | 5.732 | 95.5
7-OH-CBD | 1-disc | LQC | 6 | 5 | 6 | 5.638 | 94
7-OH-CBD | 1-disc | MQC | 101 | 1 | 1 | 86.302 | 85.4
7-OH-CBD | 1-disc | MQC | 101 | 1 | 2 | 85.078 | 84.2
7-OH-CBD | 1-disc | MQC | 101 | 1 | 3 | 128.185 | 126.9
7-OH-CBD | 1-disc | MQC | 101 | 1 | 4 | 106.49 | 105.4
7-OH-CBD | 1-disc | MQC | 101 | 1 | 5 | 108.298 | 107.2
7-OH-CBD | 1-disc | MQC | 101 | 1 | 6 | 90.921 | 90
7-OH-CBD | 1-disc | MQC | 101 | 2 | 1 | 122.789 | 121.6
7-OH-CBD | 1-disc | MQC | 101 | 2 | 2 | 101.262 | 100.3
7-OH-CBD | 1-disc | MQC | 101 | 2 | 3 | 97.285 | 96.3
7-OH-CBD | 1-disc | MQC | 101 | 2 | 4 | 93.371 | 92.4
7-OH-CBD | 1-disc | MQC | 101 | 2 | 5 | 95.231 | 94.3
7-OH-CBD | 1-disc | MQC | 101 | 2 | 6 | 105.297 | 104.3
7-OH-CBD | 1-disc | MQC | 101 | 3 | 1 | 88.42 | 87.5
7-OH-CBD | 1-disc | MQC | 101 | 3 | 2 | 121.814 | 120.6
7-OH-CBD | 1-disc | MQC | 101 | 3 | 3 | 102.483 | 101.5
7-OH-CBD | 1-disc | MQC | 101 | 3 | 4 | 115.15 | 114
7-OH-CBD | 1-disc | MQC | 101 | 3 | 5 | 88.745 | 87.9
7-OH-CBD | 1-disc | MQC | 101 | 3 | 6 | 113.933 | 112.8
7-OH-CBD | 1-disc | MQC | 101 | 4 | 1 | 93.334 | 92.4
7-OH-CBD | 1-disc | MQC | 101 | 4 | 2 | 110.209 | 109.1
7-OH-CBD | 1-disc | MQC | 101 | 4 | 3 | 113.305 | 112.2
7-OH-CBD | 1-disc | MQC | 101 | 4 | 4 | 100.343 | 99.3
7-OH-CBD | 1-disc | MQC | 101 | 4 | 5 | 111.542 | 110.4
7-OH-CBD | 1-disc | MQC | 101 | 4 | 6 | 100.45 | 99.5
7-OH-CBD | 1-disc | MQC | 101 | 5 | 1 | 116.032 | 114.9
7-OH-CBD | 1-disc | MQC | 101 | 5 | 2 | 116.548 | 115.4
7-OH-CBD | 1-disc | MQC | 101 | 5 | 3 | 118.217 | 117
7-OH-CBD | 1-disc | MQC | 101 | 5 | 4 | 75.978 | 75.2
7-OH-CBD | 1-disc | MQC | 101 | 5 | 5 | 95.76 | 94.8
7-OH-CBD | 1-disc | MQC | 101 | 5 | 6 | 90.859 | 90
7-OH-CBD | 1-disc | HQC | 150 | 1 | 1 | 123.492 | 82.3
7-OH-CBD | 1-disc | HQC | 150 | 1 | 2 | 163.573 | 109
7-OH-CBD | 1-disc | HQC | 150 | 1 | 3 | 146.487 | 97.7
7-OH-CBD | 1-disc | HQC | 150 | 1 | 4 | 129.29 | 86.2
7-OH-CBD | 1-disc | HQC | 150 | 1 | 5 | 164.957 | 110
7-OH-CBD | 1-disc | HQC | 150 | 1 | 6 | 140.078 | 93.4
7-OH-CBD | 1-disc | HQC | 150 | 2 | 1 | 134.273 | 89.5
7-OH-CBD | 1-disc | HQC | 150 | 2 | 2 | 137.737 | 91.8
7-OH-CBD | 1-disc | HQC | 150 | 2 | 3 | 131.544 | 87.7
7-OH-CBD | 1-disc | HQC | 150 | 2 | 4 | 127.025 | 84.7
7-OH-CBD | 1-disc | HQC | 150 | 2 | 5 | 141.093 | 94.1
7-OH-CBD | 1-disc | HQC | 150 | 2 | 6 | 133.833 | 89.2
7-OH-CBD | 1-disc | HQC | 150 | 3 | 1 | 153.901 | 102.6
7-OH-CBD | 1-disc | HQC | 150 | 3 | 2 | 124.816 | 83.2
7-OH-CBD | 1-disc | HQC | 150 | 3 | 3 | 142.965 | 95.3
7-OH-CBD | 1-disc | HQC | 150 | 3 | 4 | 144.91 | 96.6
7-OH-CBD | 1-disc | HQC | 150 | 3 | 5 | 128.919 | 85.9
7-OH-CBD | 1-disc | HQC | 150 | 3 | 6 | 123.619 | 82.4
7-OH-CBD | 1-disc | HQC | 150 | 4 | 1 | 145.619 | 97.1
7-OH-CBD | 1-disc | HQC | 150 | 4 | 2 | 132.809 | 88.5
7-OH-CBD | 1-disc | HQC | 150 | 4 | 3 | 126.705 | 84.5
7-OH-CBD | 1-disc | HQC | 150 | 4 | 4 | 137.597 | 91.7
7-OH-CBD | 1-disc | HQC | 150 | 4 | 5 | 135.53 | 90.4
7-OH-CBD | 1-disc | HQC | 150 | 4 | 6 | 141.541 | 94.4
7-OH-CBD | 1-disc | HQC | 150 | 5 | 1 | 143.544 | 95.7
7-OH-CBD | 1-disc | HQC | 150 | 5 | 2 | 151.577 | 101.1
7-OH-CBD | 1-disc | HQC | 150 | 5 | 3 | 135.288 | 90.2
7-OH-CBD | 1-disc | HQC | 150 | 5 | 4 | 116.945 | 78
7-OH-CBD | 1-disc | HQC | 150 | 5 | 5 | 126.576 | 84.4
7-OH-CBD | 1-disc | HQC | 150 | 5 | 6 | 149.259 | 99.5
7-COOH-CBD | 1-disc | LLOQ | 5 | 1 | 1 | 4.879 | 97.6
7-COOH-CBD | 1-disc | LLOQ | 5 | 1 | 2 | 5.778 | 115.6
7-COOH-CBD | 1-disc | LLOQ | 5 | 1 | 3 | 3.616 | 72.3
7-COOH-CBD | 1-disc | LLOQ | 5 | 1 | 4 | 6.913 | 138.3
7-COOH-CBD | 1-disc | LLOQ | 5 | 1 | 5 | 5.45 | 109
7-COOH-CBD | 1-disc | LLOQ | 5 | 1 | 6 | 4.664 | 93.3
7-COOH-CBD | 1-disc | LLOQ | 5 | 2 | 1 | 5.11 | 102.2
7-COOH-CBD | 1-disc | LLOQ | 5 | 2 | 2 | 7.076 | 141.5
7-COOH-CBD | 1-disc | LLOQ | 5 | 2 | 3 | 6.802 | 136
7-COOH-CBD | 1-disc | LLOQ | 5 | 2 | 4 | 5.67 | 113.4
7-COOH-CBD | 1-disc | LLOQ | 5 | 2 | 5 | 5.261 | 105.2
7-COOH-CBD | 1-disc | LLOQ | 5 | 2 | 6 | 6.135 | 122.7
7-COOH-CBD | 1-disc | LLOQ | 5 | 3 | 1 | 6.87 | 137.4
7-COOH-CBD | 1-disc | LLOQ | 5 | 3 | 2 | 6.236 | 124.7
7-COOH-CBD | 1-disc | LLOQ | 5 | 3 | 3 | 7.039 | 140.8
7-COOH-CBD | 1-disc | LLOQ | 5 | 3 | 4 | 4.916 | 98.3
7-COOH-CBD | 1-disc | LLOQ | 5 | 3 | 5 | 8.053 | 161.1
7-COOH-CBD | 1-disc | LLOQ | 5 | 3 | 6 | 4.837 | 96.7
7-COOH-CBD | 1-disc | LLOQ | 5 | 4 | 1 | 4.995 | 99.9
7-COOH-CBD | 1-disc | LLOQ | 5 | 4 | 2 | 5.619 | 112.4
7-COOH-CBD | 1-disc | LLOQ | 5 | 4 | 3 | 5.566 | 111.3
7-COOH-CBD | 1-disc | LLOQ | 5 | 4 | 4 | 4.866 | 97.3
7-COOH-CBD | 1-disc | LLOQ | 5 | 4 | 5 | 7.618 | 152.4
7-COOH-CBD | 1-disc | LLOQ | 5 | 4 | 6 | 5.026 | 100.5
7-COOH-CBD | 1-disc | LLOQ | 5 | 5 | 1 | 5.949 | 119
7-COOH-CBD | 1-disc | LLOQ | 5 | 5 | 2 | 4.9 | 98
7-COOH-CBD | 1-disc | LLOQ | 5 | 5 | 3 | 5.41 | 108.2
7-COOH-CBD | 1-disc | LLOQ | 5 | 5 | 4 | 4.079 | 81.6
7-COOH-CBD | 1-disc | LLOQ | 5 | 5 | 5 | 4.976 | 99.5
7-COOH-CBD | 1-disc | LLOQ | 5 | 5 | 6 | 6.24 | 124.8
7-COOH-CBD | 1-disc | LQC | 15 | 1 | 1 | 13.37 | 89.1
7-COOH-CBD | 1-disc | LQC | 15 | 1 | 2 | 17.083 | 113.9
7-COOH-CBD | 1-disc | LQC | 15 | 1 | 3 | 16.468 | 109.8
7-COOH-CBD | 1-disc | LQC | 15 | 1 | 4 | 15.344 | 102.3
7-COOH-CBD | 1-disc | LQC | 15 | 1 | 5 | 14.437 | 96.2
7-COOH-CBD | 1-disc | LQC | 15 | 1 | 6 | 13.423 | 89.5
7-COOH-CBD | 1-disc | LQC | 15 | 2 | 1 | 15.095 | 100.6
7-COOH-CBD | 1-disc | LQC | 15 | 2 | 2 | 16.258 | 108.4
7-COOH-CBD | 1-disc | LQC | 15 | 2 | 3 | 9.503 | 63.4
7-COOH-CBD | 1-disc | LQC | 15 | 2 | 4 | 15.826 | 105.5
7-COOH-CBD | 1-disc | LQC | 15 | 2 | 5 | 15.095 | 100.6
7-COOH-CBD | 1-disc | LQC | 15 | 2 | 6 | 16.824 | 112.2
7-COOH-CBD | 1-disc | LQC | 15 | 3 | 1 | 15.865 | 105.8
7-COOH-CBD | 1-disc | LQC | 15 | 3 | 2 | 18.721 | 124.8
7-COOH-CBD | 1-disc | LQC | 15 | 3 | 3 | 14.15 | 94.3
7-COOH-CBD | 1-disc | LQC | 15 | 3 | 4 | 11.6 | 77.3
7-COOH-CBD | 1-disc | LQC | 15 | 3 | 5 | 18.448 | 123
7-COOH-CBD | 1-disc | LQC | 15 | 3 | 6 | 15.641 | 104.3
7-COOH-CBD | 1-disc | LQC | 15 | 4 | 1 | 14.97 | 99.8
7-COOH-CBD | 1-disc | LQC | 15 | 4 | 2 | 11.234 | 74.9
7-COOH-CBD | 1-disc | LQC | 15 | 4 | 3 | 16.792 | 111.9
7-COOH-CBD | 1-disc | LQC | 15 | 4 | 4 | 17.472 | 116.5
7-COOH-CBD | 1-disc | LQC | 15 | 4 | 5 | 15.49 | 103.3
7-COOH-CBD | 1-disc | LQC | 15 | 4 | 6 | 12.648 | 84.3
7-COOH-CBD | 1-disc | LQC | 15 | 5 | 1 | 14.794 | 98.6
7-COOH-CBD | 1-disc | LQC | 15 | 5 | 2 | 16.892 | 112.6
7-COOH-CBD | 1-disc | LQC | 15 | 5 | 3 | 15.953 | 106.4
7-COOH-CBD | 1-disc | LQC | 15 | 5 | 4 | 16.687 | 111.2
7-COOH-CBD | 1-disc | LQC | 15 | 5 | 5 | 17.655 | 117.7
7-COOH-CBD | 1-disc | LQC | 15 | 5 | 6 | 15.038 | 100.3
7-COOH-CBD | 1-disc | MQC | 252.5 | 1 | 1 | 248.17 | 98.3
7-COOH-CBD | 1-disc | MQC | 252.5 | 1 | 2 | 304.872 | 120.7
7-COOH-CBD | 1-disc | MQC | 252.5 | 1 | 3 | 240.764 | 95.4
7-COOH-CBD | 1-disc | MQC | 252.5 | 1 | 4 | 289.785 | 114.8
7-COOH-CBD | 1-disc | MQC | 252.5 | 1 | 5 | 296.419 | 117.4
7-COOH-CBD | 1-disc | MQC | 252.5 | 1 | 6 | 233.056 | 92.3
7-COOH-CBD | 1-disc | MQC | 252.5 | 2 | 1 | 238.294 | 94.4
7-COOH-CBD | 1-disc | MQC | 252.5 | 2 | 2 | 232.097 | 91.9
7-COOH-CBD | 1-disc | MQC | 252.5 | 2 | 3 | 238.717 | 94.5
7-COOH-CBD | 1-disc | MQC | 252.5 | 2 | 4 | 304.374 | 120.5
7-COOH-CBD | 1-disc | MQC | 252.5 | 2 | 5 | 240.079 | 95.1
7-COOH-CBD | 1-disc | MQC | 252.5 | 2 | 6 | 288.345 | 114.2
7-COOH-CBD | 1-disc | MQC | 252.5 | 3 | 1 | 242.281 | 96
7-COOH-CBD | 1-disc | MQC | 252.5 | 3 | 2 | 298.093 | 118.1
7-COOH-CBD | 1-disc | MQC | 252.5 | 3 | 3 | 256.287 | 101.5
7-COOH-CBD | 1-disc | MQC | 252.5 | 3 | 4 | 250.695 | 99.3
7-COOH-CBD | 1-disc | MQC | 252.5 | 3 | 5 | 279.576 | 110.7
7-COOH-CBD | 1-disc | MQC | 252.5 | 3 | 6 | 297.274 | 117.7
7-COOH-CBD | 1-disc | MQC | 252.5 | 4 | 1 | 247.496 | 98
7-COOH-CBD | 1-disc | MQC | 252.5 | 4 | 2 | 237.037 | 93.9
7-COOH-CBD | 1-disc | MQC | 252.5 | 4 | 3 | 279.217 | 110.6
7-COOH-CBD | 1-disc | MQC | 252.5 | 4 | 4 | 233.017 | 92.3
7-COOH-CBD | 1-disc | MQC | 252.5 | 4 | 5 | 274.998 | 108.9
7-COOH-CBD | 1-disc | MQC | 252.5 | 4 | 6 | 267.665 | 106
7-COOH-CBD | 1-disc | MQC | 252.5 | 5 | 1 | 339.706 | 134.5
7-COOH-CBD | 1-disc | MQC | 252.5 | 5 | 2 | 234.636 | 92.9
7-COOH-CBD | 1-disc | MQC | 252.5 | 5 | 3 | 215.687 | 85.4
7-COOH-CBD | 1-disc | MQC | 252.5 | 5 | 4 | 168.624 | 66.8
7-COOH-CBD | 1-disc | MQC | 252.5 | 5 | 5 | 236.966 | 93.8
7-COOH-CBD | 1-disc | MQC | 252.5 | 5 | 6 | 208.482 | 82.6
7-COOH-CBD | 1-disc | HQC | 375 | 1 | 1 | 367.705 | 98.1
7-COOH-CBD | 1-disc | HQC | 375 | 1 | 2 | 317.184 | 84.6
7-COOH-CBD | 1-disc | HQC | 375 | 1 | 3 | 320.842 | 85.6
7-COOH-CBD | 1-disc | HQC | 375 | 1 | 4 | 357.182 | 95.2
7-COOH-CBD | 1-disc | HQC | 375 | 1 | 5 | 303.007 | 80.8
7-COOH-CBD | 1-disc | HQC | 375 | 1 | 6 | 372.011 | 99.2
7-COOH-CBD | 1-disc | HQC | 375 | 2 | 1 | 342.974 | 91.5
7-COOH-CBD | 1-disc | HQC | 375 | 2 | 2 | 371.989 | 99.2
7-COOH-CBD | 1-disc | HQC | 375 | 2 | 3 | 359.71 | 95.9
7-COOH-CBD | 1-disc | HQC | 375 | 2 | 4 | 318.81 | 85
7-COOH-CBD | 1-disc | HQC | 375 | 2 | 5 | 387.024 | 103.2
7-COOH-CBD | 1-disc | HQC | 375 | 2 | 6 | 371.739 | 99.1
7-COOH-CBD | 1-disc | HQC | 375 | 3 | 1 | 322.875 | 86.1
7-COOH-CBD | 1-disc | HQC | 375 | 3 | 2 | 353.81 | 94.3
7-COOH-CBD | 1-disc | HQC | 375 | 3 | 3 | 360.379 | 96.1
7-COOH-CBD | 1-disc | HQC | 375 | 3 | 4 | 286.53 | 76.4
7-COOH-CBD | 1-disc | HQC | 375 | 3 | 5 | 306.694 | 81.8
7-COOH-CBD | 1-disc | HQC | 375 | 3 | 6 | 298.701 | 79.7
7-COOH-CBD | 1-disc | HQC | 375 | 4 | 1 | 292.422 | 78
7-COOH-CBD | 1-disc | HQC | 375 | 4 | 2 | 338.979 | 90.4
7-COOH-CBD | 1-disc | HQC | 375 | 4 | 3 | 360.619 | 96.2
7-COOH-CBD | 1-disc | HQC | 375 | 4 | 4 | 343.652 | 91.6
7-COOH-CBD | 1-disc | HQC | 375 | 4 | 5 | 316.972 | 84.5
7-COOH-CBD | 1-disc | HQC | 375 | 4 | 6 | 353.808 | 94.3
7-COOH-CBD | 1-disc | HQC | 375 | 5 | 1 | 348.017 | 92.8
7-COOH-CBD | 1-disc | HQC | 375 | 5 | 2 | 308.045 | 82.1
7-COOH-CBD | 1-disc | HQC | 375 | 5 | 3 | 278.791 | 74.3
7-COOH-CBD | 1-disc | HQC | 375 | 5 | 4 | 340.652 | 90.8
7-COOH-CBD | 1-disc | HQC | 375 | 5 | 5 | 343.941 | 91.7
7-COOH-CBD | 1-disc | HQC | 375 | 5 | 6 | 402.435 | 107.3
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 1 | 1 | 0.622 | 124.4
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 1 | 2 | 0.67 | 134
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 1 | 3 | 0.53 | 106
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 1 | 4 | 0.605 | 121
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 1 | 5 | 0.78 | 156
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 1 | 6 | 0.497 | 99.4
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 2 | 1 | 0.726 | 145.2
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 2 | 2 | 0.647 | 129.4
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 2 | 3 | 0.651 | 130.2
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 2 | 4 | 0.631 | 126.2
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 2 | 5 | 0.358 | 71.6
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 2 | 6 | 0.656 | 131.2
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 3 | 1 | 0.519 | 103.8
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 3 | 2 | 0.691 | 138.2
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 3 | 3 | 0.516 | 103.2
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 3 | 4 | 0.677 | 135.4
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 3 | 5 | 0.686 | 137.2
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 3 | 6 | 0.48 | 96
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 4 | 1 | 0.461 | 92.2
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 4 | 2 | 0.625 | 125
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 4 | 3 | 0.571 | 114.2
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 4 | 4 | 0.548 | 109.6
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 4 | 5 | 0.404 | 80.8
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 4 | 6 | 0.689 | 137.8
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 5 | 1 | 0.535 | 107
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 5 | 2 | 0.451 | 90.2
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 5 | 3 | 0.379 | 75.8
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 5 | 4 | 0.635 | 127
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 5 | 5 | 0.418 | 83.6
6α-OH-CBD | 4-disc | LLOQ | 0.5 | 5 | 6 | 0.652 | 130.4
6α-OH-CBD | 4-disc | LQC | 1.5 | 1 | 1 | 2.177 | 145.1
6α-OH-CBD | 4-disc | LQC | 1.5 | 1 | 2 | 1.672 | 111.5
6α-OH-CBD | 4-disc | LQC | 1.5 | 1 | 3 | 1.699 | 113.3
6α-OH-CBD | 4-disc | LQC | 1.5 | 1 | 4 | 1.706 | 113.7
6α-OH-CBD | 4-disc | LQC | 1.5 | 1 | 5 | 1.494 | 99.6
6α-OH-CBD | 4-disc | LQC | 1.5 | 1 | 6 | 1.908 | 127.2
6α-OH-CBD | 4-disc | LQC | 1.5 | 2 | 1 | 1.207 | 80.5
6α-OH-CBD | 4-disc | LQC | 1.5 | 2 | 2 | 1.206 | 80.4
6α-OH-CBD | 4-disc | LQC | 1.5 | 2 | 3 | 1.727 | 115.1
6α-OH-CBD | 4-disc | LQC | 1.5 | 2 | 4 | 1.173 | 78.2
6α-OH-CBD | 4-disc | LQC | 1.5 | 2 | 5 | 1.844 | 122.9
6α-OH-CBD | 4-disc | LQC | 1.5 | 2 | 6 | 1.501 | 100.1
6α-OH-CBD | 4-disc | LQC | 1.5 | 3 | 1 | 1.474 | 98.3
6α-OH-CBD | 4-disc | LQC | 1.5 | 3 | 2 | 1.57 | 104.7
6α-OH-CBD | 4-disc | LQC | 1.5 | 3 | 3 | 1.556 | 103.7
6α-OH-CBD | 4-disc | LQC | 1.5 | 3 | 4 | 1.63 | 108.7
6α-OH-CBD | 4-disc | LQC | 1.5 | 3 | 5 | 1.661 | 110.7
6α-OH-CBD | 4-disc | LQC | 1.5 | 3 | 6 | 1.516 | 101.1
6α-OH-CBD | 4-disc | LQC | 1.5 | 4 | 1 | 1.529 | 101.9
6α-OH-CBD | 4-disc | LQC | 1.5 | 4 | 2 | 1.138 | 75.9
6α-OH-CBD | 4-disc | LQC | 1.5 | 4 | 3 | 1.357 | 90.5
6α-OH-CBD | 4-disc | LQC | 1.5 | 4 | 4 | 1.618 | 107.9
6α-OH-CBD | 4-disc | LQC | 1.5 | 4 | 5 | 1.512 | 100.8
6α-OH-CBD | 4-disc | LQC | 1.5 | 4 | 6 | 1.633 | 108.9
6α-OH-CBD | 4-disc | LQC | 1.5 | 5 | 1 | 1.439 | 95.9
6α-OH-CBD | 4-disc | LQC | 1.5 | 5 | 2 | 1.321 | 88.1
6α-OH-CBD | 4-disc | LQC | 1.5 | 5 | 3 | 1.491 | 99.4
6α-OH-CBD | 4-disc | LQC | 1.5 | 5 | 4 | 1.439 | 95.9
6α-OH-CBD | 4-disc | LQC | 1.5 | 5 | 5 | 1.258 | 83.9
6α-OH-CBD | 4-disc | LQC | 1.5 | 5 | 6 | 1.693 | 112.9
6α-OH-CBD | 4-disc | MQC | 50.2 | 1 | 1 | 56.753 | 113.1
6α-OH-CBD | 4-disc | MQC | 50.2 | 1 | 2 | 55.047 | 109.7
6α-OH-CBD | 4-disc | MQC | 50.2 | 1 | 3 | 54.771 | 109.1
6α-OH-CBD | 4-disc | MQC | 50.2 | 1 | 4 | 51.816 | 103.2
6α-OH-CBD | 4-disc | MQC | 50.2 | 1 | 5 | 29.283 | 58.3
6α-OH-CBD | 4-disc | MQC | 50.2 | 1 | 6 | 48.552 | 96.7
6α-OH-CBD | 4-disc | MQC | 50.2 | 2 | 1 | 61.759 | 123
6α-OH-CBD | 4-disc | MQC | 50.2 | 2 | 2 | 44.851 | 89.3
6α-OH-CBD | 4-disc | MQC | 50.2 | 2 | 3 | 68.093 | 135.6
6α-OH-CBD | 4-disc | MQC | 50.2 | 2 | 4 | 40.91 | 81.5
6α-OH-CBD | 4-disc | MQC | 50.2 | 2 | 5 | 48.15 | 95.9
6α-OH-CBD | 4-disc | MQC | 50.2 | 2 | 6 | 61.351 | 122.2
6α-OH-CBD | 4-disc | MQC | 50.2 | 3 | 1 | 49.664 | 98.9
6α-OH-CBD | 4-disc | MQC | 50.2 | 3 | 2 | 47.431 | 94.5
6α-OH-CBD | 4-disc | MQC | 50.2 | 3 | 3 | 52.666 | 104.9
6α-OH-CBD | 4-disc | MQC | 50.2 | 3 | 4 | 65.123 | 129.7
6α-OH-CBD | 4-disc | MQC | 50.2 | 3 | 5 | 42.763 | 85.2
6α-OH-CBD | 4-disc | MQC | 50.2 | 3 | 6 | 47.566 | 94.8
6α-OH-CBD | 4-disc | MQC | 50.2 | 4 | 1 | 55.47 | 110.5
6α-OH-CBD | 4-disc | MQC | 50.2 | 4 | 2 | 55.734 | 111
6α-OH-CBD | 4-disc | MQC | 50.2 | 4 | 3 | 51.326 | 102.2
6α-OH-CBD | 4-disc | MQC | 50.2 | 4 | 4 | 43.562 | 86.8
6α-OH-CBD | 4-disc | MQC | 50.2 | 4 | 5 | 50.5 | 100.6
6α-OH-CBD | 4-disc | MQC | 50.2 | 4 | 6 | 57.328 | 114.2
6α-OH-CBD | 4-disc | MQC | 50.2 | 5 | 1 | 56.465 | 112.5
6α-OH-CBD | 4-disc | MQC | 50.2 | 5 | 2 | 49.053 | 97.7
6α-OH-CBD | 4-disc | MQC | 50.2 | 5 | 3 | 52.582 | 104.7
6α-OH-CBD | 4-disc | MQC | 50.2 | 5 | 4 | 47.712 | 95
6α-OH-CBD | 4-disc | MQC | 50.2 | 5 | 5 | 47.068 | 93.8
6α-OH-CBD | 4-disc | MQC | 50.2 | 5 | 6 | 51.054 | 101.7
6α-OH-CBD | 4-disc | HQC | 75 | 1 | 1 | 63.94 | 85.3
6α-OH-CBD | 4-disc | HQC | 75 | 1 | 2 | 62.499 | 83.3
6α-OH-CBD | 4-disc | HQC | 75 | 1 | 3 | 72.16 | 96.2
6α-OH-CBD | 4-disc | HQC | 75 | 1 | 4 | 64.53 | 86
6α-OH-CBD | 4-disc | HQC | 75 | 1 | 5 | 76.4 | 101.9
6α-OH-CBD | 4-disc | HQC | 75 | 1 | 6 | 60.644 | 80.9
6α-OH-CBD | 4-disc | HQC | 75 | 2 | 1 | 67.017 | 89.4
6α-OH-CBD | 4-disc | HQC | 75 | 2 | 2 | 64.677 | 86.2
6α-OH-CBD | 4-disc | HQC | 75 | 2 | 3 | 66.382 | 88.5
6α-OH-CBD | 4-disc | HQC | 75 | 2 | 4 | 72.73 | 97
6α-OH-CBD | 4-disc | HQC | 75 | 2 | 5 | 66.084 | 88.1
6α-OH-CBD | 4-disc | HQC | 75 | 2 | 6 | 68.207 | 90.9
6α-OH-CBD | 4-disc | HQC | 75 | 3 | 1 | 76.725 | 102.3
6α-OH-CBD | 4-disc | HQC | 75 | 3 | 2 | 66.125 | 88.2
6α-OH-CBD | 4-disc | HQC | 75 | 3 | 3 | 67.373 | 89.8
6α-OH-CBD | 4-disc | HQC | 75 | 3 | 4 | 79.829 | 106.4
6α-OH-CBD | 4-disc | HQC | 75 | 3 | 5 | 55.749 | 74.3
6α-OH-CBD | 4-disc | HQC | 75 | 3 | 6 | 68.745 | 91.7
6α-OH-CBD | 4-disc | HQC | 75 | 4 | 1 | 59.046 | 78.7
6α-OH-CBD | 4-disc | HQC | 75 | 4 | 2 | 61.992 | 82.7
6α-OH-CBD | 4-disc | HQC | 75 | 4 | 3 | 60.126 | 80.2
6α-OH-CBD | 4-disc | HQC | 75 | 4 | 4 | 63.009 | 84
6α-OH-CBD | 4-disc | HQC | 75 | 4 | 5 | 69.793 | 93.1
6α-OH-CBD | 4-disc | HQC | 75 | 4 | 6 | 62.927 | 83.9
6α-OH-CBD | 4-disc | HQC | 75 | 5 | 1 | 73.398 | 97.9
6α-OH-CBD | 4-disc | HQC | 75 | 5 | 2 | 61.691 | 82.3
6α-OH-CBD | 4-disc | HQC | 75 | 5 | 3 | 82.956 | 110.6
6α-OH-CBD | 4-disc | HQC | 75 | 5 | 4 | 59.023 | 78.7
6α-OH-CBD | 4-disc | HQC | 75 | 5 | 5 | 61.345 | 81.8
6α-OH-CBD | 4-disc | HQC | 75 | 5 | 6 | 83.127 | 110.8
CBN | 4-disc | LLOQ | 0.5 | 1 | 1 | 0.679 | 135.8
CBN | 4-disc | LLOQ | 0.5 | 1 | 2 | 0.593 | 118.6
CBN | 4-disc | LLOQ | 0.5 | 1 | 3 | 0.552 | 110.4
CBN | 4-disc | LLOQ | 0.5 | 1 | 4 | 0.729 | 145.8
CBN | 4-disc | LLOQ | 0.5 | 1 | 5 | 0.644 | 128.8
CBN | 4-disc | LLOQ | 0.5 | 1 | 6 | 0.677 | 135.4
CBN | 4-disc | LLOQ | 0.5 | 2 | 1 | 0.593 | 118.6
CBN | 4-disc | LLOQ | 0.5 | 2 | 2 | 0.535 | 107
CBN | 4-disc | LLOQ | 0.5 | 2 | 3 | 0.546 | 109.2
CBN | 4-disc | LLOQ | 0.5 | 2 | 4 | 0.528 | 105.6
CBN | 4-disc | LLOQ | 0.5 | 2 | 5 | 0.442 | 88.4
CBN | 4-disc | LLOQ | 0.5 | 2 | 6 | 0.697 | 139.4
CBN | 4-disc | LLOQ | 0.5 | 3 | 1 | 0.493 | 98.6
CBN | 4-disc | LLOQ | 0.5 | 3 | 2 | 0.617 | 123.4
CBN | 4-disc | LLOQ | 0.5 | 3 | 3 | 0.52 | 104
CBN | 4-disc | LLOQ | 0.5 | 3 | 4 | 0.472 | 94.4
CBN | 4-disc | LLOQ | 0.5 | 3 | 5 | 0.487 | 97.4
CBN | 4-disc | LLOQ | 0.5 | 3 | 6 | 0.493 | 98.6
CBN | 4-disc | LLOQ | 0.5 | 4 | 1 | 0.612 | 122.4
CBN | 4-disc | LLOQ | 0.5 | 4 | 2 | 0.453 | 90.6
CBN | 4-disc | LLOQ | 0.5 | 4 | 3 | 0.499 | 99.8
CBN | 4-disc | LLOQ | 0.5 | 4 | 4 | 0.516 | 103.2
CBN | 4-disc | LLOQ | 0.5 | 4 | 5 | 0.47 | 94
CBN | 4-disc | LLOQ | 0.5 | 4 | 6 | 0.651 | 130.2
CBN | 4-disc | LLOQ | 0.5 | 5 | 1 | 0.618 | 123.6
CBN | 4-disc | LLOQ | 0.5 | 5 | 2 | 0.627 | 125.4
CBN | 4-disc | LLOQ | 0.5 | 5 | 3 | 0.642 | 128.4
CBN | 4-disc | LLOQ | 0.5 | 5 | 4 | 0.508 | 101.6
CBN | 4-disc | LLOQ | 0.5 | 5 | 5 | 0.486 | 97.2
CBN | 4-disc | LLOQ | 0.5 | 5 | 6 | 0.514 | 102.8
CBN | 4-disc | LQC | 1.5 | 1 | 1 | 1.864 | 124.3
CBN | 4-disc | LQC | 1.5 | 1 | 2 | 1.442 | 96.1
CBN | 4-disc | LQC | 1.5 | 1 | 3 | 1.677 | 111.8
CBN | 4-disc | LQC | 1.5 | 1 | 4 | 1.615 | 107.7
CBN | 4-disc | LQC | 1.5 | 1 | 5 | 1.319 | 87.9
CBN | 4-disc | LQC | 1.5 | 1 | 6 | 1.725 | 115
CBN | 4-disc | LQC | 1.5 | 2 | 1 | 1.658 | 110.5
CBN | 4-disc | LQC | 1.5 | 2 | 2 | 1.679 | 111.9
CBN | 4-disc | LQC | 1.5 | 2 | 3 | 1.311 | 87.4
CBN | 4-disc | LQC | 1.5 | 2 | 4 | 1.798 | 119.9
CBN | 4-disc | LQC | 1.5 | 2 | 5 | 1.418 | 94.5
CBN | 4-disc | LQC | 1.5 | 2 | 6 | 1.483 | 98.9
CBN | 4-disc | LQC | 1.5 | 3 | 1 | 1.463 | 97.5
CBN | 4-disc | LQC | 1.5 | 3 | 2 | 1.576 | 105.1
CBN | 4-disc | LQC | 1.5 | 3 | 3 | 1.622 | 108.1
CBN | 4-disc | LQC | 1.5 | 3 | 4 | 1.322 | 88.1
CBN | 4-disc | LQC | 1.5 | 3 | 5 | 1.657 | 110.5
CBN | 4-disc | LQC | 1.5 | 3 | 6 | 1.6 | 106.7
CBN | 4-disc | LQC | 1.5 | 4 | 1 | 1.482 | 98.8
CBN | 4-disc | LQC | 1.5 | 4 | 2 | 1.631 | 108.7
CBN | 4-disc | LQC | 1.5 | 4 | 3 | 1.361 | 90.7
CBN | 4-disc | LQC | 1.5 | 4 | 4 | 1.477 | 98.5
CBN | 4-disc | LQC | 1.5 | 4 | 5 | 1.603 | 106.9
CBN | 4-disc | LQC | 1.5 | 4 | 6 | 1.405 | 93.7
CBN | 4-disc | LQC | 1.5 | 5 | 1 | 1.184 | 78.9
CBN | 4-disc | LQC | 1.5 | 5 | 2 | 1.637 | 109.1
CBN | 4-disc | LQC | 1.5 | 5 | 3 | 1.341 | 89.4
CBN | 4-disc | LQC | 1.5 | 5 | 4 | 1.6 | 106.7
CBN | 4-disc | LQC | 1.5 | 5 | 5 | 1.765 | 117.7
CBN | 4-disc | LQC | 1.5 | 5 | 6 | 1.498 | 99.9
CBN | 4-disc | MQC | 50.2 | 1 | 1 | 43.906 | 87.5
CBN | 4-disc | MQC | 50.2 | 1 | 2 | 52.615 | 104.8
CBN | 4-disc | MQC | 50.2 | 1 | 3 | 54.041 | 107.7
CBN | 4-disc | MQC | 50.2 | 1 | 4 | 50.041 | 99.7
CBN | 4-disc | MQC | 50.2 | 1 | 5 | 54.875 | 109.3
CBN | 4-disc | MQC | 50.2 | 1 | 6 | 51.48 | 102.5
CBN | 4-disc | MQC | 50.2 | 2 | 1 | 44.679 | 89
CBN | 4-disc | MQC | 50.2 | 2 | 2 | 62.284 | 124.1
CBN | 4-disc | MQC | 50.2 | 2 | 3 | 61.42 | 122.4
CBN | 4-disc | MQC | 50.2 | 2 | 4 | 56.072 | 111.7
CBN | 4-disc | MQC | 50.2 | 2 | 5 | 62.657 | 124.8
CBN | 4-disc | MQC | 50.2 | 2 | 6 | 47.656 | 94.9
CBN | 4-disc | MQC | 50.2 | 3 | 1 | 53.543 | 106.7
CBN | 4-disc | MQC | 50.2 | 3 | 2 | 51.165 | 101.9
CBN | 4-disc | MQC | 50.2 | 3 | 3 | 44.834 | 89.3
CBN | 4-disc | MQC | 50.2 | 3 | 4 | 52.102 | 103.8
CBN | 4-disc | MQC | 50.2 | 3 | 5 | 45.285 | 90.2
CBN | 4-disc | MQC | 50.2 | 3 | 6 | 55.242 | 110
CBN | 4-disc | MQC | 50.2 | 4 | 1 | 54.975 | 109.5
CBN | 4-disc | MQC | 50.2 | 4 | 2 | 46.473 | 92.6
CBN | 4-disc | MQC | 50.2 | 4 | 3 | 51.257 | 102.1
CBN | 4-disc | MQC | 50.2 | 4 | 4 | 58.469 | 116.5
CBN | 4-disc | MQC | 50.2 | 4 | 5 | 49.821 | 99.2
CBN | 4-disc | MQC | 50.2 | 4 | 6 | 39.946 | 79.6
CBN | 4-disc | MQC | 50.2 | 5 | 1 | 50.932 | 101.5
CBN | 4-disc | MQC | 50.2 | 5 | 2 | 49.817 | 99.2
CBN | 4-disc | MQC | 50.2 | 5 | 3 | 51.815 | 103.2
CBN | 4-disc | MQC | 50.2 | 5 | 4 | 52.174 | 103.9
CBN | 4-disc | MQC | 50.2 | 5 | 5 | 45.325 | 90.3
CBN | 4-disc | MQC | 50.2 | 5 | 6 | 51.76 | 103.1
CBN | 4-disc | HQC | 75 | 1 | 1 | 79.707 | 106.3
CBN | 4-disc | HQC | 75 | 1 | 2 | 69.189 | 92.3
CBN | 4-disc | HQC | 75 | 1 | 3 | 76.748 | 102.3
CBN | 4-disc | HQC | 75 | 1 | 4 | 69.244 | 92.3
CBN | 4-disc | HQC | 75 | 1 | 5 | 75.767 | 101
CBN | 4-disc | HQC | 75 | 1 | 6 | 72.469 | 96.6
CBN | 4-disc | HQC | 75 | 2 | 1 | 71.841 | 95.8
CBN | 4-disc | HQC | 75 | 2 | 2 | 70.42 | 93.9
CBN | 4-disc | HQC | 75 | 2 | 3 | 66.351 | 88.5
CBN | 4-disc | HQC | 75 | 2 | 4 | 71.423 | 95.2
CBN | 4-disc | HQC | 75 | 2 | 5 | 69.161 | 92.2
CBN | 4-disc | HQC | 75 | 2 | 6 | 71.878 | 95.8
CBN | 4-disc | HQC | 75 | 3 | 1 | 62.847 | 83.8
CBN | 4-disc | HQC | 75 | 3 | 2 | 74.58 | 99.4
CBN | 4-disc | HQC | 75 | 3 | 3 | 73.541 | 98.1
CBN | 4-disc | HQC | 75 | 3 | 4 | 64.894 | 86.5
CBN | 4-disc | HQC | 75 | 3 | 5 | 70.395 | 93.9
CBN | 4-disc | HQC | 75 | 3 | 6 | 64.235 | 85.6
CBN | 4-disc | HQC | 75 | 4 | 1 | 65.09 | 86.8
CBN | 4-disc | HQC | 75 | 4 | 2 | 57.817 | 77.1
CBN | 4-disc | HQC | 75 | 4 | 3 | 68.449 | 91.3
CBN | 4-disc | HQC | 75 | 4 | 4 | 64.357 | 85.8
CBN | 4-disc | HQC | 75 | 4 | 5 | 68.623 | 91.5
CBN | 4-disc | HQC | 75 | 4 | 6 | 69.834 | 93.1
CBN | 4-disc | HQC | 75 | 5 | 1 | 66.901 | 89.2
CBN | 4-disc | HQC | 75 | 5 | 2 | 69.469 | 92.6
CBN | 4-disc | HQC | 75 | 5 | 3 | 75.768 | 101
CBN | 4-disc | HQC | 75 | 5 | 4 | 64.579 | 86.1
CBN | 4-disc | HQC | 75 | 5 | 5 | 66.639 | 88.9
CBN | 4-disc | HQC | 75 | 5 | 6 | 75.784 | 101
CBG | 4-disc | LLOQ | 0.5 | 1 | 1 | 0.462 | 92.4
CBG | 4-disc | LLOQ | 0.5 | 1 | 2 | 0.415 | 83
CBG | 4-disc | LLOQ | 0.5 | 1 | 3 | 0.486 | 97.2
CBG | 4-disc | LLOQ | 0.5 | 1 | 4 | 0.451 | 90.2
CBG | 4-disc | LLOQ | 0.5 | 1 | 5 | 0.672 | 134.4
CBG | 4-disc | LLOQ | 0.5 | 1 | 6 | 0.599 | 119.8
CBG | 4-disc | LLOQ | 0.5 | 2 | 1 | 0.54 | 108
CBG | 4-disc | LLOQ | 0.5 | 2 | 2 | 0.636 | 127.2
CBG | 4-disc | LLOQ | 0.5 | 2 | 3 | 0.449 | 89.8
CBG | 4-disc | LLOQ | 0.5 | 2 | 4 | 0.519 | 103.8
CBG | 4-disc | LLOQ | 0.5 | 2 | 5 | 0.355 | 71
CBG | 4-disc | LLOQ | 0.5 | 2 | 6 | 0.512 | 102.4
CBG | 4-disc | LLOQ | 0.5 | 3 | 1 | 0.635 | 127
CBG | 4-disc | LLOQ | 0.5 | 3 | 2 | 0.646 | 129.2
CBG | 4-disc | LLOQ | 0.5 | 3 | 3 | 0.557 | 111.4
CBG | 4-disc | LLOQ | 0.5 | 3 | 4 | 0.714 | 142.8
CBG | 4-disc | LLOQ | 0.5 | 3 | 5 | 0.57 | 114
CBG | 4-disc | LLOQ | 0.5 | 3 | 6 | 0.53 | 106
CBG | 4-disc | LLOQ | 0.5 | 4 | 1 | 0.721 | 144.2
CBG | 4-disc | LLOQ | 0.5 | 4 | 2 | 0.673 | 134.6
CBG | 4-disc | LLOQ | 0.5 | 4 | 3 | 0.478 | 95.6
CBG | 4-disc | LLOQ | 0.5 | 4 | 4 | 0.65 | 130
CBG | 4-disc | LLOQ | 0.5 | 4 | 5 | 0.654 | 130.8
CBG | 4-disc | LLOQ | 0.5 | 4 | 6 | 0.627 | 125.4
CBG | 4-disc | LLOQ | 0.5 | 5 | 1 | 0.692 | 138.4
CBG | 4-disc | LLOQ | 0.5 | 5 | 2 | 0.682 | 136.4
CBG | 4-disc | LLOQ | 0.5 | 5 | 3 | 0.649 | 129.8
CBG | 4-disc | LLOQ | 0.5 | 5 | 4 | 0.541 | 108.2
CBG | 4-disc | LLOQ | 0.5 | 5 | 5 | 0.532 | 106.4
CBG | 4-disc | LLOQ | 0.5 | 5 | 6 | 0.438 | 87.6
CBG | 4-disc | LQC | 1.5 | 1 | 1 | 1.424 | 94.9
CBG | 4-disc | LQC | 1.5 | 1 | 2 | 1.346 | 89.7
CBG | 4-disc | LQC | 1.5 | 1 | 3 | 1.731 | 115.4
CBG | 4-disc | LQC | 1.5 | 1 | 4 | 1.451 | 96.7
CBG | 4-disc | LQC | 1.5 | 1 | 5 | 1.469 | 97.9
CBG | 4-disc | LQC | 1.5 | 1 | 6 | 1.623 | 108.2
CBG | 4-disc | LQC | 1.5 | 2 | 1 | 1.233 | 82.2
CBG | 4-disc | LQC | 1.5 | 2 | 2 | 1.562 | 104.1
CBG | 4-disc | LQC | 1.5 | 2 | 3 | 1.249 | 83.3
CBG | 4-disc | LQC | 1.5 | 2 | 4 | 1.531 | 102.1
CBG | 4-disc | LQC | 1.5 | 2 | 5 | 1.743 | 116.2
CBG | 4-disc | LQC | 1.5 | 2 | 6 | 1.687 | 112.5
CBG | 4-disc | LQC | 1.5 | 3 | 1 | 1.781 | 118.7
CBG | 4-disc | LQC | 1.5 | 3 | 2 | 1.922 | 128.1
CBG | 4-disc | LQC | 1.5 | 3 | 3 | 1.417 | 94.5
CBG | 4-disc | LQC | 1.5 | 3 | 4 | 1.757 | 117.1
CBG | 4-disc | LQC | 1.5 | 3 | 5 | 1.415 | 94.3
CBG | 4-disc | LQC | 1.5 | 3 | 6 | 1.467 | 97.8
CBG | 4-disc | LQC | 1.5 | 4 | 1 | 1.345 | 89.7
CBG | 4-disc | LQC | 1.5 | 4 | 2 | 1.654 | 110.3
CBG | 4-disc | LQC | 1.5 | 4 | 3 | 1.4 | 93.3
CBG | 4-disc | LQC | 1.5 | 4 | 4 | 1.426 | 95.1
CBG | 4-disc | LQC | 1.5 | 4 | 5 | 1.593 | 106.2
CBG | 4-disc | LQC | 1.5 | 4 | 6 | 1.373 | 91.5
CBG | 4-disc | LQC | 1.5 | 5 | 1 | 1.588 | 105.9
CBG | 4-disc | LQC | 1.5 | 5 | 2 | 2.147 | 143.1
CBG | 4-disc | LQC | 1.5 | 5 | 3 | 1.486 | 99.1
CBG | 4-disc | LQC | 1.5 | 5 | 4 | 1.477 | 98.5
CBG | 4-disc | LQC | 1.5 | 5 | 5 | 1.226 | 81.7
CBG | 4-disc | LQC | 1.5 | 5 | 6 | 1.6 | 106.7
CBG | 4-disc | MQC | 50.2 | 1 | 1 | 50.076 | 99.8
CBG | 4-disc | MQC | 50.2 | 1 | 2 | 53.994 | 107.6
CBG | 4-disc | MQC | 50.2 | 1 | 3 | 57.434 | 114.4
CBG | 4-disc | MQC | 50.2 | 1 | 4 | 44.51 | 88.7
CBG | 4-disc | MQC | 50.2 | 1 | 5 | 50.777 | 101.1
CBG | 4-disc | MQC | 50.2 | 1 | 6 | 48.424 | 96.5
CBG | 4-disc | MQC | 50.2 | 2 | 1 | 47.183 | 94
CBG | 4-disc | MQC | 50.2 | 2 | 2 | 49.218 | 98
CBG | 4-disc | MQC | 50.2 | 2 | 3 | 52.655 | 104.9
CBG | 4-disc | MQC | 50.2 | 2 | 4 | 57.457 | 114.5
CBG | 4-disc | MQC | 50.2 | 2 | 5 | 46.091 | 91.8
CBG | 4-disc | MQC | 50.2 | 2 | 6 | 63.76 | 127
CBG | 4-disc | MQC | 50.2 | 3 | 1 | 47.475 | 94.6
CBG | 4-disc | MQC | 50.2 | 3 | 2 | 50.593 | 100.8
CBG | 4-disc | MQC | 50.2 | 3 | 3 | 39.805 | 79.3
CBG | 4-disc | MQC | 50.2 | 3 | 4 | 52.249 | 104.1
CBG | 4-disc | MQC | 50.2 | 3 | 5 | 60.053 | 119.6
CBG | 4-disc | MQC | 50.2 | 3 | 6 | 51.113 | 101.8
CBG | 4-disc | MQC | 50.2 | 4 | 1 | 52.163 | 103.9
CBG | 4-disc | MQC | 50.2 | 4 | 2 | 35.871 | 71.5
CBG | 4-disc | MQC | 50.2 | 4 | 3 | 54.978 | 109.5
CBG | 4-disc | MQC | 50.2 | 4 | 4 | 56.206 | 112
CBG | 4-disc | MQC | 50.2 | 4 | 5 | 35.823 | 71.4
CBG | 4-disc | MQC | 50.2 | 4 | 6 | 56.891 | 113.3
CBG | 4-disc | MQC | 50.2 | 5 | 1 | 50.116 | 99.8
CBG | 4-disc | MQC | 50.2 | 5 | 2 | 52.196 | 104
CBG | 4-disc | MQC | 50.2 | 5 | 3 | 58.372 | 116.3
CBG | 4-disc | MQC | 50.2 | 5 | 4 | 65.544 | 130.6
CBG | 4-disc | MQC | 50.2 | 5 | 5 | 53.295 | 106.2
CBG | 4-disc | MQC | 50.2 | 5 | 6 | 49.328 | 98.3
CBG | 4-disc | HQC | 75 | 1 | 1 | 64.704 | 86.3
CBG | 4-disc | HQC | 75 | 1 | 2 | 71.069 | 94.8
CBG | 4-disc | HQC | 75 | 1 | 3 | 60.659 | 80.9
CBG | 4-disc | HQC | 75 | 1 | 4 | 78.73 | 105
CBG | 4-disc | HQC | 75 | 1 | 5 | 64.744 | 86.3
CBG | 4-disc | HQC | 75 | 1 | 6 | 59.66 | 79.5
CBG | 4-disc | HQC | 75 | 2 | 1 | 63.782 | 85
CBG | 4-disc | HQC | 75 | 2 | 2 | 64.489 | 86
CBG | 4-disc | HQC | 75 | 2 | 3 | 74.699 | 99.6
CBG | 4-disc | HQC | 75 | 2 | 4 | 71.348 | 95.1
CBG | 4-disc | HQC | 75 | 2 | 5 | 64.101 | 85.5
CBG | 4-disc | HQC | 75 | 2 | 6 | 62.834 | 83.8
CBG | 4-disc | HQC | 75 | 3 | 1 | 78.841 | 105.1
CBG | 4-disc | HQC | 75 | 3 | 2 | 78.584 | 104.8
CBG | 4-disc | HQC | 75 | 3 | 3 | 68.094 | 90.8
CBG | 4-disc | HQC | 75 | 3 | 4 | 65.145 | 86.9
CBG | 4-disc | HQC | 75 | 3 | 5 | 67.712 | 90.3
CBG | 4-disc | HQC | 75 | 3 | 6 | 74.825 | 99.8
CBG | 4-disc | HQC | 75 | 4 | 1 | 59.132 | 78.8
CBG | 4-disc | HQC | 75 | 4 | 2 | 72.62 | 96.8
CBG | 4-disc | HQC | 75 | 4 | 3 | 67.803 | 90.4
CBG | 4-disc | HQC | 75 | 4 | 4 | 64.647 | 86.2
CBG | 4-disc | HQC | 75 | 4 | 5 | 66.773 | 89
CBG | 4-disc | HQC | 75 | 4 | 6 | 66.419 | 88.6
CBG | 4-disc | HQC | 75 | 5 | 1 | 55.746 | 74.3
CBG | 4-disc | HQC | 75 | 5 | 2 | 68.81 | 91.7
CBG | 4-disc | HQC | 75 | 5 | 3 | 67.159 | 89.5
CBG | 4-disc | HQC | 75 | 5 | 4 | 78.633 | 104.8
CBG | 4-disc | HQC | 75 | 5 | 5 | 68.26 | 91
CBG | 4-disc | HQC | 75 | 5 | 6 | 79.729 | 106.3

## HCT_Independence
Hematocrit independence — IS-norm EY% at HCT 0.26/0.41/0.55, LQC&HQC, 6 donors.
Analyte | HCT Level | HCT Value | QC Level | Donor | IS-norm EY% | Deviation vs normal (%)
Δ9-THC | Low | 0.26 | LQC | 1 | 85.2 | 10.3
Δ9-THC | Low | 0.26 | LQC | 2 | 84.9 | 9.8
Δ9-THC | Low | 0.26 | LQC | 3 | 84.9 | 9.9
Δ9-THC | Low | 0.26 | LQC | 4 | 85.3 | 10.5
Δ9-THC | Low | 0.26 | LQC | 5 | 85.8 | 11.1
Δ9-THC | Low | 0.26 | LQC | 6 | 85.1 | 10.2
Δ9-THC | Low | 0.26 | HQC | 1 | 86.1 | 11.5
Δ9-THC | Low | 0.26 | HQC | 2 | 85.2 | 10.2
Δ9-THC | Low | 0.26 | HQC | 3 | 85.2 | 10.2
Δ9-THC | Low | 0.26 | HQC | 4 | 85.9 | 11.2
Δ9-THC | Low | 0.26 | HQC | 5 | 85.5 | 10.6
Δ9-THC | Low | 0.26 | HQC | 6 | 85.4 | 10.6
Δ9-THC | Normal | 0.41 | LQC | 1 | 77.5 | 0.3
Δ9-THC | Normal | 0.41 | LQC | 2 | 77.5 | 0.3
Δ9-THC | Normal | 0.41 | LQC | 3 | 76.7 | -0.7
Δ9-THC | Normal | 0.41 | LQC | 4 | 77.7 | 0.5
Δ9-THC | Normal | 0.41 | LQC | 5 | 78 | 0.9
Δ9-THC | Normal | 0.41 | LQC | 6 | 77.7 | 0.5
Δ9-THC | Normal | 0.41 | HQC | 1 | 77 | -0.3
Δ9-THC | Normal | 0.41 | HQC | 2 | 77 | -0.3
Δ9-THC | Normal | 0.41 | HQC | 3 | 77.3 | 0.1
Δ9-THC | Normal | 0.41 | HQC | 4 | 77.7 | 0.6
Δ9-THC | Normal | 0.41 | HQC | 5 | 77.6 | 0.5
Δ9-THC | Normal | 0.41 | HQC | 6 | 77.9 | 0.9
Δ9-THC | High | 0.55 | LQC | 1 | 69.8 | -9.7
Δ9-THC | High | 0.55 | LQC | 2 | 69.5 | -10.1
Δ9-THC | High | 0.55 | LQC | 3 | 70.1 | -9.3
Δ9-THC | High | 0.55 | LQC | 4 | 69.2 | -10.4
Δ9-THC | High | 0.55 | LQC | 5 | 69.2 | -10.4
Δ9-THC | High | 0.55 | LQC | 6 | 69.3 | -10.3
Δ9-THC | High | 0.55 | HQC | 1 | 69.7 | -9.8
Δ9-THC | High | 0.55 | HQC | 2 | 68.9 | -10.9
Δ9-THC | High | 0.55 | HQC | 3 | 69.4 | -10.2
Δ9-THC | High | 0.55 | HQC | 4 | 69.3 | -10.3
Δ9-THC | High | 0.55 | HQC | 5 | 69.2 | -10.4
Δ9-THC | High | 0.55 | HQC | 6 | 69.1 | -10.6
11-OH-THC | Low | 0.26 | LQC | 1 | 88 | 9.2
11-OH-THC | Low | 0.26 | LQC | 2 | 88.6 | 10
11-OH-THC | Low | 0.26 | LQC | 3 | 87.7 | 8.8
11-OH-THC | Low | 0.26 | LQC | 4 | 88.2 | 9.4
11-OH-THC | Low | 0.26 | LQC | 5 | 88.8 | 10.2
11-OH-THC | Low | 0.26 | LQC | 6 | 88.6 | 9.9
11-OH-THC | Low | 0.26 | HQC | 1 | 88 | 9.3
11-OH-THC | Low | 0.26 | HQC | 2 | 89.2 | 10.7
11-OH-THC | Low | 0.26 | HQC | 3 | 88.8 | 10.1
11-OH-THC | Low | 0.26 | HQC | 4 | 88.5 | 9.8
11-OH-THC | Low | 0.26 | HQC | 5 | 89.2 | 10.7
11-OH-THC | Low | 0.26 | HQC | 6 | 87.7 | 8.9
11-OH-THC | Normal | 0.41 | LQC | 1 | 80.7 | 0.1
11-OH-THC | Normal | 0.41 | LQC | 2 | 81.2 | 0.7
11-OH-THC | Normal | 0.41 | LQC | 3 | 79.9 | -0.9
11-OH-THC | Normal | 0.41 | LQC | 4 | 80.3 | -0.3
11-OH-THC | Normal | 0.41 | LQC | 5 | 80 | -0.8
11-OH-THC | Normal | 0.41 | LQC | 6 | 80.8 | 0.2
11-OH-THC | Normal | 0.41 | HQC | 1 | 80.3 | -0.4
11-OH-THC | Normal | 0.41 | HQC | 2 | 80.6 | 0
11-OH-THC | Normal | 0.41 | HQC | 3 | 80.2 | -0.4
11-OH-THC | Normal | 0.41 | HQC | 4 | 80.3 | -0.3
11-OH-THC | Normal | 0.41 | HQC | 5 | 80.1 | -0.6
11-OH-THC | Normal | 0.41 | HQC | 6 | 81.2 | 0.8
11-OH-THC | High | 0.55 | LQC | 1 | 74.6 | -7.4
11-OH-THC | High | 0.55 | LQC | 2 | 74.5 | -7.5
11-OH-THC | High | 0.55 | LQC | 3 | 74.5 | -7.5
11-OH-THC | High | 0.55 | LQC | 4 | 75.9 | -5.8
11-OH-THC | High | 0.55 | LQC | 5 | 75.3 | -6.6
11-OH-THC | High | 0.55 | LQC | 6 | 74.8 | -7.2
11-OH-THC | High | 0.55 | HQC | 1 | 74.9 | -7.1
11-OH-THC | High | 0.55 | HQC | 2 | 76 | -5.6
11-OH-THC | High | 0.55 | HQC | 3 | 74.7 | -7.3
11-OH-THC | High | 0.55 | HQC | 4 | 75.5 | -6.3
11-OH-THC | High | 0.55 | HQC | 5 | 76.1 | -5.6
11-OH-THC | High | 0.55 | HQC | 6 | 75.6 | -6.2
THC-COOH | Low | 0.26 | LQC | 1 | 76.3 | 8
THC-COOH | Low | 0.26 | LQC | 2 | 76.2 | 7.8
THC-COOH | Low | 0.26 | LQC | 3 | 77 | 9
THC-COOH | Low | 0.26 | LQC | 4 | 76.4 | 8.2
THC-COOH | Low | 0.26 | LQC | 5 | 77 | 9
THC-COOH | Low | 0.26 | LQC | 6 | 76.3 | 8
THC-COOH | Low | 0.26 | HQC | 1 | 76.9 | 8.8
THC-COOH | Low | 0.26 | HQC | 2 | 76.6 | 8.5
THC-COOH | Low | 0.26 | HQC | 3 | 76.9 | 8.9
THC-COOH | Low | 0.26 | HQC | 4 | 76.3 | 8
THC-COOH | Low | 0.26 | HQC | 5 | 76.7 | 8.6
THC-COOH | Low | 0.26 | HQC | 6 | 76.1 | 7.8
THC-COOH | Normal | 0.41 | LQC | 1 | 70 | -0.9
THC-COOH | Normal | 0.41 | LQC | 2 | 70.6 | -0.1
THC-COOH | Normal | 0.41 | LQC | 3 | 69.9 | -1
THC-COOH | Normal | 0.41 | LQC | 4 | 70.6 | 0
THC-COOH | Normal | 0.41 | LQC | 5 | 70.6 | 0
THC-COOH | Normal | 0.41 | LQC | 6 | 70.1 | -0.8
THC-COOH | Normal | 0.41 | HQC | 1 | 70.8 | 0.3
THC-COOH | Normal | 0.41 | HQC | 2 | 71.1 | 0.7
THC-COOH | Normal | 0.41 | HQC | 3 | 70.3 | -0.5
THC-COOH | Normal | 0.41 | HQC | 4 | 70.3 | -0.5
THC-COOH | Normal | 0.41 | HQC | 5 | 70.9 | 0.3
THC-COOH | Normal | 0.41 | HQC | 6 | 71 | 0.5
THC-COOH | High | 0.55 | LQC | 1 | 61.8 | -12.5
THC-COOH | High | 0.55 | LQC | 2 | 61.5 | -12.9
THC-COOH | High | 0.55 | LQC | 3 | 62 | -12.2
THC-COOH | High | 0.55 | LQC | 4 | 61.6 | -12.8
THC-COOH | High | 0.55 | LQC | 5 | 62 | -12.2
THC-COOH | High | 0.55 | LQC | 6 | 61 | -13.6
THC-COOH | High | 0.55 | HQC | 1 | 61.9 | -12.3
THC-COOH | High | 0.55 | HQC | 2 | 61 | -13.6
THC-COOH | High | 0.55 | HQC | 3 | 61.6 | -12.8
THC-COOH | High | 0.55 | HQC | 4 | 61.3 | -13.2
THC-COOH | High | 0.55 | HQC | 5 | 61.8 | -12.5
THC-COOH | High | 0.55 | HQC | 6 | 61 | -13.6
CBD | Low | 0.26 | LQC | 1 | 78.1 | 7.4
CBD | Low | 0.26 | LQC | 2 | 78.9 | 8.5
CBD | Low | 0.26 | LQC | 3 | 78.8 | 8.3
CBD | Low | 0.26 | LQC | 4 | 78.8 | 8.4
CBD | Low | 0.26 | LQC | 5 | 77.8 | 7
CBD | Low | 0.26 | LQC | 6 | 79 | 8.7
CBD | Low | 0.26 | HQC | 1 | 77.9 | 7.1
CBD | Low | 0.26 | HQC | 2 | 78.7 | 8.1
CBD | Low | 0.26 | HQC | 3 | 77.9 | 7.1
CBD | Low | 0.26 | HQC | 4 | 79 | 8.6
CBD | Low | 0.26 | HQC | 5 | 78.2 | 7.5
CBD | Low | 0.26 | HQC | 6 | 78.6 | 8.1
CBD | Normal | 0.41 | LQC | 1 | 73 | 0.4
CBD | Normal | 0.41 | LQC | 2 | 72 | -1
CBD | Normal | 0.41 | LQC | 3 | 73.4 | 0.9
CBD | Normal | 0.41 | LQC | 4 | 72.2 | -0.7
CBD | Normal | 0.41 | LQC | 5 | 72.3 | -0.5
CBD | Normal | 0.41 | LQC | 6 | 72.7 | 0
CBD | Normal | 0.41 | HQC | 1 | 72.4 | -0.4
CBD | Normal | 0.41 | HQC | 2 | 72.5 | -0.3
CBD | Normal | 0.41 | HQC | 3 | 73 | 0.4
CBD | Normal | 0.41 | HQC | 4 | 72.7 | 0
CBD | Normal | 0.41 | HQC | 5 | 72.3 | -0.6
CBD | Normal | 0.41 | HQC | 6 | 72.8 | 0
CBD | High | 0.55 | LQC | 1 | 64 | -12
CBD | High | 0.55 | LQC | 2 | 64.1 | -11.9
CBD | High | 0.55 | LQC | 3 | 64.5 | -11.3
CBD | High | 0.55 | LQC | 4 | 64.6 | -11.2
CBD | High | 0.55 | LQC | 5 | 64.8 | -10.9
CBD | High | 0.55 | LQC | 6 | 64 | -12
CBD | High | 0.55 | HQC | 1 | 64.8 | -10.9
CBD | High | 0.55 | HQC | 2 | 63.7 | -12.4
CBD | High | 0.55 | HQC | 3 | 63.7 | -12.5
CBD | High | 0.55 | HQC | 4 | 64.2 | -11.7
CBD | High | 0.55 | HQC | 5 | 64.5 | -11.3
CBD | High | 0.55 | HQC | 6 | 63.8 | -12.3
7-OH-CBD | Low | 0.26 | LQC | 1 | 84.2 | 10.6
7-OH-CBD | Low | 0.26 | LQC | 2 | 84.4 | 10.8
7-OH-CBD | Low | 0.26 | LQC | 3 | 84.5 | 10.9
7-OH-CBD | Low | 0.26 | LQC | 4 | 83.7 | 9.8
7-OH-CBD | Low | 0.26 | LQC | 5 | 84 | 10.2
7-OH-CBD | Low | 0.26 | LQC | 6 | 84.2 | 10.5
7-OH-CBD | Low | 0.26 | HQC | 1 | 83.8 | 10
7-OH-CBD | Low | 0.26 | HQC | 2 | 85 | 11.5
7-OH-CBD | Low | 0.26 | HQC | 3 | 84.9 | 11.4
7-OH-CBD | Low | 0.26 | HQC | 4 | 84.5 | 10.9
7-OH-CBD | Low | 0.26 | HQC | 5 | 83.8 | 10
7-OH-CBD | Low | 0.26 | HQC | 6 | 84.9 | 11.5
7-OH-CBD | Normal | 0.41 | LQC | 1 | 76.8 | 0.8
7-OH-CBD | Normal | 0.41 | LQC | 2 | 76.8 | 0.9
7-OH-CBD | Normal | 0.41 | LQC | 3 | 76.8 | 0.8
7-OH-CBD | Normal | 0.41 | LQC | 4 | 75.8 | -0.6
7-OH-CBD | Normal | 0.41 | LQC | 5 | 76.9 | 0.9
7-OH-CBD | Normal | 0.41 | LQC | 6 | 76.1 | -0.2
7-OH-CBD | Normal | 0.41 | HQC | 1 | 76.7 | 0.7
7-OH-CBD | Normal | 0.41 | HQC | 2 | 76.8 | 0.8
7-OH-CBD | Normal | 0.41 | HQC | 3 | 76.8 | 0.8
7-OH-CBD | Normal | 0.41 | HQC | 4 | 76.2 | 0
7-OH-CBD | Normal | 0.41 | HQC | 5 | 76.5 | 0.4
7-OH-CBD | Normal | 0.41 | HQC | 6 | 75.7 | -0.6
7-OH-CBD | High | 0.55 | LQC | 1 | 67.6 | -11.3
7-OH-CBD | High | 0.55 | LQC | 2 | 68 | -10.8
7-OH-CBD | High | 0.55 | LQC | 3 | 68.5 | -10.1
7-OH-CBD | High | 0.55 | LQC | 4 | 67.7 | -11.2
7-OH-CBD | High | 0.55 | LQC | 5 | 68.6 | -10
7-OH-CBD | High | 0.55 | LQC | 6 | 68.4 | -10.3
7-OH-CBD | High | 0.55 | HQC | 1 | 67.7 | -11.1
7-OH-CBD | High | 0.55 | HQC | 2 | 68.1 | -10.7
7-OH-CBD | High | 0.55 | HQC | 3 | 68.6 | -10
7-OH-CBD | High | 0.55 | HQC | 4 | 67.9 | -10.9
7-OH-CBD | High | 0.55 | HQC | 5 | 67.4 | -11.5
7-OH-CBD | High | 0.55 | HQC | 6 | 67.5 | -11.4
7-COOH-CBD | Low | 0.26 | LQC | 1 | 81.2 | 9.4
7-COOH-CBD | Low | 0.26 | LQC | 2 | 79.9 | 7.6
7-COOH-CBD | Low | 0.26 | LQC | 3 | 81.2 | 9.5
7-COOH-CBD | Low | 0.26 | LQC | 4 | 79.9 | 7.6
7-COOH-CBD | Low | 0.26 | LQC | 5 | 80 | 7.9
7-COOH-CBD | Low | 0.26 | LQC | 6 | 81.1 | 9.3
7-COOH-CBD | Low | 0.26 | HQC | 1 | 81.2 | 9.4
7-COOH-CBD | Low | 0.26 | HQC | 2 | 80.2 | 8.1
7-COOH-CBD | Low | 0.26 | HQC | 3 | 81.1 | 9.4
7-COOH-CBD | Low | 0.26 | HQC | 4 | 80.9 | 9.1
7-COOH-CBD | Low | 0.26 | HQC | 5 | 81.1 | 9.4
7-COOH-CBD | Low | 0.26 | HQC | 6 | 81.1 | 9.4
7-COOH-CBD | Normal | 0.41 | LQC | 1 | 74.3 | 0.2
7-COOH-CBD | Normal | 0.41 | LQC | 2 | 74.8 | 0.8
7-COOH-CBD | Normal | 0.41 | LQC | 3 | 74.6 | 0.5
7-COOH-CBD | Normal | 0.41 | LQC | 4 | 74.7 | 0.6
7-COOH-CBD | Normal | 0.41 | LQC | 5 | 74.2 | 0
7-COOH-CBD | Normal | 0.41 | LQC | 6 | 74.7 | 0.7
7-COOH-CBD | Normal | 0.41 | HQC | 1 | 73.8 | -0.5
7-COOH-CBD | Normal | 0.41 | HQC | 2 | 74 | -0.3
7-COOH-CBD | Normal | 0.41 | HQC | 3 | 73.5 | -0.9
7-COOH-CBD | Normal | 0.41 | HQC | 4 | 74.6 | 0.5
7-COOH-CBD | Normal | 0.41 | HQC | 5 | 74.4 | 0.2
7-COOH-CBD | Normal | 0.41 | HQC | 6 | 74.8 | 0.8
7-COOH-CBD | High | 0.55 | LQC | 1 | 66.8 | -10
7-COOH-CBD | High | 0.55 | LQC | 2 | 66.7 | -10
7-COOH-CBD | High | 0.55 | LQC | 3 | 67.4 | -9.2
7-COOH-CBD | High | 0.55 | LQC | 4 | 66.6 | -10.2
7-COOH-CBD | High | 0.55 | LQC | 5 | 67.9 | -8.5
7-COOH-CBD | High | 0.55 | LQC | 6 | 67.5 | -9
7-COOH-CBD | High | 0.55 | HQC | 1 | 67.2 | -9.4
7-COOH-CBD | High | 0.55 | HQC | 2 | 67.5 | -9
7-COOH-CBD | High | 0.55 | HQC | 3 | 67.6 | -8.8
7-COOH-CBD | High | 0.55 | HQC | 4 | 67 | -9.7
7-COOH-CBD | High | 0.55 | HQC | 5 | 67.5 | -9
7-COOH-CBD | High | 0.55 | HQC | 6 | 66.7 | -10.1
6α-OH-CBD | Low | 0.26 | LQC | 1 | 75.5 | 6.5
6α-OH-CBD | Low | 0.26 | LQC | 2 | 76.2 | 7.5
6α-OH-CBD | Low | 0.26 | LQC | 3 | 75.9 | 7
6α-OH-CBD | Low | 0.26 | LQC | 4 | 75.5 | 6.5
6α-OH-CBD | Low | 0.26 | LQC | 5 | 75.7 | 6.8
6α-OH-CBD | Low | 0.26 | LQC | 6 | 76.4 | 7.7
6α-OH-CBD | Low | 0.26 | HQC | 1 | 75.8 | 6.9
6α-OH-CBD | Low | 0.26 | HQC | 2 | 76.4 | 7.7
6α-OH-CBD | Low | 0.26 | HQC | 3 | 75.7 | 6.8
6α-OH-CBD | Low | 0.26 | HQC | 4 | 76 | 7.1
6α-OH-CBD | Low | 0.26 | HQC | 5 | 75.5 | 6.5
6α-OH-CBD | Low | 0.26 | HQC | 6 | 75.8 | 6.9
6α-OH-CBD | Normal | 0.41 | LQC | 1 | 70.5 | -0.6
6α-OH-CBD | Normal | 0.41 | LQC | 2 | 71 | 0.1
6α-OH-CBD | Normal | 0.41 | LQC | 3 | 70.3 | -0.8
6α-OH-CBD | Normal | 0.41 | LQC | 4 | 70.5 | -0.6
6α-OH-CBD | Normal | 0.41 | LQC | 5 | 71.3 | 0.5
6α-OH-CBD | Normal | 0.41 | LQC | 6 | 70.8 | -0.2
6α-OH-CBD | Normal | 0.41 | HQC | 1 | 70.5 | -0.6
6α-OH-CBD | Normal | 0.41 | HQC | 2 | 70.8 | -0.2
6α-OH-CBD | Normal | 0.41 | HQC | 3 | 70.6 | -0.4
6α-OH-CBD | Normal | 0.41 | HQC | 4 | 70.3 | -0.9
6α-OH-CBD | Normal | 0.41 | HQC | 5 | 70.5 | -0.6
6α-OH-CBD | Normal | 0.41 | HQC | 6 | 70.7 | -0.3
6α-OH-CBD | High | 0.55 | LQC | 1 | 62.2 | -12.2
6α-OH-CBD | High | 0.55 | LQC | 2 | 62.9 | -11.3
6α-OH-CBD | High | 0.55 | LQC | 3 | 62.7 | -11.6
6α-OH-CBD | High | 0.55 | LQC | 4 | 63.1 | -11
6α-OH-CBD | High | 0.55 | LQC | 5 | 62.4 | -12
6α-OH-CBD | High | 0.55 | LQC | 6 | 63.2 | -10.9
6α-OH-CBD | High | 0.55 | HQC | 1 | 62.1 | -12.4
6α-OH-CBD | High | 0.55 | HQC | 2 | 62.5 | -11.9
6α-OH-CBD | High | 0.55 | HQC | 3 | 62.7 | -11.5
6α-OH-CBD | High | 0.55 | HQC | 4 | 62.4 | -12
6α-OH-CBD | High | 0.55 | HQC | 5 | 62.1 | -12.4
6α-OH-CBD | High | 0.55 | HQC | 6 | 62.8 | -11.4
CBN | Low | 0.26 | LQC | 1 | 90.1 | 11.1
CBN | Low | 0.26 | LQC | 2 | 90.9 | 12
CBN | Low | 0.26 | LQC | 3 | 91.4 | 12.7
CBN | Low | 0.26 | LQC | 4 | 90.1 | 11
CBN | Low | 0.26 | LQC | 5 | 90.2 | 11.2
CBN | Low | 0.26 | LQC | 6 | 90.7 | 11.7
CBN | Low | 0.26 | HQC | 1 | 90.7 | 11.7
CBN | Low | 0.26 | HQC | 2 | 90.9 | 12
CBN | Low | 0.26 | HQC | 3 | 90.9 | 12
CBN | Low | 0.26 | HQC | 4 | 90.8 | 11.8
CBN | Low | 0.26 | HQC | 5 | 90.2 | 11.1
CBN | Low | 0.26 | HQC | 6 | 90.7 | 11.8
CBN | Normal | 0.41 | LQC | 1 | 80.5 | -0.8
CBN | Normal | 0.41 | LQC | 2 | 81.6 | 0.6
CBN | Normal | 0.41 | LQC | 3 | 81.1 | -0.1
CBN | Normal | 0.41 | LQC | 4 | 80.6 | -0.7
CBN | Normal | 0.41 | LQC | 5 | 81.6 | 0.6
CBN | Normal | 0.41 | LQC | 6 | 81.2 | 0.1
CBN | Normal | 0.41 | HQC | 1 | 81.3 | 0.2
CBN | Normal | 0.41 | HQC | 2 | 80.5 | -0.8
CBN | Normal | 0.41 | HQC | 3 | 81.3 | 0.2
CBN | Normal | 0.41 | HQC | 4 | 81.1 | -0.1
CBN | Normal | 0.41 | HQC | 5 | 81.7 | 0.7
CBN | Normal | 0.41 | HQC | 6 | 80.7 | -0.5
CBN | High | 0.55 | LQC | 1 | 75.2 | -7.3
CBN | High | 0.55 | LQC | 2 | 75.8 | -6.6
CBN | High | 0.55 | LQC | 3 | 74.8 | -7.8
CBN | High | 0.55 | LQC | 4 | 75.8 | -6.6
CBN | High | 0.55 | LQC | 5 | 75.8 | -6.6
CBN | High | 0.55 | LQC | 6 | 76 | -6.3
CBN | High | 0.55 | HQC | 1 | 75 | -7.6
CBN | High | 0.55 | HQC | 2 | 74.8 | -7.8
CBN | High | 0.55 | HQC | 3 | 75.5 | -6.9
CBN | High | 0.55 | HQC | 4 | 76 | -6.3
CBN | High | 0.55 | HQC | 5 | 74.9 | -7.7
CBN | High | 0.55 | HQC | 6 | 75.6 | -6.8
CBG | Low | 0.26 | LQC | 1 | 82.4 | 10.8
CBG | Low | 0.26 | LQC | 2 | 81.7 | 9.9
CBG | Low | 0.26 | LQC | 3 | 81.8 | 10.1
CBG | Low | 0.26 | LQC | 4 | 82.4 | 10.8
CBG | Low | 0.26 | LQC | 5 | 81.6 | 9.8
CBG | Low | 0.26 | LQC | 6 | 81.2 | 9.3
CBG | Low | 0.26 | HQC | 1 | 81.7 | 9.9
CBG | Low | 0.26 | HQC | 2 | 82 | 10.3
CBG | Low | 0.26 | HQC | 3 | 81.7 | 10
CBG | Low | 0.26 | HQC | 4 | 81.7 | 9.9
CBG | Low | 0.26 | HQC | 5 | 81 | 9
CBG | Low | 0.26 | HQC | 6 | 81.6 | 9.8
CBG | Normal | 0.41 | LQC | 1 | 74 | -0.4
CBG | Normal | 0.41 | LQC | 2 | 73.7 | -0.9
CBG | Normal | 0.41 | LQC | 3 | 74.3 | -0.1
CBG | Normal | 0.41 | LQC | 4 | 74.6 | 0.3
CBG | Normal | 0.41 | LQC | 5 | 73.8 | -0.7
CBG | Normal | 0.41 | LQC | 6 | 74.6 | 0.3
CBG | Normal | 0.41 | HQC | 1 | 74.9 | 0.8
CBG | Normal | 0.41 | HQC | 2 | 73.7 | -0.8
CBG | Normal | 0.41 | HQC | 3 | 75 | 0.9
CBG | Normal | 0.41 | HQC | 4 | 74.6 | 0.4
CBG | Normal | 0.41 | HQC | 5 | 74.4 | 0.1
CBG | Normal | 0.41 | HQC | 6 | 74.1 | -0.4
CBG | High | 0.55 | LQC | 1 | 67.7 | -8.9
CBG | High | 0.55 | LQC | 2 | 68.8 | -7.4
CBG | High | 0.55 | LQC | 3 | 67.6 | -9
CBG | High | 0.55 | LQC | 4 | 67.6 | -9
CBG | High | 0.55 | LQC | 5 | 68.3 | -8.1
CBG | High | 0.55 | LQC | 6 | 67.6 | -9
CBG | High | 0.55 | HQC | 1 | 68.4 | -8.1
CBG | High | 0.55 | HQC | 2 | 68.8 | -7.5
CBG | High | 0.55 | HQC | 3 | 67.9 | -8.7
CBG | High | 0.55 | HQC | 4 | 67.4 | -9.3
CBG | High | 0.55 | HQC | 5 | 68.3 | -8.1
CBG | High | 0.55 | HQC | 6 | 67.5 | -9.2

## Stability
Stability — All Conditions (6 Replicates per Analyte per QC Level)
Condition | Analyte | QC Level | Nominal (ng/mL) | Replicate | Analyte Area (cts) | IS Area (cts) | Calculated (ng/mL) | % of Reference
RT 30 days | CBD | LQC | 6 | 1 | 1512 | 47827 | 5.662 | 94.4
RT 30 days | CBD | LQC | 6 | 2 | 1503 | 47707 | 5.72 | 95.3
RT 30 days | CBD | LQC | 6 | 3 | 1498 | 48957 | 5.723 | 95.4
RT 30 days | CBD | LQC | 6 | 4 | 1482 | 47421 | 5.665 | 94.4
RT 30 days | CBD | LQC | 6 | 5 | 1444 | 45612 | 5.648 | 94.1
RT 30 days | CBD | LQC | 6 | 6 | 1487 | 46735 | 5.602 | 93.4
RT 30 days | CBD | HQC | 160 | 1 | 40340 | 46237 | 148.876 | 93
RT 30 days | CBD | HQC | 160 | 2 | 38558 | 48084 | 149.901 | 93.7
RT 30 days | CBD | HQC | 160 | 3 | 38630 | 48138 | 149.422 | 93.4
RT 30 days | CBD | HQC | 160 | 4 | 39603 | 47250 | 150.804 | 94.3
RT 30 days | CBD | HQC | 160 | 5 | 39415 | 50311 | 149.336 | 93.3
RT 30 days | CBD | HQC | 160 | 6 | 38381 | 49026 | 148.651 | 92.9
RT 30 days | 7-OH-CBD | LQC | 6 | 1 | 1320 | 41758 | 5.574 | 92.9
RT 30 days | 7-OH-CBD | LQC | 6 | 2 | 1356 | 44844 | 5.573 | 92.9
RT 30 days | 7-OH-CBD | LQC | 6 | 3 | 1312 | 43655 | 5.556 | 92.6
RT 30 days | 7-OH-CBD | LQC | 6 | 4 | 1365 | 43473 | 5.522 | 92
RT 30 days | 7-OH-CBD | LQC | 6 | 5 | 1362 | 41983 | 5.581 | 93
RT 30 days | 7-OH-CBD | LQC | 6 | 6 | 1359 | 43225 | 5.541 | 92.3
RT 30 days | 7-OH-CBD | HQC | 160 | 1 | 35822 | 45065 | 149.849 | 93.7
RT 30 days | 7-OH-CBD | HQC | 160 | 2 | 36520 | 44378 | 147.861 | 92.4
RT 30 days | 7-OH-CBD | HQC | 160 | 3 | 35115 | 43787 | 147.609 | 92.3
RT 30 days | 7-OH-CBD | HQC | 160 | 4 | 36438 | 44929 | 146.546 | 91.6
RT 30 days | 7-OH-CBD | HQC | 160 | 5 | 36209 | 45148 | 148.213 | 92.6
RT 30 days | 7-OH-CBD | HQC | 160 | 6 | 36759 | 44413 | 147.363 | 92.1
RT 30 days | 7-COOH-CBD | LQC | 15 | 1 | 3831 | 54115 | 13.975 | 93.2
RT 30 days | 7-COOH-CBD | LQC | 15 | 2 | 4033 | 52117 | 14.095 | 94
RT 30 days | 7-COOH-CBD | LQC | 15 | 3 | 4010 | 49312 | 13.993 | 93.3
RT 30 days | 7-COOH-CBD | LQC | 15 | 4 | 3998 | 53998 | 14.03 | 93.5
RT 30 days | 7-COOH-CBD | LQC | 15 | 5 | 4055 | 51321 | 13.867 | 92.4
RT 30 days | 7-COOH-CBD | LQC | 15 | 6 | 4072 | 51283 | 14.026 | 93.5
RT 30 days | 7-COOH-CBD | HQC | 400 | 1 | 106007 | 53309 | 373.162 | 93.3
RT 30 days | 7-COOH-CBD | HQC | 400 | 2 | 104092 | 51469 | 371.578 | 92.9
RT 30 days | 7-COOH-CBD | HQC | 400 | 3 | 107588 | 52352 | 373.904 | 93.5
RT 30 days | 7-COOH-CBD | HQC | 400 | 4 | 106198 | 50086 | 371.925 | 93
RT 30 days | 7-COOH-CBD | HQC | 400 | 5 | 106582 | 50915 | 371.522 | 92.9
RT 30 days | 7-COOH-CBD | HQC | 400 | 6 | 108032 | 54550 | 374.306 | 93.6
RT 30 days | 6α-OH-CBD | LQC | 1.5 | 1 | 277 | 37926 | 1.363 | 90.8
RT 30 days | 6α-OH-CBD | LQC | 1.5 | 2 | 299 | 38059 | 1.359 | 90.6
RT 30 days | 6α-OH-CBD | LQC | 1.5 | 3 | 285 | 38297 | 1.356 | 90.4
RT 30 days | 6α-OH-CBD | LQC | 1.5 | 4 | 287 | 39129 | 1.343 | 89.5
RT 30 days | 6α-OH-CBD | LQC | 1.5 | 5 | 296 | 37101 | 1.37 | 91.3
RT 30 days | 6α-OH-CBD | LQC | 1.5 | 6 | 293 | 38579 | 1.34 | 89.3
RT 30 days | 6α-OH-CBD | HQC | 80 | 1 | 15160 | 37440 | 71.762 | 89.7
RT 30 days | 6α-OH-CBD | HQC | 80 | 2 | 15226 | 36467 | 72.115 | 90.1
RT 30 days | 6α-OH-CBD | HQC | 80 | 3 | 14872 | 38467 | 71.71 | 89.6
RT 30 days | 6α-OH-CBD | HQC | 80 | 4 | 15431 | 37226 | 73.603 | 92
RT 30 days | 6α-OH-CBD | HQC | 80 | 5 | 15472 | 36783 | 73.07 | 91.3
RT 30 days | 6α-OH-CBD | HQC | 80 | 6 | 15535 | 36411 | 73.427 | 91.8
RT 30 days | CBN | LQC | 1.5 | 1 | 307 | 40813 | 1.411 | 94
RT 30 days | CBN | LQC | 1.5 | 2 | 310 | 40542 | 1.388 | 92.6
RT 30 days | CBN | LQC | 1.5 | 3 | 310 | 40360 | 1.411 | 94
RT 30 days | CBN | LQC | 1.5 | 4 | 309 | 40304 | 1.41 | 94
RT 30 days | CBN | LQC | 1.5 | 5 | 312 | 40492 | 1.433 | 95.6
RT 30 days | CBN | LQC | 1.5 | 6 | 319 | 38986 | 1.416 | 94.4
RT 30 days | CBN | HQC | 80 | 1 | 17193 | 39146 | 75.909 | 94.9
RT 30 days | CBN | HQC | 80 | 2 | 17348 | 40854 | 75.35 | 94.2
RT 30 days | CBN | HQC | 80 | 3 | 16399 | 39216 | 74.856 | 93.6
RT 30 days | CBN | HQC | 80 | 4 | 16672 | 39919 | 74.428 | 93
RT 30 days | CBN | HQC | 80 | 5 | 16757 | 40860 | 75.248 | 94.1
RT 30 days | CBN | HQC | 80 | 6 | 17779 | 39724 | 76.131 | 95.2
RT 30 days | CBG | LQC | 1.5 | 1 | 255 | 34219 | 1.334 | 89
RT 30 days | CBG | LQC | 1.5 | 2 | 263 | 34796 | 1.332 | 88.8
RT 30 days | CBG | LQC | 1.5 | 3 | 256 | 34933 | 1.332 | 88.8
RT 30 days | CBG | LQC | 1.5 | 4 | 259 | 34593 | 1.306 | 87.1
RT 30 days | CBG | LQC | 1.5 | 5 | 260 | 33866 | 1.329 | 88.6
RT 30 days | CBG | LQC | 1.5 | 6 | 260 | 34195 | 1.331 | 88.7
RT 30 days | CBG | HQC | 80 | 1 | 13914 | 33959 | 70.761 | 88.5
RT 30 days | CBG | HQC | 80 | 2 | 14158 | 35985 | 71.117 | 88.9
RT 30 days | CBG | HQC | 80 | 3 | 13753 | 34146 | 69.746 | 87.2
RT 30 days | CBG | HQC | 80 | 4 | 14851 | 35468 | 71.083 | 88.9
RT 30 days | CBG | HQC | 80 | 5 | 14135 | 36033 | 71.124 | 88.9
RT 30 days | CBG | HQC | 80 | 6 | 14097 | 34713 | 71.181 | 89
RT 30 days | Δ9-THC | LQC | 1.5 | 1 | 342 | 44722 | 1.372 | 91.5
RT 30 days | Δ9-THC | LQC | 1.5 | 2 | 337 | 47708 | 1.384 | 92.3
RT 30 days | Δ9-THC | LQC | 1.5 | 3 | 347 | 43113 | 1.392 | 92.8
RT 30 days | Δ9-THC | LQC | 1.5 | 4 | 345 | 45075 | 1.398 | 93.2
RT 30 days | Δ9-THC | LQC | 1.5 | 5 | 341 | 45795 | 1.373 | 91.5
RT 30 days | Δ9-THC | LQC | 1.5 | 6 | 345 | 45053 | 1.386 | 92.4
RT 30 days | Δ9-THC | HQC | 80 | 1 | 18290 | 45741 | 75.341 | 94.2
RT 30 days | Δ9-THC | HQC | 80 | 2 | 18904 | 44225 | 73.898 | 92.4
RT 30 days | Δ9-THC | HQC | 80 | 3 | 18604 | 44865 | 73.175 | 91.5
RT 30 days | Δ9-THC | HQC | 80 | 4 | 18632 | 43595 | 74.399 | 93
RT 30 days | Δ9-THC | HQC | 80 | 5 | 19083 | 44235 | 73.41 | 91.8
RT 30 days | Δ9-THC | HQC | 80 | 6 | 18967 | 43524 | 74.059 | 92.6
RT 30 days | 11-OH-THC | LQC | 1.5 | 1 | 321 | 43127 | 1.403 | 93.5
RT 30 days | 11-OH-THC | LQC | 1.5 | 2 | 318 | 43383 | 1.367 | 91.2
RT 30 days | 11-OH-THC | LQC | 1.5 | 3 | 320 | 42845 | 1.38 | 92
RT 30 days | 11-OH-THC | LQC | 1.5 | 4 | 317 | 38460 | 1.373 | 91.5
RT 30 days | 11-OH-THC | LQC | 1.5 | 5 | 332 | 40637 | 1.371 | 91.4
RT 30 days | 11-OH-THC | LQC | 1.5 | 6 | 320 | 41519 | 1.354 | 90.3
RT 30 days | 11-OH-THC | HQC | 80 | 1 | 17899 | 40432 | 74.336 | 92.9
RT 30 days | 11-OH-THC | HQC | 80 | 2 | 17428 | 40928 | 73.287 | 91.6
RT 30 days | 11-OH-THC | HQC | 80 | 3 | 17339 | 41344 | 73.426 | 91.8
RT 30 days | 11-OH-THC | HQC | 80 | 4 | 17421 | 42123 | 72.998 | 91.2
RT 30 days | 11-OH-THC | HQC | 80 | 5 | 18231 | 40648 | 74.442 | 93.1
RT 30 days | 11-OH-THC | HQC | 80 | 6 | 17125 | 41834 | 71.85 | 89.8
RT 30 days | THC-COOH | LQC | 6 | 1 | 1600 | 49190 | 5.704 | 95.1
RT 30 days | THC-COOH | LQC | 6 | 2 | 1618 | 49233 | 5.66 | 94.3
RT 30 days | THC-COOH | LQC | 6 | 3 | 1633 | 49099 | 5.708 | 95.1
RT 30 days | THC-COOH | LQC | 6 | 4 | 1624 | 51056 | 5.705 | 95.1
RT 30 days | THC-COOH | LQC | 6 | 5 | 1609 | 49271 | 5.641 | 94
RT 30 days | THC-COOH | LQC | 6 | 6 | 1614 | 49972 | 5.723 | 95.4
RT 30 days | THC-COOH | HQC | 160 | 1 | 43684 | 49230 | 153.618 | 96
RT 30 days | THC-COOH | HQC | 160 | 2 | 43469 | 49717 | 151.373 | 94.6
RT 30 days | THC-COOH | HQC | 160 | 3 | 44077 | 51057 | 152.932 | 95.6
RT 30 days | THC-COOH | HQC | 160 | 4 | 42483 | 50886 | 151.712 | 94.8
RT 30 days | THC-COOH | HQC | 160 | 5 | 42926 | 49830 | 152.172 | 95.1
RT 30 days | THC-COOH | HQC | 160 | 6 | 44585 | 48936 | 152.583 | 95.4
RT 60 days | CBD | LQC | 6 | 1 | 1445 | 46484 | 5.363 | 89.4
RT 60 days | CBD | LQC | 6 | 2 | 1457 | 48778 | 5.461 | 91
RT 60 days | CBD | LQC | 6 | 3 | 1432 | 46880 | 5.417 | 90.3
RT 60 days | CBD | LQC | 6 | 4 | 1416 | 49216 | 5.381 | 89.7
RT 60 days | CBD | LQC | 6 | 5 | 1429 | 47598 | 5.373 | 89.6
RT 60 days | CBD | LQC | 6 | 6 | 1428 | 46973 | 5.387 | 89.8
RT 60 days | CBD | HQC | 160 | 1 | 37911 | 47367 | 144.834 | 90.5
RT 60 days | CBD | HQC | 160 | 2 | 37262 | 46192 | 144.815 | 90.5
RT 60 days | CBD | HQC | 160 | 3 | 38080 | 47733 | 143.442 | 89.7
RT 60 days | CBD | HQC | 160 | 4 | 38594 | 49070 | 146.612 | 91.6
RT 60 days | CBD | HQC | 160 | 5 | 38122 | 46748 | 144.453 | 90.3
RT 60 days | CBD | HQC | 160 | 6 | 37466 | 48402 | 144.063 | 90
RT 60 days | 7-OH-CBD | LQC | 6 | 1 | 1294 | 45753 | 5.344 | 89.1
RT 60 days | 7-OH-CBD | LQC | 6 | 2 | 1286 | 44789 | 5.337 | 89
RT 60 days | 7-OH-CBD | LQC | 6 | 3 | 1295 | 44014 | 5.328 | 88.8
RT 60 days | 7-OH-CBD | LQC | 6 | 4 | 1291 | 44028 | 5.275 | 87.9
RT 60 days | 7-OH-CBD | LQC | 6 | 5 | 1359 | 45097 | 5.393 | 89.9
RT 60 days | 7-OH-CBD | LQC | 6 | 6 | 1284 | 44997 | 5.275 | 87.9
RT 60 days | 7-OH-CBD | HQC | 160 | 1 | 34630 | 43075 | 144.864 | 90.5
RT 60 days | 7-OH-CBD | HQC | 160 | 2 | 33878 | 41570 | 140.91 | 88.1
RT 60 days | 7-OH-CBD | HQC | 160 | 3 | 34357 | 44172 | 140.684 | 87.9
RT 60 days | 7-OH-CBD | HQC | 160 | 4 | 34690 | 45087 | 144.42 | 90.3
RT 60 days | 7-OH-CBD | HQC | 160 | 5 | 33289 | 44562 | 140.486 | 87.8
RT 60 days | 7-OH-CBD | HQC | 160 | 6 | 34743 | 45349 | 144.356 | 90.2
RT 60 days | 7-COOH-CBD | LQC | 15 | 1 | 3801 | 53830 | 13.239 | 88.3
RT 60 days | 7-COOH-CBD | LQC | 15 | 2 | 3881 | 49844 | 13.637 | 90.9
RT 60 days | 7-COOH-CBD | LQC | 15 | 3 | 3844 | 52063 | 13.471 | 89.8
RT 60 days | 7-COOH-CBD | LQC | 15 | 4 | 3889 | 50556 | 13.554 | 90.4
RT 60 days | 7-COOH-CBD | LQC | 15 | 5 | 3929 | 52695 | 13.486 | 89.9
RT 60 days | 7-COOH-CBD | LQC | 15 | 6 | 3920 | 49925 | 13.319 | 88.8
RT 60 days | 7-COOH-CBD | HQC | 400 | 1 | 106586 | 50988 | 360.394 | 90.1
RT 60 days | 7-COOH-CBD | HQC | 400 | 2 | 103592 | 53594 | 359.616 | 89.9
RT 60 days | 7-COOH-CBD | HQC | 400 | 3 | 105004 | 54373 | 366.602 | 91.7
RT 60 days | 7-COOH-CBD | HQC | 400 | 4 | 106922 | 52872 | 362.69 | 90.7
RT 60 days | 7-COOH-CBD | HQC | 400 | 5 | 104300 | 52927 | 355.734 | 88.9
RT 60 days | 7-COOH-CBD | HQC | 400 | 6 | 97803 | 50400 | 352.883 | 88.2
RT 60 days | 6α-OH-CBD | LQC | 1.5 | 1 | 282 | 38708 | 1.304 | 87
RT 60 days | 6α-OH-CBD | LQC | 1.5 | 2 | 269 | 39609 | 1.309 | 87.3
RT 60 days | 6α-OH-CBD | LQC | 1.5 | 3 | 273 | 37945 | 1.286 | 85.7
RT 60 days | 6α-OH-CBD | LQC | 1.5 | 4 | 275 | 35957 | 1.308 | 87.2
RT 60 days | 6α-OH-CBD | LQC | 1.5 | 5 | 274 | 38661 | 1.291 | 86.1
RT 60 days | 6α-OH-CBD | LQC | 1.5 | 6 | 268 | 37492 | 1.296 | 86.4
RT 60 days | 6α-OH-CBD | HQC | 80 | 1 | 14440 | 38943 | 69.716 | 87.1
RT 60 days | 6α-OH-CBD | HQC | 80 | 2 | 14574 | 37476 | 70.112 | 87.6
RT 60 days | 6α-OH-CBD | HQC | 80 | 3 | 14548 | 36977 | 69.685 | 87.1
RT 60 days | 6α-OH-CBD | HQC | 80 | 4 | 14544 | 39941 | 68.924 | 86.2
RT 60 days | 6α-OH-CBD | HQC | 80 | 5 | 14578 | 38211 | 69.272 | 86.6
RT 60 days | 6α-OH-CBD | HQC | 80 | 6 | 14878 | 38606 | 69.606 | 87
RT 60 days | CBN | LQC | 1.5 | 1 | 298 | 39401 | 1.349 | 89.9
RT 60 days | CBN | LQC | 1.5 | 2 | 302 | 38424 | 1.325 | 88.3
RT 60 days | CBN | LQC | 1.5 | 3 | 310 | 39741 | 1.378 | 91.9
RT 60 days | CBN | LQC | 1.5 | 4 | 308 | 43202 | 1.36 | 90.7
RT 60 days | CBN | LQC | 1.5 | 5 | 292 | 39006 | 1.354 | 90.3
RT 60 days | CBN | LQC | 1.5 | 6 | 294 | 39213 | 1.359 | 90.6
RT 60 days | CBN | HQC | 80 | 1 | 16483 | 38875 | 71.852 | 89.8
RT 60 days | CBN | HQC | 80 | 2 | 16675 | 39991 | 72.958 | 91.2
RT 60 days | CBN | HQC | 80 | 3 | 16555 | 39104 | 72.376 | 90.5
RT 60 days | CBN | HQC | 80 | 4 | 16132 | 38921 | 72.71 | 90.9
RT 60 days | CBN | HQC | 80 | 5 | 16226 | 38561 | 71.687 | 89.6
RT 60 days | CBN | HQC | 80 | 6 | 16575 | 38545 | 73.701 | 92.1
RT 60 days | CBG | LQC | 1.5 | 1 | 249 | 34556 | 1.291 | 86
RT 60 days | CBG | LQC | 1.5 | 2 | 247 | 35044 | 1.288 | 85.9
RT 60 days | CBG | LQC | 1.5 | 3 | 253 | 34954 | 1.303 | 86.8
RT 60 days | CBG | LQC | 1.5 | 4 | 254 | 34475 | 1.287 | 85.8
RT 60 days | CBG | LQC | 1.5 | 5 | 255 | 34110 | 1.306 | 87
RT 60 days | CBG | LQC | 1.5 | 6 | 258 | 33480 | 1.309 | 87.3
RT 60 days | CBG | HQC | 80 | 1 | 13262 | 35519 | 68.821 | 86
RT 60 days | CBG | HQC | 80 | 2 | 13391 | 33518 | 68.48 | 85.6
RT 60 days | CBG | HQC | 80 | 3 | 13727 | 34177 | 69.46 | 86.8
RT 60 days | CBG | HQC | 80 | 4 | 13088 | 34939 | 68.48 | 85.6
RT 60 days | CBG | HQC | 80 | 5 | 13239 | 35043 | 68.828 | 86
RT 60 days | CBG | HQC | 80 | 6 | 13344 | 35915 | 69.014 | 86.3
RT 60 days | Δ9-THC | LQC | 1.5 | 1 | 341 | 43678 | 1.343 | 89.5
RT 60 days | Δ9-THC | LQC | 1.5 | 2 | 335 | 42108 | 1.351 | 90.1
RT 60 days | Δ9-THC | LQC | 1.5 | 3 | 339 | 44762 | 1.34 | 89.3
RT 60 days | Δ9-THC | LQC | 1.5 | 4 | 332 | 45101 | 1.324 | 88.3
RT 60 days | Δ9-THC | LQC | 1.5 | 5 | 341 | 45297 | 1.348 | 89.8
RT 60 days | Δ9-THC | LQC | 1.5 | 6 | 331 | 44429 | 1.327 | 88.4
RT 60 days | Δ9-THC | HQC | 80 | 1 | 18063 | 44507 | 71.32 | 89.2
RT 60 days | Δ9-THC | HQC | 80 | 2 | 18163 | 46019 | 72.514 | 90.6
RT 60 days | Δ9-THC | HQC | 80 | 3 | 17449 | 44522 | 71.893 | 89.9
RT 60 days | Δ9-THC | HQC | 80 | 4 | 17606 | 42811 | 70.324 | 87.9
RT 60 days | Δ9-THC | HQC | 80 | 5 | 18007 | 46961 | 71.053 | 88.8
RT 60 days | Δ9-THC | HQC | 80 | 6 | 17148 | 45970 | 70.884 | 88.6
RT 60 days | 11-OH-THC | LQC | 1.5 | 1 | 301 | 42841 | 1.314 | 87.6
RT 60 days | 11-OH-THC | LQC | 1.5 | 2 | 310 | 42369 | 1.326 | 88.4
RT 60 days | 11-OH-THC | LQC | 1.5 | 3 | 312 | 44479 | 1.319 | 88
RT 60 days | 11-OH-THC | LQC | 1.5 | 4 | 305 | 41498 | 1.314 | 87.6
RT 60 days | 11-OH-THC | LQC | 1.5 | 5 | 312 | 41065 | 1.322 | 88.1
RT 60 days | 11-OH-THC | LQC | 1.5 | 6 | 309 | 42523 | 1.305 | 87
RT 60 days | 11-OH-THC | HQC | 80 | 1 | 16580 | 41443 | 70.645 | 88.3
RT 60 days | 11-OH-THC | HQC | 80 | 2 | 16440 | 41514 | 69.236 | 86.5
RT 60 days | 11-OH-THC | HQC | 80 | 3 | 16875 | 40992 | 70.449 | 88.1
RT 60 days | 11-OH-THC | HQC | 80 | 4 | 17278 | 42451 | 70.868 | 88.6
RT 60 days | 11-OH-THC | HQC | 80 | 5 | 15839 | 40640 | 69.379 | 86.7
RT 60 days | 11-OH-THC | HQC | 80 | 6 | 16720 | 42714 | 70.966 | 88.7
RT 60 days | THC-COOH | LQC | 6 | 1 | 1620 | 48536 | 5.511 | 91.9
RT 60 days | THC-COOH | LQC | 6 | 2 | 1569 | 50142 | 5.503 | 91.7
RT 60 days | THC-COOH | LQC | 6 | 3 | 1532 | 50291 | 5.522 | 92
RT 60 days | THC-COOH | LQC | 6 | 4 | 1592 | 52446 | 5.522 | 92
RT 60 days | THC-COOH | LQC | 6 | 5 | 1544 | 49335 | 5.584 | 93.1
RT 60 days | THC-COOH | LQC | 6 | 6 | 1576 | 50072 | 5.489 | 91.5
RT 60 days | THC-COOH | HQC | 160 | 1 | 40457 | 51988 | 144.08 | 90
RT 60 days | THC-COOH | HQC | 160 | 2 | 39870 | 48684 | 145.934 | 91.2
RT 60 days | THC-COOH | HQC | 160 | 3 | 40644 | 50095 | 147.766 | 92.4
RT 60 days | THC-COOH | HQC | 160 | 4 | 41960 | 49563 | 144.662 | 90.4
RT 60 days | THC-COOH | HQC | 160 | 5 | 40982 | 48045 | 146.18 | 91.4
RT 60 days | THC-COOH | HQC | 160 | 6 | 41118 | 46494 | 146.16 | 91.4
RT 90 days | CBD | LQC | 6 | 1 | 1430 | 48868 | 5.244 | 87.4
RT 90 days | CBD | LQC | 6 | 2 | 1375 | 47664 | 5.315 | 88.6
RT 90 days | CBD | LQC | 6 | 3 | 1423 | 48073 | 5.391 | 89.9
RT 90 days | CBD | LQC | 6 | 4 | 1383 | 48247 | 5.255 | 87.6
RT 90 days | CBD | LQC | 6 | 5 | 1378 | 47317 | 5.226 | 87.1
RT 90 days | CBD | LQC | 6 | 6 | 1382 | 47110 | 5.227 | 87.1
RT 90 days | CBD | HQC | 160 | 1 | 35146 | 49876 | 139.803 | 87.4
RT 90 days | CBD | HQC | 160 | 2 | 36010 | 49555 | 141.69 | 88.6
RT 90 days | CBD | HQC | 160 | 3 | 35941 | 47536 | 139.68 | 87.3
RT 90 days | CBD | HQC | 160 | 4 | 37872 | 46614 | 139.07 | 86.9
RT 90 days | CBD | HQC | 160 | 5 | 37833 | 49586 | 141.471 | 88.4
RT 90 days | CBD | HQC | 160 | 6 | 36906 | 47345 | 138.578 | 86.6
RT 90 days | 7-OH-CBD | LQC | 6 | 1 | 1244 | 44815 | 5.15 | 85.8
RT 90 days | 7-OH-CBD | LQC | 6 | 2 | 1270 | 44813 | 5.194 | 86.6
RT 90 days | 7-OH-CBD | LQC | 6 | 3 | 1236 | 45326 | 5.195 | 86.6
RT 90 days | 7-OH-CBD | LQC | 6 | 4 | 1266 | 44678 | 5.246 | 87.4
RT 90 days | 7-OH-CBD | LQC | 6 | 5 | 1290 | 42568 | 5.231 | 87.2
RT 90 days | 7-OH-CBD | LQC | 6 | 6 | 1287 | 43402 | 5.204 | 86.7
RT 90 days | 7-OH-CBD | HQC | 160 | 1 | 32571 | 42388 | 138.061 | 86.3
RT 90 days | 7-OH-CBD | HQC | 160 | 2 | 35025 | 42535 | 139.883 | 87.4
RT 90 days | 7-OH-CBD | HQC | 160 | 3 | 33370 | 45940 | 136.96 | 85.6
RT 90 days | 7-OH-CBD | HQC | 160 | 4 | 33958 | 43376 | 138.906 | 86.8
RT 90 days | 7-OH-CBD | HQC | 160 | 5 | 33805 | 45262 | 138.988 | 86.9
RT 90 days | 7-OH-CBD | HQC | 160 | 6 | 33764 | 43584 | 139.249 | 87
RT 90 days | 7-COOH-CBD | LQC | 15 | 1 | 3687 | 49687 | 13.135 | 87.6
RT 90 days | 7-COOH-CBD | LQC | 15 | 2 | 3783 | 52231 | 13.192 | 87.9
RT 90 days | 7-COOH-CBD | LQC | 15 | 3 | 3731 | 52469 | 13.097 | 87.3
RT 90 days | 7-COOH-CBD | LQC | 15 | 4 | 3672 | 52264 | 12.993 | 86.6
RT 90 days | 7-COOH-CBD | LQC | 15 | 5 | 3853 | 49698 | 13.148 | 87.7
RT 90 days | 7-COOH-CBD | LQC | 15 | 6 | 3852 | 52346 | 13.157 | 87.7
RT 90 days | 7-COOH-CBD | HQC | 400 | 1 | 98759 | 53371 | 355.016 | 88.8
RT 90 days | 7-COOH-CBD | HQC | 400 | 2 | 99231 | 51155 | 344.732 | 86.2
RT 90 days | 7-COOH-CBD | HQC | 400 | 3 | 101330 | 51018 | 351.008 | 87.8
RT 90 days | 7-COOH-CBD | HQC | 400 | 4 | 97186 | 51173 | 346.562 | 86.6
RT 90 days | 7-COOH-CBD | HQC | 400 | 5 | 97794 | 50172 | 345.976 | 86.5
RT 90 days | 7-COOH-CBD | HQC | 400 | 6 | 106339 | 50716 | 352.879 | 88.2
RT 90 days | 6α-OH-CBD | LQC | 1.5 | 1 | 268 | 38182 | 1.292 | 86.1
RT 90 days | 6α-OH-CBD | LQC | 1.5 | 2 | 273 | 37431 | 1.295 | 86.3
RT 90 days | 6α-OH-CBD | LQC | 1.5 | 3 | 283 | 37905 | 1.318 | 87.9
RT 90 days | 6α-OH-CBD | LQC | 1.5 | 4 | 279 | 37965 | 1.284 | 85.6
RT 90 days | 6α-OH-CBD | LQC | 1.5 | 5 | 274 | 39790 | 1.284 | 85.6
RT 90 days | 6α-OH-CBD | LQC | 1.5 | 6 | 275 | 38624 | 1.284 | 85.6
RT 90 days | 6α-OH-CBD | HQC | 80 | 1 | 14357 | 36448 | 68.986 | 86.2
RT 90 days | 6α-OH-CBD | HQC | 80 | 2 | 14602 | 37926 | 68.48 | 85.6
RT 90 days | 6α-OH-CBD | HQC | 80 | 3 | 14582 | 36680 | 68.682 | 85.9
RT 90 days | 6α-OH-CBD | HQC | 80 | 4 | 14829 | 38553 | 68.979 | 86.2
RT 90 days | 6α-OH-CBD | HQC | 80 | 5 | 14560 | 38453 | 69.132 | 86.4
RT 90 days | 6α-OH-CBD | HQC | 80 | 6 | 14495 | 38424 | 68.48 | 85.6
RT 90 days | CBN | LQC | 1.5 | 1 | 289 | 38672 | 1.325 | 88.3
RT 90 days | CBN | LQC | 1.5 | 2 | 300 | 39958 | 1.335 | 89
RT 90 days | CBN | LQC | 1.5 | 3 | 299 | 40030 | 1.322 | 88.1
RT 90 days | CBN | LQC | 1.5 | 4 | 290 | 40099 | 1.315 | 87.6
RT 90 days | CBN | LQC | 1.5 | 5 | 294 | 39678 | 1.316 | 87.7
RT 90 days | CBN | LQC | 1.5 | 6 | 288 | 41305 | 1.315 | 87.7
RT 90 days | CBN | HQC | 80 | 1 | 16070 | 39542 | 70.348 | 87.9
RT 90 days | CBN | HQC | 80 | 2 | 15302 | 41073 | 70.619 | 88.3
RT 90 days | CBN | HQC | 80 | 3 | 15756 | 40925 | 70.668 | 88.3
RT 90 days | CBN | HQC | 80 | 4 | 16261 | 39462 | 71.231 | 89
RT 90 days | CBN | HQC | 80 | 5 | 15923 | 39622 | 72.1 | 90.1
RT 90 days | CBN | HQC | 80 | 6 | 15769 | 41642 | 71.504 | 89.4
RT 90 days | CBG | LQC | 1.5 | 1 | 249 | 34743 | 1.306 | 87.1
RT 90 days | CBG | LQC | 1.5 | 2 | 249 | 34086 | 1.29 | 86
RT 90 days | CBG | LQC | 1.5 | 3 | 267 | 35213 | 1.334 | 88.9
RT 90 days | CBG | LQC | 1.5 | 4 | 248 | 35895 | 1.274 | 84.9
RT 90 days | CBG | LQC | 1.5 | 5 | 246 | 35614 | 1.287 | 85.8
RT 90 days | CBG | LQC | 1.5 | 6 | 257 | 35151 | 1.305 | 87
RT 90 days | CBG | HQC | 80 | 1 | 13419 | 37234 | 69.48 | 86.9
RT 90 days | CBG | HQC | 80 | 2 | 13329 | 34432 | 68.912 | 86.1
RT 90 days | CBG | HQC | 80 | 3 | 13794 | 36082 | 68.838 | 86
RT 90 days | CBG | HQC | 80 | 4 | 13573 | 34242 | 68.885 | 86.1
RT 90 days | CBG | HQC | 80 | 5 | 13510 | 35575 | 68.897 | 86.1
RT 90 days | CBG | HQC | 80 | 6 | 13812 | 36408 | 68.48 | 85.6
RT 90 days | Δ9-THC | LQC | 1.5 | 1 | 326 | 44546 | 1.288 | 85.9
RT 90 days | Δ9-THC | LQC | 1.5 | 2 | 339 | 45770 | 1.3 | 86.7
RT 90 days | Δ9-THC | LQC | 1.5 | 3 | 317 | 44065 | 1.294 | 86.2
RT 90 days | Δ9-THC | LQC | 1.5 | 4 | 334 | 44961 | 1.287 | 85.8
RT 90 days | Δ9-THC | LQC | 1.5 | 5 | 324 | 45261 | 1.289 | 86
RT 90 days | Δ9-THC | LQC | 1.5 | 6 | 327 | 47956 | 1.311 | 87.4
RT 90 days | Δ9-THC | HQC | 80 | 1 | 17542 | 46221 | 68.782 | 86
RT 90 days | Δ9-THC | HQC | 80 | 2 | 17622 | 45683 | 70.396 | 88
RT 90 days | Δ9-THC | HQC | 80 | 3 | 17776 | 46297 | 69.528 | 86.9
RT 90 days | Δ9-THC | HQC | 80 | 4 | 17871 | 46248 | 69.471 | 86.8
RT 90 days | Δ9-THC | HQC | 80 | 5 | 17602 | 45758 | 70.075 | 87.6
RT 90 days | Δ9-THC | HQC | 80 | 6 | 17184 | 46411 | 69.221 | 86.5
RT 90 days | 11-OH-THC | LQC | 1.5 | 1 | 302 | 41570 | 1.284 | 85.6
RT 90 days | 11-OH-THC | LQC | 1.5 | 2 | 305 | 42208 | 1.294 | 86.3
RT 90 days | 11-OH-THC | LQC | 1.5 | 3 | 306 | 42814 | 1.284 | 85.6
RT 90 days | 11-OH-THC | LQC | 1.5 | 4 | 309 | 41663 | 1.306 | 87
RT 90 days | 11-OH-THC | LQC | 1.5 | 5 | 302 | 43722 | 1.284 | 85.6
RT 90 days | 11-OH-THC | LQC | 1.5 | 6 | 303 | 42663 | 1.284 | 85.6
RT 90 days | 11-OH-THC | HQC | 80 | 1 | 16463 | 42211 | 68.48 | 85.6
RT 90 days | 11-OH-THC | HQC | 80 | 2 | 15947 | 42145 | 68.48 | 85.6
RT 90 days | 11-OH-THC | HQC | 80 | 3 | 16104 | 40336 | 68.975 | 86.2
RT 90 days | 11-OH-THC | HQC | 80 | 4 | 16131 | 43709 | 68.48 | 85.6
RT 90 days | 11-OH-THC | HQC | 80 | 5 | 15730 | 44054 | 68.48 | 85.6
RT 90 days | 11-OH-THC | HQC | 80 | 6 | 16005 | 42480 | 68.48 | 85.6
RT 90 days | THC-COOH | LQC | 6 | 1 | 1491 | 50920 | 5.274 | 87.9
RT 90 days | THC-COOH | LQC | 6 | 2 | 1515 | 48335 | 5.298 | 88.3
RT 90 days | THC-COOH | LQC | 6 | 3 | 1560 | 49871 | 5.363 | 89.4
RT 90 days | THC-COOH | LQC | 6 | 4 | 1465 | 48016 | 5.271 | 87.8
RT 90 days | THC-COOH | LQC | 6 | 5 | 1493 | 50316 | 5.326 | 88.8
RT 90 days | THC-COOH | LQC | 6 | 6 | 1544 | 48372 | 5.347 | 89.1
RT 90 days | THC-COOH | HQC | 160 | 1 | 40223 | 51452 | 141.963 | 88.7
RT 90 days | THC-COOH | HQC | 160 | 2 | 40516 | 50740 | 142.729 | 89.2
RT 90 days | THC-COOH | HQC | 160 | 3 | 40409 | 51727 | 142.993 | 89.4
RT 90 days | THC-COOH | HQC | 160 | 4 | 41337 | 49883 | 143.087 | 89.4
RT 90 days | THC-COOH | HQC | 160 | 5 | 39710 | 52341 | 141.119 | 88.2
RT 90 days | THC-COOH | HQC | 160 | 6 | 38953 | 50166 | 140.047 | 87.5
40C 7 days | CBD | LQC | 6 | 1 | 1448 | 47185 | 5.53 | 92.2
40C 7 days | CBD | LQC | 6 | 2 | 1358 | 47435 | 5.38 | 89.7
40C 7 days | CBD | LQC | 6 | 3 | 1444 | 48948 | 5.397 | 90
40C 7 days | CBD | LQC | 6 | 4 | 1452 | 46793 | 5.507 | 91.8
40C 7 days | CBD | LQC | 6 | 5 | 1486 | 46554 | 5.484 | 91.4
40C 7 days | CBD | LQC | 6 | 6 | 1461 | 47724 | 5.532 | 92.2
40C 7 days | CBD | HQC | 160 | 1 | 38515 | 45451 | 146.545 | 91.6
40C 7 days | CBD | HQC | 160 | 2 | 38150 | 46749 | 145.243 | 90.8
40C 7 days | CBD | HQC | 160 | 3 | 38897 | 48799 | 148.869 | 93
40C 7 days | CBD | HQC | 160 | 4 | 39474 | 49746 | 147.077 | 91.9
40C 7 days | CBD | HQC | 160 | 5 | 39115 | 47192 | 146.327 | 91.5
40C 7 days | CBD | HQC | 160 | 6 | 39202 | 49117 | 146.815 | 91.8
40C 7 days | 7-OH-CBD | LQC | 6 | 1 | 1357 | 43387 | 5.463 | 91
40C 7 days | 7-OH-CBD | LQC | 6 | 2 | 1298 | 46374 | 5.417 | 90.3
40C 7 days | 7-OH-CBD | LQC | 6 | 3 | 1319 | 44226 | 5.5 | 91.7
40C 7 days | 7-OH-CBD | LQC | 6 | 4 | 1315 | 43633 | 5.38 | 89.7
40C 7 days | 7-OH-CBD | LQC | 6 | 5 | 1318 | 43343 | 5.434 | 90.6
40C 7 days | 7-OH-CBD | LQC | 6 | 6 | 1350 | 45976 | 5.522 | 92
40C 7 days | 7-OH-CBD | HQC | 160 | 1 | 35391 | 44137 | 144.215 | 90.1
40C 7 days | 7-OH-CBD | HQC | 160 | 2 | 33593 | 43705 | 142.686 | 89.2
40C 7 days | 7-OH-CBD | HQC | 160 | 3 | 34825 | 44740 | 144.736 | 90.5
40C 7 days | 7-OH-CBD | HQC | 160 | 4 | 34605 | 42597 | 146.429 | 91.5
40C 7 days | 7-OH-CBD | HQC | 160 | 5 | 36132 | 45197 | 144.484 | 90.3
40C 7 days | 7-OH-CBD | HQC | 160 | 6 | 34806 | 45234 | 143.498 | 89.7
40C 7 days | 7-COOH-CBD | LQC | 15 | 1 | 3992 | 53194 | 13.716 | 91.4
40C 7 days | 7-COOH-CBD | LQC | 15 | 2 | 3987 | 53996 | 13.523 | 90.2
40C 7 days | 7-COOH-CBD | LQC | 15 | 3 | 3963 | 52535 | 13.654 | 91
40C 7 days | 7-COOH-CBD | LQC | 15 | 4 | 3915 | 51255 | 13.776 | 91.8
40C 7 days | 7-COOH-CBD | LQC | 15 | 5 | 3990 | 53585 | 13.736 | 91.6
40C 7 days | 7-COOH-CBD | LQC | 15 | 6 | 3736 | 53043 | 13.689 | 91.3
40C 7 days | 7-COOH-CBD | HQC | 400 | 1 | 105934 | 49647 | 370.609 | 92.7
40C 7 days | 7-COOH-CBD | HQC | 400 | 2 | 106416 | 50249 | 362.298 | 90.6
40C 7 days | 7-COOH-CBD | HQC | 400 | 3 | 104837 | 52468 | 364.686 | 91.2
40C 7 days | 7-COOH-CBD | HQC | 400 | 4 | 103475 | 52103 | 368.148 | 92
40C 7 days | 7-COOH-CBD | HQC | 400 | 5 | 103486 | 51332 | 370.05 | 92.5
40C 7 days | 7-COOH-CBD | HQC | 400 | 6 | 106574 | 54029 | 364.831 | 91.2
40C 7 days | 6α-OH-CBD | LQC | 1.5 | 1 | 278 | 37654 | 1.327 | 88.5
40C 7 days | 6α-OH-CBD | LQC | 1.5 | 2 | 288 | 38377 | 1.36 | 90.6
40C 7 days | 6α-OH-CBD | LQC | 1.5 | 3 | 282 | 38235 | 1.346 | 89.7
40C 7 days | 6α-OH-CBD | LQC | 1.5 | 4 | 280 | 37929 | 1.329 | 88.6
40C 7 days | 6α-OH-CBD | LQC | 1.5 | 5 | 286 | 38051 | 1.342 | 89.4
40C 7 days | 6α-OH-CBD | LQC | 1.5 | 6 | 270 | 38077 | 1.331 | 88.7
40C 7 days | 6α-OH-CBD | HQC | 80 | 1 | 15389 | 38342 | 71.691 | 89.6
40C 7 days | 6α-OH-CBD | HQC | 80 | 2 | 14510 | 40063 | 68.983 | 86.2
40C 7 days | 6α-OH-CBD | HQC | 80 | 3 | 15314 | 36972 | 71.827 | 89.8
40C 7 days | 6α-OH-CBD | HQC | 80 | 4 | 15333 | 37383 | 70.292 | 87.9
40C 7 days | 6α-OH-CBD | HQC | 80 | 5 | 15186 | 38214 | 70.905 | 88.6
40C 7 days | 6α-OH-CBD | HQC | 80 | 6 | 15154 | 38148 | 71.392 | 89.2
40C 7 days | CBN | LQC | 1.5 | 1 | 314 | 39940 | 1.409 | 93.9
40C 7 days | CBN | LQC | 1.5 | 2 | 316 | 41149 | 1.39 | 92.7
40C 7 days | CBN | LQC | 1.5 | 3 | 317 | 38811 | 1.384 | 92.3
40C 7 days | CBN | LQC | 1.5 | 4 | 298 | 40314 | 1.36 | 90.6
40C 7 days | CBN | LQC | 1.5 | 5 | 307 | 40341 | 1.375 | 91.6
40C 7 days | CBN | LQC | 1.5 | 6 | 310 | 41678 | 1.381 | 92.1
40C 7 days | CBN | HQC | 80 | 1 | 16641 | 41002 | 73.181 | 91.5
40C 7 days | CBN | HQC | 80 | 2 | 16335 | 40621 | 72.459 | 90.6
40C 7 days | CBN | HQC | 80 | 3 | 16240 | 41430 | 73.142 | 91.4
40C 7 days | CBN | HQC | 80 | 4 | 16357 | 39820 | 73.452 | 91.8
40C 7 days | CBN | HQC | 80 | 5 | 16617 | 38502 | 72.556 | 90.7
40C 7 days | CBN | HQC | 80 | 6 | 15566 | 38968 | 72.739 | 90.9
40C 7 days | CBG | LQC | 1.5 | 1 | 261 | 33795 | 1.3 | 86.6
40C 7 days | CBG | LQC | 1.5 | 2 | 268 | 34373 | 1.321 | 88.1
40C 7 days | CBG | LQC | 1.5 | 3 | 246 | 34704 | 1.321 | 88.1
40C 7 days | CBG | LQC | 1.5 | 4 | 253 | 33735 | 1.338 | 89.2
40C 7 days | CBG | LQC | 1.5 | 5 | 264 | 37220 | 1.321 | 88.1
40C 7 days | CBG | LQC | 1.5 | 6 | 260 | 35541 | 1.315 | 87.7
40C 7 days | CBG | HQC | 80 | 1 | 13759 | 35255 | 70.289 | 87.9
40C 7 days | CBG | HQC | 80 | 2 | 13552 | 34313 | 69.716 | 87.1
40C 7 days | CBG | HQC | 80 | 3 | 13161 | 34910 | 68.579 | 85.7
40C 7 days | CBG | HQC | 80 | 4 | 13817 | 35163 | 68.99 | 86.2
40C 7 days | CBG | HQC | 80 | 5 | 13510 | 33570 | 70.401 | 88
40C 7 days | CBG | HQC | 80 | 6 | 13702 | 34105 | 70.416 | 88
40C 7 days | Δ9-THC | LQC | 1.5 | 1 | 344 | 43867 | 1.366 | 91.1
40C 7 days | Δ9-THC | LQC | 1.5 | 2 | 339 | 44296 | 1.341 | 89.4
40C 7 days | Δ9-THC | LQC | 1.5 | 3 | 336 | 45515 | 1.337 | 89.2
40C 7 days | Δ9-THC | LQC | 1.5 | 4 | 352 | 45261 | 1.367 | 91.1
40C 7 days | Δ9-THC | LQC | 1.5 | 5 | 343 | 44497 | 1.361 | 90.7
40C 7 days | Δ9-THC | LQC | 1.5 | 6 | 344 | 45057 | 1.366 | 91.1
40C 7 days | Δ9-THC | HQC | 80 | 1 | 17675 | 44224 | 71.971 | 90
40C 7 days | Δ9-THC | HQC | 80 | 2 | 18349 | 43544 | 73.748 | 92.2
40C 7 days | Δ9-THC | HQC | 80 | 3 | 18034 | 46047 | 71.959 | 89.9
40C 7 days | Δ9-THC | HQC | 80 | 4 | 18582 | 44212 | 73.384 | 91.7
40C 7 days | Δ9-THC | HQC | 80 | 5 | 17722 | 47509 | 71.875 | 89.8
40C 7 days | Δ9-THC | HQC | 80 | 6 | 18454 | 45741 | 72.705 | 90.9
40C 7 days | 11-OH-THC | LQC | 1.5 | 1 | 315 | 41858 | 1.336 | 89
40C 7 days | 11-OH-THC | LQC | 1.5 | 2 | 327 | 43052 | 1.336 | 89.1
40C 7 days | 11-OH-THC | LQC | 1.5 | 3 | 313 | 40504 | 1.331 | 88.7
40C 7 days | 11-OH-THC | LQC | 1.5 | 4 | 333 | 41122 | 1.373 | 91.5
40C 7 days | 11-OH-THC | LQC | 1.5 | 5 | 324 | 41395 | 1.36 | 90.7
40C 7 days | 11-OH-THC | LQC | 1.5 | 6 | 319 | 41585 | 1.36 | 90.7
40C 7 days | 11-OH-THC | HQC | 80 | 1 | 16536 | 40891 | 70.945 | 88.7
40C 7 days | 11-OH-THC | HQC | 80 | 2 | 16393 | 42037 | 70.487 | 88.1
40C 7 days | 11-OH-THC | HQC | 80 | 3 | 16970 | 40300 | 71.527 | 89.4
40C 7 days | 11-OH-THC | HQC | 80 | 4 | 17084 | 39822 | 71.955 | 89.9
40C 7 days | 11-OH-THC | HQC | 80 | 5 | 17549 | 40679 | 72.236 | 90.3
40C 7 days | 11-OH-THC | HQC | 80 | 6 | 16883 | 41230 | 71.659 | 89.6
40C 7 days | THC-COOH | LQC | 6 | 1 | 1593 | 50098 | 5.605 | 93.4
40C 7 days | THC-COOH | LQC | 6 | 2 | 1515 | 50220 | 5.523 | 92
40C 7 days | THC-COOH | LQC | 6 | 3 | 1547 | 48824 | 5.568 | 92.8
40C 7 days | THC-COOH | LQC | 6 | 4 | 1606 | 51418 | 5.643 | 94
40C 7 days | THC-COOH | LQC | 6 | 5 | 1586 | 49588 | 5.513 | 91.9
40C 7 days | THC-COOH | LQC | 6 | 6 | 1582 | 53015 | 5.582 | 93
40C 7 days | THC-COOH | HQC | 160 | 1 | 42038 | 49598 | 148.925 | 93.1
40C 7 days | THC-COOH | HQC | 160 | 2 | 42512 | 50920 | 148.413 | 92.8
40C 7 days | THC-COOH | HQC | 160 | 3 | 40155 | 48269 | 147.472 | 92.2
40C 7 days | THC-COOH | HQC | 160 | 4 | 40439 | 48565 | 149.395 | 93.4
40C 7 days | THC-COOH | HQC | 160 | 5 | 42015 | 53240 | 149.218 | 93.3
40C 7 days | THC-COOH | HQC | 160 | 6 | 42132 | 50106 | 149.886 | 93.7
Freeze-thaw | CBD | LQC | 6 | 1 | 1516 | 47637 | 5.716 | 95.3
Freeze-thaw | CBD | LQC | 6 | 2 | 1526 | 47166 | 5.682 | 94.7
Freeze-thaw | CBD | LQC | 6 | 3 | 1507 | 47056 | 5.654 | 94.2
Freeze-thaw | CBD | LQC | 6 | 4 | 1481 | 48035 | 5.613 | 93.6
Freeze-thaw | CBD | LQC | 6 | 5 | 1510 | 48592 | 5.726 | 95.4
Freeze-thaw | CBD | LQC | 6 | 6 | 1522 | 49734 | 5.71 | 95.2
Freeze-thaw | CBD | HQC | 160 | 1 | 39114 | 46362 | 151.056 | 94.4
Freeze-thaw | CBD | HQC | 160 | 2 | 39619 | 47239 | 150.579 | 94.1
Freeze-thaw | CBD | HQC | 160 | 3 | 39846 | 48661 | 151.126 | 94.5
Freeze-thaw | CBD | HQC | 160 | 4 | 38643 | 48080 | 151.775 | 94.9
Freeze-thaw | CBD | HQC | 160 | 5 | 38639 | 47820 | 149.621 | 93.5
Freeze-thaw | CBD | HQC | 160 | 6 | 40993 | 49910 | 151.946 | 95
Freeze-thaw | 7-OH-CBD | LQC | 6 | 1 | 1358 | 45705 | 5.58 | 93
Freeze-thaw | 7-OH-CBD | LQC | 6 | 2 | 1345 | 43298 | 5.573 | 92.9
Freeze-thaw | 7-OH-CBD | LQC | 6 | 3 | 1355 | 44194 | 5.57 | 92.8
Freeze-thaw | 7-OH-CBD | LQC | 6 | 4 | 1386 | 44499 | 5.592 | 93.2
Freeze-thaw | 7-OH-CBD | LQC | 6 | 5 | 1307 | 42301 | 5.645 | 94.1
Freeze-thaw | 7-OH-CBD | LQC | 6 | 6 | 1364 | 42436 | 5.644 | 94.1
Freeze-thaw | 7-OH-CBD | HQC | 160 | 1 | 36120 | 41938 | 147.375 | 92.1
Freeze-thaw | 7-OH-CBD | HQC | 160 | 2 | 34769 | 43340 | 150.252 | 93.9
Freeze-thaw | 7-OH-CBD | HQC | 160 | 3 | 35039 | 44448 | 148.305 | 92.7
Freeze-thaw | 7-OH-CBD | HQC | 160 | 4 | 36639 | 43982 | 149.393 | 93.4
Freeze-thaw | 7-OH-CBD | HQC | 160 | 5 | 37019 | 42886 | 149.298 | 93.3
Freeze-thaw | 7-OH-CBD | HQC | 160 | 6 | 35530 | 43968 | 148.884 | 93.1
Freeze-thaw | 7-COOH-CBD | LQC | 15 | 1 | 4030 | 50803 | 14.052 | 93.7
Freeze-thaw | 7-COOH-CBD | LQC | 15 | 2 | 4010 | 50145 | 14.357 | 95.7
Freeze-thaw | 7-COOH-CBD | LQC | 15 | 3 | 4071 | 51716 | 14.323 | 95.5
Freeze-thaw | 7-COOH-CBD | LQC | 15 | 4 | 4238 | 51181 | 14.262 | 95.1
Freeze-thaw | 7-COOH-CBD | LQC | 15 | 5 | 4247 | 51340 | 14.211 | 94.7
Freeze-thaw | 7-COOH-CBD | LQC | 15 | 6 | 4058 | 53003 | 14.03 | 93.5
Freeze-thaw | 7-COOH-CBD | HQC | 400 | 1 | 109769 | 54111 | 381.3 | 95.3
Freeze-thaw | 7-COOH-CBD | HQC | 400 | 2 | 106169 | 52621 | 373.564 | 93.4
Freeze-thaw | 7-COOH-CBD | HQC | 400 | 3 | 109750 | 52184 | 383.978 | 96
Freeze-thaw | 7-COOH-CBD | HQC | 400 | 4 | 106598 | 50178 | 377.095 | 94.3
Freeze-thaw | 7-COOH-CBD | HQC | 400 | 5 | 107531 | 51111 | 380.918 | 95.2
Freeze-thaw | 7-COOH-CBD | HQC | 400 | 6 | 110286 | 53603 | 378.42 | 94.6
Freeze-thaw | 6α-OH-CBD | LQC | 1.5 | 1 | 285 | 37252 | 1.389 | 92.6
Freeze-thaw | 6α-OH-CBD | LQC | 1.5 | 2 | 293 | 37681 | 1.396 | 93.1
Freeze-thaw | 6α-OH-CBD | LQC | 1.5 | 3 | 296 | 38549 | 1.394 | 92.9
Freeze-thaw | 6α-OH-CBD | LQC | 1.5 | 4 | 301 | 37364 | 1.405 | 93.7
Freeze-thaw | 6α-OH-CBD | LQC | 1.5 | 5 | 298 | 38265 | 1.376 | 91.8
Freeze-thaw | 6α-OH-CBD | LQC | 1.5 | 6 | 297 | 38531 | 1.382 | 92.2
Freeze-thaw | 6α-OH-CBD | HQC | 80 | 1 | 15123 | 37170 | 73.092 | 91.4
Freeze-thaw | 6α-OH-CBD | HQC | 80 | 2 | 14864 | 40031 | 71.917 | 89.9
Freeze-thaw | 6α-OH-CBD | HQC | 80 | 3 | 15521 | 37726 | 73.198 | 91.5
Freeze-thaw | 6α-OH-CBD | HQC | 80 | 4 | 15432 | 38850 | 73.958 | 92.4
Freeze-thaw | 6α-OH-CBD | HQC | 80 | 5 | 15375 | 38279 | 73.237 | 91.5
Freeze-thaw | 6α-OH-CBD | HQC | 80 | 6 | 15926 | 38485 | 74.539 | 93.2
Freeze-thaw | CBN | LQC | 1.5 | 1 | 319 | 41220 | 1.428 | 95.2
Freeze-thaw | CBN | LQC | 1.5 | 2 | 308 | 40415 | 1.401 | 93.4
Freeze-thaw | CBN | LQC | 1.5 | 3 | 321 | 40609 | 1.411 | 94.1
Freeze-thaw | CBN | LQC | 1.5 | 4 | 325 | 40321 | 1.419 | 94.6
Freeze-thaw | CBN | LQC | 1.5 | 5 | 319 | 42705 | 1.422 | 94.8
Freeze-thaw | CBN | LQC | 1.5 | 6 | 312 | 40783 | 1.394 | 92.9
Freeze-thaw | CBN | HQC | 80 | 1 | 17024 | 40700 | 76.656 | 95.8
Freeze-thaw | CBN | HQC | 80 | 2 | 16388 | 39782 | 74.753 | 93.4
Freeze-thaw | CBN | HQC | 80 | 3 | 17255 | 40612 | 75.241 | 94.1
Freeze-thaw | CBN | HQC | 80 | 4 | 17051 | 38756 | 75.979 | 95
Freeze-thaw | CBN | HQC | 80 | 5 | 17009 | 39365 | 76.586 | 95.7
Freeze-thaw | CBN | HQC | 80 | 6 | 17048 | 39183 | 75.691 | 94.6
Freeze-thaw | CBG | LQC | 1.5 | 1 | 257 | 35381 | 1.339 | 89.3
Freeze-thaw | CBG | LQC | 1.5 | 2 | 259 | 35657 | 1.335 | 89
Freeze-thaw | CBG | LQC | 1.5 | 3 | 267 | 34636 | 1.364 | 90.9
Freeze-thaw | CBG | LQC | 1.5 | 4 | 256 | 33709 | 1.354 | 90.2
Freeze-thaw | CBG | LQC | 1.5 | 5 | 255 | 36148 | 1.321 | 88
Freeze-thaw | CBG | LQC | 1.5 | 6 | 269 | 34486 | 1.375 | 91.7
Freeze-thaw | CBG | HQC | 80 | 1 | 14716 | 36825 | 72.623 | 90.8
Freeze-thaw | CBG | HQC | 80 | 2 | 14293 | 35931 | 72.384 | 90.5
Freeze-thaw | CBG | HQC | 80 | 3 | 13875 | 34498 | 72.077 | 90.1
Freeze-thaw | CBG | HQC | 80 | 4 | 13859 | 34845 | 71.518 | 89.4
Freeze-thaw | CBG | HQC | 80 | 5 | 14093 | 35869 | 72.018 | 90
Freeze-thaw | CBG | HQC | 80 | 6 | 14131 | 33961 | 72.577 | 90.7
Freeze-thaw | Δ9-THC | LQC | 1.5 | 1 | 351 | 42816 | 1.397 | 93.1
Freeze-thaw | Δ9-THC | LQC | 1.5 | 2 | 358 | 46770 | 1.405 | 93.7
Freeze-thaw | Δ9-THC | LQC | 1.5 | 3 | 358 | 47169 | 1.393 | 92.8
Freeze-thaw | Δ9-THC | LQC | 1.5 | 4 | 350 | 44735 | 1.415 | 94.3
Freeze-thaw | Δ9-THC | LQC | 1.5 | 5 | 365 | 46273 | 1.402 | 93.5
Freeze-thaw | Δ9-THC | LQC | 1.5 | 6 | 352 | 44428 | 1.429 | 95.2
Freeze-thaw | Δ9-THC | HQC | 80 | 1 | 19043 | 42930 | 75.146 | 93.9
Freeze-thaw | Δ9-THC | HQC | 80 | 2 | 18336 | 44454 | 74.929 | 93.7
Freeze-thaw | Δ9-THC | HQC | 80 | 3 | 18603 | 44376 | 75.425 | 94.3
Freeze-thaw | Δ9-THC | HQC | 80 | 4 | 18949 | 44967 | 72.885 | 91.1
Freeze-thaw | Δ9-THC | HQC | 80 | 5 | 19012 | 44465 | 76.287 | 95.4
Freeze-thaw | Δ9-THC | HQC | 80 | 6 | 19098 | 43831 | 75.392 | 94.2
Freeze-thaw | 11-OH-THC | LQC | 1.5 | 1 | 329 | 43339 | 1.4 | 93.3
Freeze-thaw | 11-OH-THC | LQC | 1.5 | 2 | 322 | 41041 | 1.386 | 92.4
Freeze-thaw | 11-OH-THC | LQC | 1.5 | 3 | 332 | 42401 | 1.386 | 92.4
Freeze-thaw | 11-OH-THC | LQC | 1.5 | 4 | 323 | 43630 | 1.378 | 91.9
Freeze-thaw | 11-OH-THC | LQC | 1.5 | 5 | 313 | 41184 | 1.389 | 92.6
Freeze-thaw | 11-OH-THC | LQC | 1.5 | 6 | 328 | 42804 | 1.401 | 93.4
Freeze-thaw | 11-OH-THC | HQC | 80 | 1 | 16991 | 42563 | 74.13 | 92.7
Freeze-thaw | 11-OH-THC | HQC | 80 | 2 | 16987 | 42867 | 73.688 | 92.1
Freeze-thaw | 11-OH-THC | HQC | 80 | 3 | 17670 | 40518 | 74.297 | 92.9
Freeze-thaw | 11-OH-THC | HQC | 80 | 4 | 17831 | 40236 | 74.054 | 92.6
Freeze-thaw | 11-OH-THC | HQC | 80 | 5 | 17406 | 41115 | 73.536 | 91.9
Freeze-thaw | 11-OH-THC | HQC | 80 | 6 | 17692 | 40890 | 74.931 | 93.7
Freeze-thaw | THC-COOH | LQC | 6 | 1 | 1620 | 50901 | 5.736 | 95.6
Freeze-thaw | THC-COOH | LQC | 6 | 2 | 1605 | 48924 | 5.691 | 94.9
Freeze-thaw | THC-COOH | LQC | 6 | 3 | 1603 | 46202 | 5.651 | 94.2
Freeze-thaw | THC-COOH | LQC | 6 | 4 | 1593 | 46928 | 5.764 | 96.1
Freeze-thaw | THC-COOH | LQC | 6 | 5 | 1619 | 52674 | 5.747 | 95.8
Freeze-thaw | THC-COOH | LQC | 6 | 6 | 1636 | 49762 | 5.67 | 94.5
Freeze-thaw | THC-COOH | HQC | 160 | 1 | 42245 | 48961 | 151.336 | 94.6
Freeze-thaw | THC-COOH | HQC | 160 | 2 | 43726 | 50159 | 150.868 | 94.3
Freeze-thaw | THC-COOH | HQC | 160 | 3 | 42805 | 49671 | 152.812 | 95.5
Freeze-thaw | THC-COOH | HQC | 160 | 4 | 42163 | 51148 | 149.939 | 93.7
Freeze-thaw | THC-COOH | HQC | 160 | 5 | 43046 | 54151 | 152.045 | 95
Freeze-thaw | THC-COOH | HQC | 160 | 6 | 43450 | 49445 | 151.175 | 94.5
Autosampler 24h | CBD | LQC | 6 | 1 | 1557 | 46205 | 5.803 | 96.7
Autosampler 24h | CBD | LQC | 6 | 2 | 1549 | 50337 | 5.757 | 96
Autosampler 24h | CBD | LQC | 6 | 3 | 1525 | 49598 | 5.77 | 96.2
Autosampler 24h | CBD | LQC | 6 | 4 | 1519 | 48604 | 5.845 | 97.4
Autosampler 24h | CBD | LQC | 6 | 5 | 1511 | 49285 | 5.759 | 96
Autosampler 24h | CBD | LQC | 6 | 6 | 1535 | 46808 | 5.773 | 96.2
Autosampler 24h | CBD | HQC | 160 | 1 | 41663 | 46667 | 157.748 | 98.6
Autosampler 24h | CBD | HQC | 160 | 2 | 41717 | 47900 | 156.743 | 98
Autosampler 24h | CBD | HQC | 160 | 3 | 40234 | 48534 | 152.918 | 95.6
Autosampler 24h | CBD | HQC | 160 | 4 | 43126 | 49481 | 156.084 | 97.6
Autosampler 24h | CBD | HQC | 160 | 5 | 41183 | 48406 | 155.456 | 97.2
Autosampler 24h | CBD | HQC | 160 | 6 | 40443 | 48422 | 154.014 | 96.3
Autosampler 24h | 7-OH-CBD | LQC | 6 | 1 | 1373 | 45507 | 5.764 | 96.1
Autosampler 24h | 7-OH-CBD | LQC | 6 | 2 | 1409 | 45506 | 5.82 | 97
Autosampler 24h | 7-OH-CBD | LQC | 6 | 3 | 1398 | 43304 | 5.632 | 93.9
Autosampler 24h | 7-OH-CBD | LQC | 6 | 4 | 1383 | 42856 | 5.753 | 95.9
Autosampler 24h | 7-OH-CBD | LQC | 6 | 5 | 1402 | 44625 | 5.802 | 96.7
Autosampler 24h | 7-OH-CBD | LQC | 6 | 6 | 1368 | 43551 | 5.733 | 95.5
Autosampler 24h | 7-OH-CBD | HQC | 160 | 1 | 38674 | 43631 | 155.919 | 97.4
Autosampler 24h | 7-OH-CBD | HQC | 160 | 2 | 37516 | 45142 | 155.774 | 97.4
Autosampler 24h | 7-OH-CBD | HQC | 160 | 3 | 36785 | 43972 | 154.059 | 96.3
Autosampler 24h | 7-OH-CBD | HQC | 160 | 4 | 38163 | 42651 | 153.938 | 96.2
Autosampler 24h | 7-OH-CBD | HQC | 160 | 5 | 37259 | 44230 | 152.727 | 95.5
Autosampler 24h | 7-OH-CBD | HQC | 160 | 6 | 36829 | 44671 | 153.043 | 95.7
Autosampler 24h | 7-COOH-CBD | LQC | 15 | 1 | 4201 | 50180 | 14.571 | 97.1
Autosampler 24h | 7-COOH-CBD | LQC | 15 | 2 | 4184 | 50647 | 14.725 | 98.2
Autosampler 24h | 7-COOH-CBD | LQC | 15 | 3 | 4203 | 52242 | 14.519 | 96.8
Autosampler 24h | 7-COOH-CBD | LQC | 15 | 4 | 4259 | 52858 | 14.516 | 96.8
Autosampler 24h | 7-COOH-CBD | LQC | 15 | 5 | 4223 | 51673 | 14.504 | 96.7
Autosampler 24h | 7-COOH-CBD | LQC | 15 | 6 | 4160 | 51820 | 14.719 | 98.1
Autosampler 24h | 7-COOH-CBD | HQC | 400 | 1 | 107672 | 51837 | 389.464 | 97.4
Autosampler 24h | 7-COOH-CBD | HQC | 400 | 2 | 111585 | 51276 | 388.464 | 97.1
Autosampler 24h | 7-COOH-CBD | HQC | 400 | 3 | 112616 | 52257 | 386.479 | 96.6
Autosampler 24h | 7-COOH-CBD | HQC | 400 | 4 | 109830 | 52693 | 383.523 | 95.9
Autosampler 24h | 7-COOH-CBD | HQC | 400 | 5 | 109767 | 52831 | 388.046 | 97
Autosampler 24h | 7-COOH-CBD | HQC | 400 | 6 | 113661 | 51275 | 389.383 | 97.3
Autosampler 24h | 6α-OH-CBD | LQC | 1.5 | 1 | 299 | 39816 | 1.427 | 95.1
Autosampler 24h | 6α-OH-CBD | LQC | 1.5 | 2 | 297 | 37980 | 1.417 | 94.5
Autosampler 24h | 6α-OH-CBD | LQC | 1.5 | 3 | 296 | 36403 | 1.433 | 95.6
Autosampler 24h | 6α-OH-CBD | LQC | 1.5 | 4 | 305 | 38938 | 1.407 | 93.8
Autosampler 24h | 6α-OH-CBD | LQC | 1.5 | 5 | 299 | 37833 | 1.421 | 94.7
Autosampler 24h | 6α-OH-CBD | LQC | 1.5 | 6 | 304 | 37083 | 1.439 | 95.9
Autosampler 24h | 6α-OH-CBD | HQC | 80 | 1 | 16204 | 38495 | 76.58 | 95.7
Autosampler 24h | 6α-OH-CBD | HQC | 80 | 2 | 16350 | 35329 | 76.599 | 95.7
Autosampler 24h | 6α-OH-CBD | HQC | 80 | 3 | 16034 | 36191 | 75.344 | 94.2
Autosampler 24h | 6α-OH-CBD | HQC | 80 | 4 | 16490 | 37864 | 76.882 | 96.1
Autosampler 24h | 6α-OH-CBD | HQC | 80 | 5 | 16730 | 39299 | 77.063 | 96.3
Autosampler 24h | 6α-OH-CBD | HQC | 80 | 6 | 15647 | 38074 | 76.403 | 95.5
Autosampler 24h | CBN | LQC | 1.5 | 1 | 335 | 39123 | 1.45 | 96.7
Autosampler 24h | CBN | LQC | 1.5 | 2 | 326 | 40128 | 1.451 | 96.7
Autosampler 24h | CBN | LQC | 1.5 | 3 | 320 | 41762 | 1.454 | 96.9
Autosampler 24h | CBN | LQC | 1.5 | 4 | 316 | 39227 | 1.456 | 97.1
Autosampler 24h | CBN | LQC | 1.5 | 5 | 319 | 39961 | 1.445 | 96.3
Autosampler 24h | CBN | LQC | 1.5 | 6 | 320 | 38537 | 1.443 | 96.2
Autosampler 24h | CBN | HQC | 80 | 1 | 17024 | 40790 | 76.354 | 95.4
Autosampler 24h | CBN | HQC | 80 | 2 | 16777 | 39678 | 76.39 | 95.5
Autosampler 24h | CBN | HQC | 80 | 3 | 16729 | 41394 | 77.767 | 97.2
Autosampler 24h | CBN | HQC | 80 | 4 | 17487 | 39833 | 77.609 | 97
Autosampler 24h | CBN | HQC | 80 | 5 | 17196 | 39217 | 77.685 | 97.1
Autosampler 24h | CBN | HQC | 80 | 6 | 17128 | 41119 | 76.903 | 96.1
Autosampler 24h | CBG | LQC | 1.5 | 1 | 270 | 34944 | 1.4 | 93.3
Autosampler 24h | CBG | LQC | 1.5 | 2 | 256 | 33359 | 1.353 | 90.2
Autosampler 24h | CBG | LQC | 1.5 | 3 | 261 | 35462 | 1.378 | 91.8
Autosampler 24h | CBG | LQC | 1.5 | 4 | 271 | 34560 | 1.386 | 92.4
Autosampler 24h | CBG | LQC | 1.5 | 5 | 271 | 35816 | 1.396 | 93.1
Autosampler 24h | CBG | LQC | 1.5 | 6 | 262 | 33041 | 1.362 | 90.8
Autosampler 24h | CBG | HQC | 80 | 1 | 14254 | 34101 | 73.125 | 91.4
Autosampler 24h | CBG | HQC | 80 | 2 | 14161 | 34795 | 73.336 | 91.7
Autosampler 24h | CBG | HQC | 80 | 3 | 14060 | 35402 | 73.404 | 91.8
Autosampler 24h | CBG | HQC | 80 | 4 | 14524 | 34910 | 74.314 | 92.9
Autosampler 24h | CBG | HQC | 80 | 5 | 14666 | 35212 | 73.799 | 92.2
Autosampler 24h | CBG | HQC | 80 | 6 | 14452 | 34442 | 73.49 | 91.9
Autosampler 24h | Δ9-THC | LQC | 1.5 | 1 | 359 | 45127 | 1.426 | 95.1
Autosampler 24h | Δ9-THC | LQC | 1.5 | 2 | 350 | 44849 | 1.44 | 96
Autosampler 24h | Δ9-THC | LQC | 1.5 | 3 | 362 | 45108 | 1.435 | 95.7
Autosampler 24h | Δ9-THC | LQC | 1.5 | 4 | 354 | 46480 | 1.428 | 95.2
Autosampler 24h | Δ9-THC | LQC | 1.5 | 5 | 362 | 45499 | 1.457 | 97.1
Autosampler 24h | Δ9-THC | LQC | 1.5 | 6 | 376 | 43603 | 1.469 | 97.9
Autosampler 24h | Δ9-THC | HQC | 80 | 1 | 20075 | 44215 | 77.73 | 97.2
Autosampler 24h | Δ9-THC | HQC | 80 | 2 | 19432 | 45337 | 76.768 | 96
Autosampler 24h | Δ9-THC | HQC | 80 | 3 | 19640 | 45948 | 76.857 | 96.1
Autosampler 24h | Δ9-THC | HQC | 80 | 4 | 18481 | 46526 | 75.614 | 94.5
Autosampler 24h | Δ9-THC | HQC | 80 | 5 | 19205 | 46318 | 77.755 | 97.2
Autosampler 24h | Δ9-THC | HQC | 80 | 6 | 19236 | 43706 | 76.644 | 95.8
Autosampler 24h | 11-OH-THC | LQC | 1.5 | 1 | 344 | 42851 | 1.427 | 95.1
Autosampler 24h | 11-OH-THC | LQC | 1.5 | 2 | 340 | 41386 | 1.451 | 96.7
Autosampler 24h | 11-OH-THC | LQC | 1.5 | 3 | 337 | 41119 | 1.407 | 93.8
Autosampler 24h | 11-OH-THC | LQC | 1.5 | 4 | 339 | 40652 | 1.454 | 96.9
Autosampler 24h | 11-OH-THC | LQC | 1.5 | 5 | 332 | 42305 | 1.431 | 95.4
Autosampler 24h | 11-OH-THC | LQC | 1.5 | 6 | 338 | 41482 | 1.437 | 95.8
Autosampler 24h | 11-OH-THC | HQC | 80 | 1 | 17956 | 42007 | 76.737 | 95.9
Autosampler 24h | 11-OH-THC | HQC | 80 | 2 | 17860 | 41903 | 76.236 | 95.3
Autosampler 24h | 11-OH-THC | HQC | 80 | 3 | 18134 | 43745 | 76.122 | 95.2
Autosampler 24h | 11-OH-THC | HQC | 80 | 4 | 17839 | 42761 | 76.094 | 95.1
Autosampler 24h | 11-OH-THC | HQC | 80 | 5 | 18176 | 42403 | 77.215 | 96.5
Autosampler 24h | 11-OH-THC | HQC | 80 | 6 | 17714 | 40284 | 76.662 | 95.8
Autosampler 24h | THC-COOH | LQC | 6 | 1 | 1671 | 50185 | 5.756 | 95.9
Autosampler 24h | THC-COOH | LQC | 6 | 2 | 1692 | 50829 | 5.956 | 99.3
Autosampler 24h | THC-COOH | LQC | 6 | 3 | 1588 | 49294 | 5.721 | 95.4
Autosampler 24h | THC-COOH | LQC | 6 | 4 | 1616 | 49781 | 5.79 | 96.5
Autosampler 24h | THC-COOH | LQC | 6 | 5 | 1577 | 50972 | 5.711 | 95.2
Autosampler 24h | THC-COOH | LQC | 6 | 6 | 1617 | 52696 | 5.84 | 97.3
Autosampler 24h | THC-COOH | HQC | 160 | 1 | 43487 | 49628 | 151.429 | 94.6
Autosampler 24h | THC-COOH | HQC | 160 | 2 | 45004 | 48401 | 158.402 | 99
Autosampler 24h | THC-COOH | HQC | 160 | 3 | 45618 | 49862 | 154.637 | 96.6
Autosampler 24h | THC-COOH | HQC | 160 | 4 | 43454 | 50329 | 153.025 | 95.6
Autosampler 24h | THC-COOH | HQC | 160 | 5 | 44469 | 50127 | 157.047 | 98.2
Autosampler 24h | THC-COOH | HQC | 160 | 6 | 43308 | 50349 | 154.752 | 96.7
Pre-spotting 2h | CBD | LQC | 6 | 1 | 1543 | 47538 | 5.808 | 96.8
Pre-spotting 2h | CBD | LQC | 6 | 2 | 1521 | 49405 | 5.77 | 96.2
Pre-spotting 2h | CBD | LQC | 6 | 3 | 1472 | 49574 | 5.67 | 94.5
Pre-spotting 2h | CBD | LQC | 6 | 4 | 1551 | 47256 | 5.734 | 95.6
Pre-spotting 2h | CBD | LQC | 6 | 5 | 1519 | 51915 | 5.649 | 94.1
Pre-spotting 2h | CBD | LQC | 6 | 6 | 1535 | 47931 | 5.785 | 96.4
Pre-spotting 2h | CBD | HQC | 160 | 1 | 41255 | 48805 | 154.722 | 96.7
Pre-spotting 2h | CBD | HQC | 160 | 2 | 40560 | 47753 | 154.058 | 96.3
Pre-spotting 2h | CBD | HQC | 160 | 3 | 40988 | 49414 | 153.62 | 96
Pre-spotting 2h | CBD | HQC | 160 | 4 | 39578 | 48818 | 149.394 | 93.4
Pre-spotting 2h | CBD | HQC | 160 | 5 | 41449 | 47515 | 153.891 | 96.2
Pre-spotting 2h | CBD | HQC | 160 | 6 | 41320 | 47035 | 155.378 | 97.1
Pre-spotting 2h | 7-OH-CBD | LQC | 6 | 1 | 1360 | 42421 | 5.668 | 94.5
Pre-spotting 2h | 7-OH-CBD | LQC | 6 | 2 | 1414 | 45090 | 5.773 | 96.2
Pre-spotting 2h | 7-OH-CBD | LQC | 6 | 3 | 1383 | 44590 | 5.766 | 96.1
Pre-spotting 2h | 7-OH-CBD | LQC | 6 | 4 | 1414 | 44915 | 5.749 | 95.8
Pre-spotting 2h | 7-OH-CBD | LQC | 6 | 5 | 1414 | 44810 | 5.792 | 96.5
Pre-spotting 2h | 7-OH-CBD | LQC | 6 | 6 | 1381 | 43969 | 5.785 | 96.4
Pre-spotting 2h | 7-OH-CBD | HQC | 160 | 1 | 36539 | 43147 | 152.391 | 95.2
Pre-spotting 2h | 7-OH-CBD | HQC | 160 | 2 | 36791 | 43093 | 152.419 | 95.3
Pre-spotting 2h | 7-OH-CBD | HQC | 160 | 3 | 36348 | 44244 | 149.75 | 93.6
Pre-spotting 2h | 7-OH-CBD | HQC | 160 | 4 | 37367 | 44754 | 151.393 | 94.6
Pre-spotting 2h | 7-OH-CBD | HQC | 160 | 5 | 35949 | 43127 | 150.186 | 93.9
Pre-spotting 2h | 7-OH-CBD | HQC | 160 | 6 | 36573 | 44948 | 151.205 | 94.5
Pre-spotting 2h | 7-COOH-CBD | LQC | 15 | 1 | 4183 | 49290 | 14.377 | 95.8
Pre-spotting 2h | 7-COOH-CBD | LQC | 15 | 2 | 4116 | 52081 | 14.236 | 94.9
Pre-spotting 2h | 7-COOH-CBD | LQC | 15 | 3 | 3975 | 50230 | 14.611 | 97.4
Pre-spotting 2h | 7-COOH-CBD | LQC | 15 | 4 | 4247 | 54434 | 14.442 | 96.3
Pre-spotting 2h | 7-COOH-CBD | LQC | 15 | 5 | 4216 | 51335 | 14.445 | 96.3
Pre-spotting 2h | 7-COOH-CBD | LQC | 15 | 6 | 4144 | 52644 | 14.275 | 95.2
Pre-spotting 2h | 7-COOH-CBD | HQC | 400 | 1 | 112560 | 55731 | 388.475 | 97.1
Pre-spotting 2h | 7-COOH-CBD | HQC | 400 | 2 | 106531 | 51965 | 382.444 | 95.6
Pre-spotting 2h | 7-COOH-CBD | HQC | 400 | 3 | 108793 | 51446 | 381.73 | 95.4
Pre-spotting 2h | 7-COOH-CBD | HQC | 400 | 4 | 113001 | 50888 | 384.986 | 96.2
Pre-spotting 2h | 7-COOH-CBD | HQC | 400 | 5 | 109945 | 51987 | 384.401 | 96.1
Pre-spotting 2h | 7-COOH-CBD | HQC | 400 | 6 | 109040 | 53115 | 384.997 | 96.2
Pre-spotting 2h | 6α-OH-CBD | LQC | 1.5 | 1 | 288 | 37699 | 1.404 | 93.6
Pre-spotting 2h | 6α-OH-CBD | LQC | 1.5 | 2 | 300 | 37529 | 1.401 | 93.4
Pre-spotting 2h | 6α-OH-CBD | LQC | 1.5 | 3 | 303 | 38191 | 1.417 | 94.4
Pre-spotting 2h | 6α-OH-CBD | LQC | 1.5 | 4 | 297 | 38385 | 1.411 | 94.1
Pre-spotting 2h | 6α-OH-CBD | LQC | 1.5 | 5 | 285 | 37424 | 1.417 | 94.4
Pre-spotting 2h | 6α-OH-CBD | LQC | 1.5 | 6 | 302 | 39405 | 1.411 | 94.1
Pre-spotting 2h | 6α-OH-CBD | HQC | 80 | 1 | 16071 | 39221 | 76.088 | 95.1
Pre-spotting 2h | 6α-OH-CBD | HQC | 80 | 2 | 15832 | 37389 | 75.509 | 94.4
Pre-spotting 2h | 6α-OH-CBD | HQC | 80 | 3 | 15949 | 39189 | 75.824 | 94.8
Pre-spotting 2h | 6α-OH-CBD | HQC | 80 | 4 | 15991 | 38000 | 75.021 | 93.8
Pre-spotting 2h | 6α-OH-CBD | HQC | 80 | 5 | 15507 | 35731 | 74.272 | 92.8
Pre-spotting 2h | 6α-OH-CBD | HQC | 80 | 6 | 15656 | 37497 | 74.44 | 93.1
Pre-spotting 2h | CBN | LQC | 1.5 | 1 | 319 | 41510 | 1.428 | 95.2
Pre-spotting 2h | CBN | LQC | 1.5 | 2 | 313 | 40647 | 1.438 | 95.9
Pre-spotting 2h | CBN | LQC | 1.5 | 3 | 326 | 39921 | 1.449 | 96.6
Pre-spotting 2h | CBN | LQC | 1.5 | 4 | 320 | 38384 | 1.419 | 94.6
Pre-spotting 2h | CBN | LQC | 1.5 | 5 | 315 | 39061 | 1.439 | 96
Pre-spotting 2h | CBN | LQC | 1.5 | 6 | 332 | 38763 | 1.45 | 96.7
Pre-spotting 2h | CBN | HQC | 80 | 1 | 16774 | 40663 | 75.791 | 94.7
Pre-spotting 2h | CBN | HQC | 80 | 2 | 17080 | 39434 | 76.941 | 96.2
Pre-spotting 2h | CBN | HQC | 80 | 3 | 17261 | 38957 | 76.511 | 95.6
Pre-spotting 2h | CBN | HQC | 80 | 4 | 16636 | 39762 | 76.087 | 95.1
Pre-spotting 2h | CBN | HQC | 80 | 5 | 17311 | 39801 | 74.52 | 93.2
Pre-spotting 2h | CBN | HQC | 80 | 6 | 17079 | 39979 | 77.08 | 96.4
Pre-spotting 2h | CBG | LQC | 1.5 | 1 | 267 | 35497 | 1.4 | 93.3
Pre-spotting 2h | CBG | LQC | 1.5 | 2 | 280 | 35650 | 1.415 | 94.3
Pre-spotting 2h | CBG | LQC | 1.5 | 3 | 277 | 35017 | 1.406 | 93.7
Pre-spotting 2h | CBG | LQC | 1.5 | 4 | 272 | 34677 | 1.408 | 93.9
Pre-spotting 2h | CBG | LQC | 1.5 | 5 | 267 | 36409 | 1.401 | 93.4
Pre-spotting 2h | CBG | LQC | 1.5 | 6 | 269 | 35149 | 1.379 | 91.9
Pre-spotting 2h | CBG | HQC | 80 | 1 | 14550 | 35221 | 74.769 | 93.5
Pre-spotting 2h | CBG | HQC | 80 | 2 | 14580 | 33543 | 75.752 | 94.7
Pre-spotting 2h | CBG | HQC | 80 | 3 | 14522 | 33504 | 74.563 | 93.2
Pre-spotting 2h | CBG | HQC | 80 | 4 | 14887 | 35019 | 75.88 | 94.9
Pre-spotting 2h | CBG | HQC | 80 | 5 | 14302 | 34439 | 73.699 | 92.1
Pre-spotting 2h | CBG | HQC | 80 | 6 | 14683 | 36166 | 74.403 | 93
Pre-spotting 2h | Δ9-THC | LQC | 1.5 | 1 | 345 | 45288 | 1.36 | 90.7
Pre-spotting 2h | Δ9-THC | LQC | 1.5 | 2 | 344 | 46849 | 1.391 | 92.7
Pre-spotting 2h | Δ9-THC | LQC | 1.5 | 3 | 342 | 43528 | 1.364 | 91
Pre-spotting 2h | Δ9-THC | LQC | 1.5 | 4 | 347 | 46155 | 1.382 | 92.1
Pre-spotting 2h | Δ9-THC | LQC | 1.5 | 5 | 340 | 45177 | 1.371 | 91.4
Pre-spotting 2h | Δ9-THC | LQC | 1.5 | 6 | 354 | 43119 | 1.404 | 93.6
Pre-spotting 2h | Δ9-THC | HQC | 80 | 1 | 18792 | 45007 | 74.845 | 93.6
Pre-spotting 2h | Δ9-THC | HQC | 80 | 2 | 18872 | 44788 | 73.098 | 91.4
Pre-spotting 2h | Δ9-THC | HQC | 80 | 3 | 17611 | 46794 | 71.737 | 89.7
Pre-spotting 2h | Δ9-THC | HQC | 80 | 4 | 18629 | 44497 | 73.164 | 91.5
Pre-spotting 2h | Δ9-THC | HQC | 80 | 5 | 18850 | 45540 | 72.188 | 90.2
Pre-spotting 2h | Δ9-THC | HQC | 80 | 6 | 18014 | 45221 | 72.36 | 90.4
Pre-spotting 2h | 11-OH-THC | LQC | 1.5 | 1 | 319 | 42020 | 1.369 | 91.3
Pre-spotting 2h | 11-OH-THC | LQC | 1.5 | 2 | 322 | 43970 | 1.37 | 91.3
Pre-spotting 2h | 11-OH-THC | LQC | 1.5 | 3 | 318 | 40587 | 1.373 | 91.6
Pre-spotting 2h | 11-OH-THC | LQC | 1.5 | 4 | 325 | 43620 | 1.37 | 91.3
Pre-spotting 2h | 11-OH-THC | LQC | 1.5 | 5 | 313 | 41999 | 1.359 | 90.6
Pre-spotting 2h | 11-OH-THC | LQC | 1.5 | 6 | 320 | 42963 | 1.372 | 91.4
Pre-spotting 2h | 11-OH-THC | HQC | 80 | 1 | 16936 | 42290 | 72.378 | 90.5
Pre-spotting 2h | 11-OH-THC | HQC | 80 | 2 | 17063 | 44292 | 72.672 | 90.8
Pre-spotting 2h | 11-OH-THC | HQC | 80 | 3 | 17318 | 41292 | 72.459 | 90.6
Pre-spotting 2h | 11-OH-THC | HQC | 80 | 4 | 17527 | 41290 | 72.588 | 90.7
Pre-spotting 2h | 11-OH-THC | HQC | 80 | 5 | 16810 | 41344 | 73.752 | 92.2
Pre-spotting 2h | 11-OH-THC | HQC | 80 | 6 | 17570 | 41252 | 73.453 | 91.8
Pre-spotting 2h | THC-COOH | LQC | 6 | 1 | 1552 | 51160 | 5.635 | 93.9
Pre-spotting 2h | THC-COOH | LQC | 6 | 2 | 1602 | 50018 | 5.79 | 96.5
Pre-spotting 2h | THC-COOH | LQC | 6 | 3 | 1615 | 50892 | 5.725 | 95.4
Pre-spotting 2h | THC-COOH | LQC | 6 | 4 | 1582 | 52569 | 5.782 | 96.4
Pre-spotting 2h | THC-COOH | LQC | 6 | 5 | 1618 | 50135 | 5.841 | 97.3
Pre-spotting 2h | THC-COOH | LQC | 6 | 6 | 1623 | 49754 | 5.672 | 94.5
Pre-spotting 2h | THC-COOH | HQC | 160 | 1 | 41982 | 51964 | 153.084 | 95.7
Pre-spotting 2h | THC-COOH | HQC | 160 | 2 | 43440 | 48371 | 153.261 | 95.8
Pre-spotting 2h | THC-COOH | HQC | 160 | 3 | 41505 | 49453 | 154.639 | 96.6
Pre-spotting 2h | THC-COOH | HQC | 160 | 4 | 43564 | 50567 | 153.585 | 96
Pre-spotting 2h | THC-COOH | HQC | 160 | 5 | 44232 | 48607 | 155.174 | 97
Pre-spotting 2h | THC-COOH | HQC | 160 | 6 | 43929 | 47687 | 155.108 | 96.9
Pre-spotting 4h | CBD | LQC | 6 | 1 | 1485 | 47898 | 5.594 | 93.2
Pre-spotting 4h | CBD | LQC | 6 | 2 | 1477 | 47156 | 5.595 | 93.2
Pre-spotting 4h | CBD | LQC | 6 | 3 | 1486 | 46644 | 5.561 | 92.7
Pre-spotting 4h | CBD | LQC | 6 | 4 | 1479 | 48831 | 5.557 | 92.6
Pre-spotting 4h | CBD | LQC | 6 | 5 | 1489 | 47154 | 5.525 | 92.1
Pre-spotting 4h | CBD | LQC | 6 | 6 | 1481 | 47420 | 5.567 | 92.8
Pre-spotting 4h | CBD | HQC | 160 | 1 | 39459 | 47686 | 151.094 | 94.4
Pre-spotting 4h | CBD | HQC | 160 | 2 | 40849 | 50170 | 152.052 | 95
Pre-spotting 4h | CBD | HQC | 160 | 3 | 39020 | 46991 | 150.048 | 93.8
Pre-spotting 4h | CBD | HQC | 160 | 4 | 40789 | 48522 | 150.703 | 94.2
Pre-spotting 4h | CBD | HQC | 160 | 5 | 40590 | 46124 | 150.895 | 94.3
Pre-spotting 4h | CBD | HQC | 160 | 6 | 39360 | 47520 | 148.424 | 92.8
Pre-spotting 4h | 7-OH-CBD | LQC | 6 | 1 | 1390 | 47557 | 5.667 | 94.4
Pre-spotting 4h | 7-OH-CBD | LQC | 6 | 2 | 1367 | 44611 | 5.546 | 92.4
Pre-spotting 4h | 7-OH-CBD | LQC | 6 | 3 | 1349 | 46053 | 5.538 | 92.3
Pre-spotting 4h | 7-OH-CBD | LQC | 6 | 4 | 1365 | 44761 | 5.584 | 93.1
Pre-spotting 4h | 7-OH-CBD | LQC | 6 | 5 | 1346 | 43293 | 5.603 | 93.4
Pre-spotting 4h | 7-OH-CBD | LQC | 6 | 6 | 1352 | 44502 | 5.495 | 91.6
Pre-spotting 4h | 7-OH-CBD | HQC | 160 | 1 | 36123 | 45915 | 148.838 | 93
Pre-spotting 4h | 7-OH-CBD | HQC | 160 | 2 | 35950 | 42751 | 149.151 | 93.2
Pre-spotting 4h | 7-OH-CBD | HQC | 160 | 3 | 36020 | 44971 | 149.191 | 93.2
Pre-spotting 4h | 7-OH-CBD | HQC | 160 | 4 | 35705 | 42263 | 148.054 | 92.5
Pre-spotting 4h | 7-OH-CBD | HQC | 160 | 5 | 36609 | 44658 | 149.479 | 93.4
Pre-spotting 4h | 7-OH-CBD | HQC | 160 | 6 | 35552 | 44008 | 148.234 | 92.6
Pre-spotting 4h | 7-COOH-CBD | LQC | 15 | 1 | 4086 | 52569 | 14.078 | 93.9
Pre-spotting 4h | 7-COOH-CBD | LQC | 15 | 2 | 4080 | 49088 | 13.906 | 92.7
Pre-spotting 4h | 7-COOH-CBD | LQC | 15 | 3 | 4052 | 50833 | 13.913 | 92.8
Pre-spotting 4h | 7-COOH-CBD | LQC | 15 | 4 | 4050 | 53433 | 13.925 | 92.8
Pre-spotting 4h | 7-COOH-CBD | LQC | 15 | 5 | 4138 | 50153 | 14.076 | 93.8
Pre-spotting 4h | 7-COOH-CBD | LQC | 15 | 6 | 4043 | 51500 | 14.017 | 93.4
Pre-spotting 4h | 7-COOH-CBD | HQC | 400 | 1 | 103403 | 52552 | 366.178 | 91.5
Pre-spotting 4h | 7-COOH-CBD | HQC | 400 | 2 | 107928 | 49941 | 367.757 | 91.9
Pre-spotting 4h | 7-COOH-CBD | HQC | 400 | 3 | 105557 | 54073 | 370.454 | 92.6
Pre-spotting 4h | 7-COOH-CBD | HQC | 400 | 4 | 107774 | 50490 | 366.435 | 91.6
Pre-spotting 4h | 7-COOH-CBD | HQC | 400 | 5 | 107401 | 51258 | 372.219 | 93.1
Pre-spotting 4h | 7-COOH-CBD | HQC | 400 | 6 | 110609 | 53510 | 374.671 | 93.7
Pre-spotting 4h | 6α-OH-CBD | LQC | 1.5 | 1 | 284 | 36689 | 1.357 | 90.5
Pre-spotting 4h | 6α-OH-CBD | LQC | 1.5 | 2 | 283 | 37047 | 1.366 | 91
Pre-spotting 4h | 6α-OH-CBD | LQC | 1.5 | 3 | 288 | 37951 | 1.367 | 91.1
Pre-spotting 4h | 6α-OH-CBD | LQC | 1.5 | 4 | 280 | 37772 | 1.352 | 90.1
Pre-spotting 4h | 6α-OH-CBD | LQC | 1.5 | 5 | 280 | 39022 | 1.374 | 91.6
Pre-spotting 4h | 6α-OH-CBD | LQC | 1.5 | 6 | 281 | 36521 | 1.349 | 90
Pre-spotting 4h | 6α-OH-CBD | HQC | 80 | 1 | 15051 | 36567 | 72.579 | 90.7
Pre-spotting 4h | 6α-OH-CBD | HQC | 80 | 2 | 15183 | 38204 | 71.911 | 89.9
Pre-spotting 4h | 6α-OH-CBD | HQC | 80 | 3 | 15058 | 38174 | 72.116 | 90.1
Pre-spotting 4h | 6α-OH-CBD | HQC | 80 | 4 | 15645 | 39653 | 72.443 | 90.6
Pre-spotting 4h | 6α-OH-CBD | HQC | 80 | 5 | 15340 | 38819 | 71.696 | 89.6
Pre-spotting 4h | 6α-OH-CBD | HQC | 80 | 6 | 15561 | 38808 | 71.806 | 89.8
Pre-spotting 4h | CBN | LQC | 1.5 | 1 | 312 | 39045 | 1.378 | 91.9
Pre-spotting 4h | CBN | LQC | 1.5 | 2 | 312 | 40573 | 1.397 | 93.2
Pre-spotting 4h | CBN | LQC | 1.5 | 3 | 317 | 38620 | 1.399 | 93.3
Pre-spotting 4h | CBN | LQC | 1.5 | 4 | 313 | 39761 | 1.41 | 94
Pre-spotting 4h | CBN | LQC | 1.5 | 5 | 308 | 40534 | 1.372 | 91.5
Pre-spotting 4h | CBN | LQC | 1.5 | 6 | 308 | 41169 | 1.393 | 92.8
Pre-spotting 4h | CBN | HQC | 80 | 1 | 16632 | 38939 | 74.74 | 93.4
Pre-spotting 4h | CBN | HQC | 80 | 2 | 16407 | 40234 | 74.19 | 92.7
Pre-spotting 4h | CBN | HQC | 80 | 3 | 16394 | 40859 | 74.222 | 92.8
Pre-spotting 4h | CBN | HQC | 80 | 4 | 16685 | 41819 | 73.126 | 91.4
Pre-spotting 4h | CBN | HQC | 80 | 5 | 16184 | 40717 | 73.36 | 91.7
Pre-spotting 4h | CBN | HQC | 80 | 6 | 16520 | 40305 | 74.516 | 93.1
Pre-spotting 4h | CBG | LQC | 1.5 | 1 | 259 | 34371 | 1.348 | 89.9
Pre-spotting 4h | CBG | LQC | 1.5 | 2 | 259 | 34938 | 1.348 | 89.8
Pre-spotting 4h | CBG | LQC | 1.5 | 3 | 264 | 36353 | 1.368 | 91.2
Pre-spotting 4h | CBG | LQC | 1.5 | 4 | 268 | 35454 | 1.362 | 90.8
Pre-spotting 4h | CBG | LQC | 1.5 | 5 | 268 | 35675 | 1.36 | 90.7
Pre-spotting 4h | CBG | LQC | 1.5 | 6 | 266 | 35790 | 1.342 | 89.5
Pre-spotting 4h | CBG | HQC | 80 | 1 | 14226 | 35034 | 71.927 | 89.9
Pre-spotting 4h | CBG | HQC | 80 | 2 | 14733 | 35976 | 72.328 | 90.4
Pre-spotting 4h | CBG | HQC | 80 | 3 | 14127 | 33743 | 71.91 | 89.9
Pre-spotting 4h | CBG | HQC | 80 | 4 | 14636 | 34789 | 73.003 | 91.3
Pre-spotting 4h | CBG | HQC | 80 | 5 | 13819 | 35630 | 71.51 | 89.4
Pre-spotting 4h | CBG | HQC | 80 | 6 | 14206 | 36813 | 73.726 | 92.2
Pre-spotting 4h | Δ9-THC | LQC | 1.5 | 1 | 326 | 46005 | 1.307 | 87.1
Pre-spotting 4h | Δ9-THC | LQC | 1.5 | 2 | 319 | 45029 | 1.313 | 87.6
Pre-spotting 4h | Δ9-THC | LQC | 1.5 | 3 | 328 | 44193 | 1.314 | 87.6
Pre-spotting 4h | Δ9-THC | LQC | 1.5 | 4 | 340 | 43529 | 1.333 | 88.9
Pre-spotting 4h | Δ9-THC | LQC | 1.5 | 5 | 334 | 44672 | 1.31 | 87.4
Pre-spotting 4h | Δ9-THC | LQC | 1.5 | 6 | 337 | 45667 | 1.318 | 87.9
Pre-spotting 4h | Δ9-THC | HQC | 80 | 1 | 17254 | 44283 | 69.848 | 87.3
Pre-spotting 4h | Δ9-THC | HQC | 80 | 2 | 17903 | 48450 | 69.502 | 86.9
Pre-spotting 4h | Δ9-THC | HQC | 80 | 3 | 17569 | 43452 | 69.133 | 86.4
Pre-spotting 4h | Δ9-THC | HQC | 80 | 4 | 17672 | 45642 | 70.464 | 88.1
Pre-spotting 4h | Δ9-THC | HQC | 80 | 5 | 17662 | 45383 | 69.992 | 87.5
Pre-spotting 4h | Δ9-THC | HQC | 80 | 6 | 17000 | 44731 | 68.525 | 85.7
Pre-spotting 4h | 11-OH-THC | LQC | 1.5 | 1 | 308 | 45259 | 1.298 | 86.6
Pre-spotting 4h | 11-OH-THC | LQC | 1.5 | 2 | 307 | 43006 | 1.292 | 86.1
Pre-spotting 4h | 11-OH-THC | LQC | 1.5 | 3 | 318 | 42662 | 1.301 | 86.8
Pre-spotting 4h | 11-OH-THC | LQC | 1.5 | 4 | 297 | 42517 | 1.284 | 85.6
Pre-spotting 4h | 11-OH-THC | LQC | 1.5 | 5 | 305 | 40996 | 1.303 | 86.9
Pre-spotting 4h | 11-OH-THC | LQC | 1.5 | 6 | 309 | 43065 | 1.292 | 86.1
Pre-spotting 4h | 11-OH-THC | HQC | 80 | 1 | 15944 | 42059 | 68.48 | 85.6
Pre-spotting 4h | 11-OH-THC | HQC | 80 | 2 | 16061 | 41763 | 68.48 | 85.6
Pre-spotting 4h | 11-OH-THC | HQC | 80 | 3 | 16283 | 42472 | 70.797 | 88.5
Pre-spotting 4h | 11-OH-THC | HQC | 80 | 4 | 16819 | 43115 | 68.73 | 85.9
Pre-spotting 4h | 11-OH-THC | HQC | 80 | 5 | 16695 | 39947 | 68.48 | 85.6
Pre-spotting 4h | 11-OH-THC | HQC | 80 | 6 | 16523 | 41051 | 69.389 | 86.7
Pre-spotting 4h | THC-COOH | LQC | 6 | 1 | 1525 | 50004 | 5.386 | 89.8
Pre-spotting 4h | THC-COOH | LQC | 6 | 2 | 1574 | 48681 | 5.583 | 93.1
Pre-spotting 4h | THC-COOH | LQC | 6 | 3 | 1518 | 50331 | 5.481 | 91.3
Pre-spotting 4h | THC-COOH | LQC | 6 | 4 | 1569 | 48705 | 5.509 | 91.8
Pre-spotting 4h | THC-COOH | LQC | 6 | 5 | 1572 | 49694 | 5.579 | 93
Pre-spotting 4h | THC-COOH | LQC | 6 | 6 | 1533 | 49610 | 5.452 | 90.9
Pre-spotting 4h | THC-COOH | HQC | 160 | 1 | 43753 | 48554 | 151.386 | 94.6
Pre-spotting 4h | THC-COOH | HQC | 160 | 2 | 43283 | 50364 | 148.991 | 93.1
Pre-spotting 4h | THC-COOH | HQC | 160 | 3 | 41656 | 50505 | 148.069 | 92.5
Pre-spotting 4h | THC-COOH | HQC | 160 | 4 | 41576 | 51425 | 149.325 | 93.3
Pre-spotting 4h | THC-COOH | HQC | 160 | 5 | 40403 | 49736 | 145.574 | 91
Pre-spotting 4h | THC-COOH | HQC | 160 | 6 | 41173 | 51608 | 147.426 | 92.1

## Dilution_Integrity
Dilution integrity — 3×ULOQ, 1:5 with blank extract, 6 replicates.
Analyte | Protocol | 3xULOQ (ng/mL) | Dilution | Replicate | Measured (ng/mL) | Bias (%)
Δ9-THC | 4-disc | 300 | 1:5 | 1 | 328.7 | 9.6
Δ9-THC | 4-disc | 300 | 1:5 | 2 | 329.2 | 9.7
Δ9-THC | 4-disc | 300 | 1:5 | 3 | 305.5 | 1.8
Δ9-THC | 4-disc | 300 | 1:5 | 4 | 312.5 | 4.2
Δ9-THC | 4-disc | 300 | 1:5 | 5 | 326.8 | 8.9
Δ9-THC | 4-disc | 300 | 1:5 | 6 | 313.3 | 4.4
11-OH-THC | 4-disc | 300 | 1:5 | 1 | 303.9 | 1.3
11-OH-THC | 4-disc | 300 | 1:5 | 2 | 318.6 | 6.2
11-OH-THC | 4-disc | 300 | 1:5 | 3 | 308.2 | 2.7
11-OH-THC | 4-disc | 300 | 1:5 | 4 | 309.8 | 3.3
11-OH-THC | 4-disc | 300 | 1:5 | 5 | 315.8 | 5.3
11-OH-THC | 4-disc | 300 | 1:5 | 6 | 292.1 | -2.6
THC-COOH | 1-disc | 600 | 1:5 | 1 | 617.4 | 2.9
THC-COOH | 1-disc | 600 | 1:5 | 2 | 631.7 | 5.3
THC-COOH | 1-disc | 600 | 1:5 | 3 | 575.2 | -4.1
THC-COOH | 1-disc | 600 | 1:5 | 4 | 598.2 | -0.3
THC-COOH | 1-disc | 600 | 1:5 | 5 | 580.7 | -3.2
THC-COOH | 1-disc | 600 | 1:5 | 6 | 621.9 | 3.6
CBD | 1-disc | 600 | 1:5 | 1 | 565.3 | -5.8
CBD | 1-disc | 600 | 1:5 | 2 | 610.4 | 1.7
CBD | 1-disc | 600 | 1:5 | 3 | 602.1 | 0.4
CBD | 1-disc | 600 | 1:5 | 4 | 617.4 | 2.9
CBD | 1-disc | 600 | 1:5 | 5 | 600.8 | 0.1
CBD | 1-disc | 600 | 1:5 | 6 | 609.7 | 1.6
7-OH-CBD | 1-disc | 600 | 1:5 | 1 | 542.3 | -9.6
7-OH-CBD | 1-disc | 600 | 1:5 | 2 | 557.2 | -7.1
7-OH-CBD | 1-disc | 600 | 1:5 | 3 | 587.4 | -2.1
7-OH-CBD | 1-disc | 600 | 1:5 | 4 | 589.6 | -1.7
7-OH-CBD | 1-disc | 600 | 1:5 | 5 | 585.7 | -2.4
7-OH-CBD | 1-disc | 600 | 1:5 | 6 | 606.8 | 1.1
7-COOH-CBD | 1-disc | 1500 | 1:5 | 1 | 1539.5 | 2.6
7-COOH-CBD | 1-disc | 1500 | 1:5 | 2 | 1397 | -6.9
7-COOH-CBD | 1-disc | 1500 | 1:5 | 3 | 1534.4 | 2.3
7-COOH-CBD | 1-disc | 1500 | 1:5 | 4 | 1560.8 | 4.1
7-COOH-CBD | 1-disc | 1500 | 1:5 | 5 | 1549.2 | 3.3
7-COOH-CBD | 1-disc | 1500 | 1:5 | 6 | 1517 | 1.1
6α-OH-CBD | 4-disc | 300 | 1:5 | 1 | 300.2 | 0.1
6α-OH-CBD | 4-disc | 300 | 1:5 | 2 | 305.4 | 1.8
6α-OH-CBD | 4-disc | 300 | 1:5 | 3 | 289.6 | -3.5
6α-OH-CBD | 4-disc | 300 | 1:5 | 4 | 298.6 | -0.5
6α-OH-CBD | 4-disc | 300 | 1:5 | 5 | 322.3 | 7.4
6α-OH-CBD | 4-disc | 300 | 1:5 | 6 | 288.3 | -3.9
CBN | 4-disc | 300 | 1:5 | 1 | 288.5 | -3.8
CBN | 4-disc | 300 | 1:5 | 2 | 305.9 | 2
CBN | 4-disc | 300 | 1:5 | 3 | 304.2 | 1.4
CBN | 4-disc | 300 | 1:5 | 4 | 281.4 | -6.2
CBN | 4-disc | 300 | 1:5 | 5 | 302.6 | 0.9
CBN | 4-disc | 300 | 1:5 | 6 | 290.3 | -3.2
CBG | 4-disc | 300 | 1:5 | 1 | 297.1 | -1
CBG | 4-disc | 300 | 1:5 | 2 | 296.1 | -1.3
CBG | 4-disc | 300 | 1:5 | 3 | 294.8 | -1.7
CBG | 4-disc | 300 | 1:5 | 4 | 292 | -2.7
CBG | 4-disc | 300 | 1:5 | 5 | 265.2 | -11.6
CBG | 4-disc | 300 | 1:5 | 6 | 296.5 | -1.2

## Selectivity_Carryover
Selectivity (6 blank donors) & carryover (post-ULOQ; +5000 ng/mL for 7-COOH-CBD).
Test | Analyte | Donor/Run | Interference (% of LLOQ) | IS signal (% of mean IS) | Result
Selectivity | Δ9-THC | Blank donor 1 | 6.1 | 0.4 | Pass
Selectivity | Δ9-THC | Blank donor 2 | 7.2 | 1.9 | Pass
Selectivity | Δ9-THC | Blank donor 3 | 7.1 | 0.6 | Pass
Selectivity | Δ9-THC | Blank donor 4 | 7.3 | 1.3 | Pass
Selectivity | Δ9-THC | Blank donor 5 | 2 | 0.9 | Pass
Selectivity | Δ9-THC | Blank donor 6 | 0.2 | 1 | Pass
Selectivity | 11-OH-THC | Blank donor 1 | 1.5 | 0.2 | Pass
Selectivity | 11-OH-THC | Blank donor 2 | 7.6 | 0.2 | Pass
Selectivity | 11-OH-THC | Blank donor 3 | 1.5 | 0.7 | Pass
Selectivity | 11-OH-THC | Blank donor 4 | 0.8 | 2.7 | Pass
Selectivity | 11-OH-THC | Blank donor 5 | 3.4 | 0.3 | Pass
Selectivity | 11-OH-THC | Blank donor 6 | 5.7 | 2.4 | Pass
Selectivity | THC-COOH | Blank donor 1 | 7.8 | 0.2 | Pass
Selectivity | THC-COOH | Blank donor 2 | 1.3 | 0.3 | Pass
Selectivity | THC-COOH | Blank donor 3 | 5.5 | 0.1 | Pass
Selectivity | THC-COOH | Blank donor 4 | 2.5 | 0.3 | Pass
Selectivity | THC-COOH | Blank donor 5 | 0.2 | 2.1 | Pass
Selectivity | THC-COOH | Blank donor 6 | 4.4 | 2.2 | Pass
Selectivity | CBD | Blank donor 1 | 3.5 | 2 | Pass
Selectivity | CBD | Blank donor 2 | 6.6 | 2.5 | Pass
Selectivity | CBD | Blank donor 3 | 2.5 | 1.7 | Pass
Selectivity | CBD | Blank donor 4 | 3.7 | 1 | Pass
Selectivity | CBD | Blank donor 5 | 3.9 | 0.8 | Pass
Selectivity | CBD | Blank donor 6 | 4.5 | 0.3 | Pass
Selectivity | 7-OH-CBD | Blank donor 1 | 7.9 | 1.2 | Pass
Selectivity | 7-OH-CBD | Blank donor 2 | 2.1 | 0.2 | Pass
Selectivity | 7-OH-CBD | Blank donor 3 | 4.5 | 0.5 | Pass
Selectivity | 7-OH-CBD | Blank donor 4 | 5.4 | 2.6 | Pass
Selectivity | 7-OH-CBD | Blank donor 5 | 2.9 | 1.1 | Pass
Selectivity | 7-OH-CBD | Blank donor 6 | 0.8 | 0.8 | Pass
Selectivity | 7-COOH-CBD | Blank donor 1 | 5.7 | 3 | Pass
Selectivity | 7-COOH-CBD | Blank donor 2 | 4.7 | 1.1 | Pass
Selectivity | 7-COOH-CBD | Blank donor 3 | 7 | 2.5 | Pass
Selectivity | 7-COOH-CBD | Blank donor 4 | 4.5 | 2.7 | Pass
Selectivity | 7-COOH-CBD | Blank donor 5 | 3.8 | 2.1 | Pass
Selectivity | 7-COOH-CBD | Blank donor 6 | 7.2 | 0.4 | Pass
Selectivity | 6α-OH-CBD | Blank donor 1 | 7.2 | 0.2 | Pass
Selectivity | 6α-OH-CBD | Blank donor 2 | 2.1 | 1.4 | Pass
Selectivity | 6α-OH-CBD | Blank donor 3 | 6.7 | 0.5 | Pass
Selectivity | 6α-OH-CBD | Blank donor 4 | 5.8 | 0.9 | Pass
Selectivity | 6α-OH-CBD | Blank donor 5 | 2.2 | 0.2 | Pass
Selectivity | 6α-OH-CBD | Blank donor 6 | 5.1 | 0.8 | Pass
Selectivity | CBN | Blank donor 1 | 1.9 | 2.9 | Pass
Selectivity | CBN | Blank donor 2 | 5.9 | 1.1 | Pass
Selectivity | CBN | Blank donor 3 | 3.6 | 2 | Pass
Selectivity | CBN | Blank donor 4 | 6 | 2.1 | Pass
Selectivity | CBN | Blank donor 5 | 2.4 | 1 | Pass
Selectivity | CBN | Blank donor 6 | 5.2 | 1.5 | Pass
Selectivity | CBG | Blank donor 1 | 3 | 0.5 | Pass
Selectivity | CBG | Blank donor 2 | 2.3 | 1.5 | Pass
Selectivity | CBG | Blank donor 3 | 2.4 | 2.6 | Pass
Selectivity | CBG | Blank donor 4 | 7.3 | 2.7 | Pass
Selectivity | CBG | Blank donor 5 | 2.7 | 1 | Pass
Selectivity | CBG | Blank donor 6 | 7.9 | 0.2 | Pass
Carryover post-ULOQ | Δ9-THC | Blank after ULOQ | 2.9 | 1.1 | Pass
Carryover post-ULOQ | 11-OH-THC | Blank after ULOQ | 1.4 | 1.7 | Pass
Carryover post-ULOQ | THC-COOH | Blank after ULOQ | 1.4 | 1.8 | Pass
Carryover post-ULOQ | CBD | Blank after ULOQ | 1.4 | 1.7 | Pass
Carryover post-ULOQ | 7-OH-CBD | Blank after ULOQ | 2.6 | 0.7 | Pass
Carryover post-ULOQ | 7-COOH-CBD | Blank after ULOQ | 0.3 | 0.3 | Pass
Carryover post-ULOQ | 6α-OH-CBD | Blank after ULOQ | 2.4 | 1.8 | Pass
Carryover post-ULOQ | CBN | Blank after ULOQ | 1.7 | 0.5 | Pass
Carryover post-ULOQ | CBG | Blank after ULOQ | 0.6 | 1.3 | Pass
Carryover 5000 ng/mL | 7-COOH-CBD | Blank after 5000 | 0.8 | 0.8 | Pass

## Application_hemaPEN_vs_SPE
Application to authentic whole blood specimens — hemaPEN cDBS vs SPE plasma-based reference (whole-blood-equivalent).
Specimen | Analyte | Protocol | hemaPEN cDBS (ng/mL) | SPE plasma ref, WB-equiv (ng/mL) | Bias (%) | Detected
1 | Δ9-THC | Four-disc | 1.4 | 1.37 | 2.3 | Yes
2 | Δ9-THC | Four-disc | 2.2 | 2.14 | 3 | Yes
3 | Δ9-THC | Four-disc | <LLOQ |  |  | No
4 | Δ9-THC | Four-disc | 23.9 | 23.3 | 2.4 | Yes
5 | Δ9-THC | Four-disc | 13.2 | 12.5 | 5.5 | Yes
6 | Δ9-THC | Four-disc | 41.9 | 40.5 | 3.4 | Yes
7 | Δ9-THC | Four-disc | 3.8 | 3.58 | 6.1 | Yes
8 | Δ9-THC | Four-disc | 22.7 | 21.7 | 4.4 | Yes
9 | Δ9-THC | Four-disc | <LLOQ |  |  | No
10 | Δ9-THC | Four-disc | 42.3 | 39.7 | 6.5 | Yes
11 | Δ9-THC | Four-disc | <LLOQ |  |  | No
12 | Δ9-THC | Four-disc | <LLOQ |  |  | No
1 | 11-OH-THC | Four-disc | <LLOQ |  |  | No
2 | 11-OH-THC | Four-disc | 0.6 | 0.63 | -4.7 | Yes
3 | 11-OH-THC | Four-disc | 14 | 15.2 | -7.9 | Yes
4 | 11-OH-THC | Four-disc | <LLOQ |  |  | No
5 | 11-OH-THC | Four-disc | 1 | 1.05 | -4.8 | Yes
6 | 11-OH-THC | Four-disc | 9.7 | 10.49 | -7.5 | Yes
7 | 11-OH-THC | Four-disc | <LLOQ |  |  | No
8 | 11-OH-THC | Four-disc | <LLOQ |  |  | No
9 | 11-OH-THC | Four-disc | 4.3 | 4.4 | -2.2 | Yes
10 | 11-OH-THC | Four-disc | <LLOQ |  |  | No
11 | 11-OH-THC | Four-disc | 14.8 | 16.1 | -7.9 | Yes
12 | 11-OH-THC | Four-disc | <LLOQ |  |  | No
1 | THC-COOH | Single-disc | 5.8 | 5.57 | 4.2 | Yes
2 | THC-COOH | Single-disc | 75.6 | 74.7 | 1.2 | Yes
3 | THC-COOH | Single-disc | 30.3 | 29.5 | 2.7 | Yes
4 | THC-COOH | Single-disc | 71.1 | 69.7 | 2 | Yes
5 | THC-COOH | Single-disc | 18.1 | 17.8 | 1.5 | Yes
6 | THC-COOH | Single-disc | 134.7 | 131.8 | 2.2 | Yes
7 | THC-COOH | Single-disc | 125.2 | 120.6 | 3.8 | Yes
8 | THC-COOH | Single-disc | 81.3 | 79.5 | 2.3 | Yes
9 | THC-COOH | Single-disc | 47.2 | 45.9 | 2.8 | Yes
10 | THC-COOH | Single-disc | 131.1 | 128.9 | 1.7 | Yes
11 | THC-COOH | Single-disc | 84.7 | 81.8 | 3.5 | Yes
12 | THC-COOH | Single-disc | 143.7 | 137.6 | 4.4 | Yes
1 | CBD | Single-disc | <LLOQ |  |  | No
2 | CBD | Single-disc | <LLOQ |  |  | No
3 | CBD | Single-disc | <LLOQ |  |  | No
4 | CBD | Single-disc | 2.3 | 2.14 | 7.4 | Yes
5 | CBD | Single-disc | 9 | 8.71 | 3.3 | Yes
6 | CBD | Single-disc | <LLOQ |  |  | No
7 | CBD | Single-disc | <LLOQ |  |  | No
8 | CBD | Single-disc | <LLOQ |  |  | No
9 | CBD | Single-disc | 18.6 | 17.3 | 7.6 | Yes
10 | CBD | Single-disc | <LLOQ |  |  | No
11 | CBD | Single-disc | <LLOQ |  |  | No
12 | CBD | Single-disc | <LLOQ |  |  | No
1 | 7-OH-CBD | Single-disc | <LLOQ |  |  | No
2 | 7-OH-CBD | Single-disc | <LLOQ |  |  | No
3 | 7-OH-CBD | Single-disc | <LLOQ |  |  | No
4 | 7-OH-CBD | Single-disc | <LLOQ |  |  | No
5 | 7-OH-CBD | Single-disc | 4.1 | 4.18 | -1.9 | Yes
6 | 7-OH-CBD | Single-disc | <LLOQ |  |  | No
7 | 7-OH-CBD | Single-disc | 16.3 | 16.4 | -0.8 | Yes
8 | 7-OH-CBD | Single-disc | <LLOQ |  |  | No
9 | 7-OH-CBD | Single-disc | <LLOQ |  |  | No
10 | 7-OH-CBD | Single-disc | 31.2 | 31.6 | -1.4 | Yes
11 | 7-OH-CBD | Single-disc | <LLOQ |  |  | No
12 | 7-OH-CBD | Single-disc | <LLOQ |  |  | No
1 | 7-COOH-CBD | Single-disc | 8.4 | 8.15 | 3.1 | Yes
2 | 7-COOH-CBD | Single-disc | <LLOQ |  |  | No
3 | 7-COOH-CBD | Single-disc | <LLOQ |  |  | No
4 | 7-COOH-CBD | Single-disc | 22.9 | 22.5 | 1.6 | Yes
5 | 7-COOH-CBD | Single-disc | 91.2 | 89.9 | 1.5 | Yes
6 | 7-COOH-CBD | Single-disc | <LLOQ |  |  | No
7 | 7-COOH-CBD | Single-disc | <LLOQ |  |  | No
8 | 7-COOH-CBD | Single-disc | <LLOQ |  |  | No
9 | 7-COOH-CBD | Single-disc | <LLOQ |  |  | No
10 | 7-COOH-CBD | Single-disc | <LLOQ |  |  | No
11 | 7-COOH-CBD | Single-disc | 187.3 | 181.7 | 3.1 | Yes
12 | 7-COOH-CBD | Single-disc | <LLOQ |  |  | No
1 | 6α-OH-CBD | Four-disc | <LLOQ |  |  | No
2 | 6α-OH-CBD | Four-disc | <LLOQ |  |  | No
3 | 6α-OH-CBD | Four-disc | <LLOQ |  |  | No
4 | 6α-OH-CBD | Four-disc | <LLOQ |  |  | No
5 | 6α-OH-CBD | Four-disc | 0.6 | 0.57 | 4.5 | Yes
6 | 6α-OH-CBD | Four-disc | <LLOQ |  |  | No
7 | 6α-OH-CBD | Four-disc | <LLOQ |  |  | No
8 | 6α-OH-CBD | Four-disc | <LLOQ |  |  | No
9 | 6α-OH-CBD | Four-disc | 4.8 | 4.68 | 2.5 | Yes
10 | 6α-OH-CBD | Four-disc | <LLOQ |  |  | No
11 | 6α-OH-CBD | Four-disc | <LLOQ |  |  | No
12 | 6α-OH-CBD | Four-disc | <LLOQ |  |  | No
1 | CBN | Four-disc | <LLOQ |  |  | No
2 | CBN | Four-disc | <LLOQ |  |  | No
3 | CBN | Four-disc | 0.7 | 0.77 | -8.6 | Yes
4 | CBN | Four-disc | <LLOQ |  |  | No
5 | CBN | Four-disc | <LLOQ |  |  | No
6 | CBN | Four-disc | 5.09 | 5.62 | -9.5 | Yes
7 | CBN | Four-disc | <LLOQ |  |  | No
8 | CBN | Four-disc | 1.56 | 1.64 | -5.1 | Yes
9 | CBN | Four-disc | 2.52 | 2.6 | -3.2 | Yes
10 | CBN | Four-disc | <LLOQ |  |  | No
11 | CBN | Four-disc | 5.1 | 5.64 | -9.6 | Yes
12 | CBN | Four-disc | <LLOQ |  |  | No
1 | CBG | Four-disc | <LLOQ |  |  | No
2 | CBG | Four-disc | <LLOQ |  |  | No
3 | CBG | Four-disc | <LLOQ |  |  | No
4 | CBG | Four-disc | 0.5 | 0.49 | 2 | Yes
5 | CBG | Four-disc | <LLOQ |  |  | No
6 | CBG | Four-disc | <LLOQ |  |  | No
7 | CBG | Four-disc | <LLOQ |  |  | No
8 | CBG | Four-disc | <LLOQ |  |  | No
9 | CBG | Four-disc | <LLOQ |  |  | No
10 | CBG | Four-disc | <LLOQ |  |  | No
11 | CBG | Four-disc | 1.8 | 1.74 | 3.6 | Yes
12 | CBG | Four-disc | <LLOQ |  |  | No